## Example RACiR Analysis
|  |  | [Header] |  |  | Example Empty Chamber Correction for RACiR
|  |  | File opened | 2016-10-03 08:55 |  | Steps
|  |  | Console s/n | 68C-571037 |  | 1. Plot A against reference CO2 for the empty chamber as in Fig. S1
|  |  | Console ver | Bluestem v.1.0.4 |  | 2. Note the data outside the linear response range, in this case outside the reference [CO2] range of ~ 25 ppm to ~ 460 ppm (data within these range are highlighted in green below)
|  |  | Scripts ver | 2016.19 |  | 3. Plot data within linear range and calculate the slope and intercept of the response as in Fig. S2
|  |  | Head s/n | 68H-581037 |  | 4. The equation in Fig. S2 provides the correction factor for the RACiR response. A needs to be corrected, then Ci needs to be corrected with the recalculated A
|  |  | Head ver | 1.0.3 |  | 5. Fig. S3 shows the corrected A (Column A) in an empty chamber - note the non-zero values for A at each end of the CO2 ramp; the correction does not work in this region.
|  |  | Head cal | {"flowbzero": "0.21169", "co2aspan2": "0", "h2oazero": "1.01655", "h2obspan2a": "0.0639438", "co2bspan2b": "0.160286", "h2oaspan2b": "0.0637607", "h2oaspan1": "0.996511", "ssb_ref": "36008.3", "ssa_ref": "38337.8", "co2azero": "0.880048", "h2obspan2": "0", "co2aspan2b": "0.159807", "chamberpressurezero": "2.54269", "co2bspan2a": "0.161614", "h2oaspanconc1": "12.19", "oxygen": "21", "tbzero": "0.12393", "co2bspan2": "0", "h2oaspan2a": "0.0639839", "h2obspan2b": "0.0637178", "tazero": "0.118774", "co2bzero": "0.891439", "flowazero": "0.38396", "h2obzero": "1.02592", "flowmeterzero": "1.01549", "co2bspanconc2": "0", "co2bspan1": "0.991781", "co2aspanconc2": "0", "co2bspanconc1": "1002", "h2obspan1": "0.996465", "co2aspan2a": "0.161122", "h2obspanconc2": "0", "h2oaspanconc2": "0", "h2obspanconc1": "12.19", "h2oaspan2": "0", "co2aspanconc1": "1002", "co2aspan1": "0.991835"} |  | 6. Fig. S4 shows the uncorrected and corrected A in a Populus deltoides leaf plotted against reference [CO2], note that the data that lie outside the range of the correction (see point 2 above) lie outside the expected A-Ci response - these data need to be excluded from analysis (data within range are highlighted in green)
|  |  | Chamber type | 1789690 |  | 7. Fig. S5 shows the uncorrected and corrected A in a Populus deltoides leaf plotted against the uncorrected and corrected Ci - note the behaviour at each end of the ramp; these data are not corrected by the linear regression and must be removed.
|  |  | Chamber s/n | MPF-551036 |  | 8. Fig. S6 shows the corrected A versus Ci response in a Populus deltoides leaf that has been cleaned of data outside the range of the correction factor. These data are ready to be used for A-Ci analysis
|  |  | Chamber rev | 0
|  |  | Chamber cal | 0
|  |  | Fluorometer | MPF-551036
|  |  | Flr. Version | 1.0.3
|  |  | 8:55 | tests
|  |  | Const:S | 6
|  |  | Const:K | 0.5
|  |  | Const:ChType | 2: Fluorometer
|  |  | Const:Geometry | 0: Broadleaf
|  |  | Const:CustomBLC | 2
|  |  | Const:Oxygen | 21
|  |  | LTConst:deltaTw | 0
|  |  | LTConst:fT1 | 1
|  |  | LTConst:fT2 | 0
|  |  | LTConst:fTeb | 0
|  |  | LQConst:Leaf | standard
|  |  | LQConst:Ambient | Sun+Sky
|  |  | LQConst:abs_blueFlrLED | 0.87
|  |  | LQConst:abs_blueLED | 0.87
|  |  | LQConst:abs_redFlrLED | 0.84
|  |  | LQConst:abs_ambient | 0.49
|  |  | LQConst:abs_greenLED | 0.7
|  |  | LQConst:abs_whiteLED | 0.75
|  |  | LQConst:abs_redLED | 0.84
|  |  | LQConst:k_blueFlrLED | 0.25
|  |  | LQConst:k_blueLED | 0.26
|  |  | LQConst:k_redFlrLED | 0.19
|  |  | LQConst:k_ambient | 0.39
|  |  | LQConst:k_greenLED | 0.23
|  |  | LQConst:k_whiteLED | 0.21
|  |  | LQConst:k_redLED | 0.18
|  |  | QConst:fQ_Amb_in | 0
|  |  | QConst:fQ_Amb_out | 0
|  |  | QConst:fQ_HeadLS | 0
|  |  | QConst:fQ_ConsoleLS | 0
|  |  | QConst:fQ_Flr | 1
[Data]
|  | Chamber (empty or with leaf) | Sys | Sys | Sys | Sys | GasEx | GasEx | GasEx | GasEx | GasEx | GasEx | GasEx | GasEx | GasEx | GasEx | GasEx | GasEx | GasEx | GasEx | GasEx | GasEx | GasEx | GasEx | GasEx | GasEx | GasEx | FLR | FLR | FLR | FLR | FLR | FLR | FLR | FLR | FLR | FLR | FLR | FLR | FLR | FLR | FLR | FLR | FLR | FLR | FLR | FLR | FLR | FLR | FLR | FLR | LeafQ | LeafQ | LeafQ | LeafQ | Meas | Meas | Meas | Meas | Meas | Meas | Meas | Meas | Meas | Meas | Meas | Meas | Meas | Meas | FlrLS | FlrLS | FlrLS | FlrLS | FlrLS | FlrLS | FlrLS | FlrLS | FlrLS | FlrLS | FlrLS | FlrLS | FlrStats | FlrStats | FlrStats | FlrStats | FlrStats | Status | Status | Status | Status | Status | Status | Status | Status | Status | Status | Status | Status | Status | Status | Status | Status | Status | Status | Status | Status
Corrected A | Corrected Ci |  | obs | time | elapsed | date | TIME | E | A | Ci | Pci | Pca | gsw | gbw | gtw | gtc | Rabs | TleafEB | TleafCnd | SVPleaf | RHcham | VPcham | SVPcham | VPDleaf | blfa_1 | blfa_2 | blfa_3 | DarkAdaptedID | Fo | Fm | Fv | Fv/Fm | Adark | LightAdaptedID | Fs | Fm' | PhiPS2 | PS2/1 | Qabs_fs | Afs | ETR | Fv'/Fm' | PhiCO2 | NPQ | DarkPulseID | Fo' | Fv' | qP | qN | qP_Fo | qN_Fo | Qin | Qabs | alpha | convert | TIME | CO2_s | CO2_r | H2O_s | H2O_r | Flow | Pa | ΔPcham | Tair | Tleaf | Tleaf2 | Fan_speed | Qamb_in | Qamb_out | Q | f_red | f_blue | f_farred | F | Q_modavg | F_dc | Tled | TDigital | TPreamp | TPwrSpy | TDrive | F_avg | dF/dt | dF_dc/dt | F_dc_avg | period | DIAG | Flow_s | Flow_r | Txchg | Tirga | Tchopper | Ts | Tr | CO2_% | Desiccant_% | Humidifier_% | Heatx_setpoint | CO2_r_setpoint | H2O_r_setpoint | SS_s | SS_r | MatchH2O | MatchCO2 | MatchValveR | MatchValveS
µmol m⁻² s⁻¹ | µmol mol⁻¹ |  |  | s | s | s | s | mol m⁻² s⁻¹ | µmol m⁻² s⁻¹ | µmol mol⁻¹ | Pa | Pa | mol m⁻² s⁻¹ | mol m⁻² s⁻¹ | mol m⁻² s⁻¹ | mol m⁻² s⁻¹ | W m⁻² | °C | °C | kPa | % | kPa | kPa | kPa | [8 empty] | µmol m⁻² s⁻¹ | [5 empty] | µmol m⁻² s⁻¹ | µmol m⁻² s⁻¹ | µmol m⁻² s⁻¹ | [10 empty] | µmol m⁻² s⁻¹ | µmol m⁻² s⁻¹ |  | J/µmol | s | µmol mol⁻¹ | µmol mol⁻¹ | mmol mol⁻¹ | mmol mol⁻¹ | µmol s⁻¹ | kPa | kPa | °C | °C | °C | rpm | µmol m⁻² s⁻¹ | µmol m⁻² s⁻¹ | µmol m⁻² s⁻¹ |  |  |  |  | µmol m⁻² s⁻¹ |  | °C | °C | °C | °C | °C |  | min⁻¹ | min⁻¹ |  | s |  | µmol s⁻¹ | µmol s⁻¹ | °C | °C | °C | °C | °C | % | % | % | °C | µmol mol⁻¹ | mmol mol⁻¹ | % | % | mmol/mol | µmol/mol | % | %
8.3658 | 498.5808 | Empty Chamber Correction for 500 to 0 | 1 | 1475420561.1 | 0 | 20161002 09:02:41 | 1475420560.6 | -0.00021362 | -0.71804 | 500.8622 | 42.2202 | 42.2028 | -4.5255 | 2.0105 | -0.014552 | 3.9819 | 165.2398 | 26.2882 | 25.0111 | 3.1818 | 62.9775 | 1.9823 | 3.1476 | 1.1995 | -0.03502 | 0.039313 | 3.0354 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.00033446 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.06 | 843.0507 | 0.843 | 0.196 | 1475420560.6 | 500.655 | 499.83 | 23.5159 | 23.7245 | 599.99 | 84.0953 | 0.19982 | 24.83 | 25.0111 | 999.9 | 10042.5 | 0 | 5.1025 | 1000.06 | 0.9 | 0.1 | 0 | 2.1644 | 0 | 138.171 | 40.187 | 43.25 | 42.062 | 41.812 | 40.25 | 2.2892 | 0.43685 | -1.4405 | 137.9105 | 15 | 2 | 571.324 | 711.31 | 21.9913 | 27.3152 | 30.0003 | 27.2955 | 27.2615 | 31.1343 | 19.4269 | 0.75899 | 21.9707 | 500 | 23.657 | 100.63 | 101.376 | -0.016 | 5.393 | 100 | 100
8.4424 | 498.5528 | Empty Chamber Correction for 500 to 0 | 2 | 1475420563.1 | 2 | 20161002 09:02:43 | 1475420562.6 | -0.00020962 | -0.64102 | 500.8344 | 42.2177 | 42.2019 | -4.5246 | 2.0102 | -0.014267 | 3.9814 | 165.2398 | 26.2864 | 25.0167 | 3.1828 | 62.9799 | 1.9823 | 3.1475 | 1.2006 | -0.035011 | 0.039303 | 3.0347 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.00042581 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.06 | 843.0507 | 0.843 | 0.196 | 1475420562.6 | 500.647 | 499.901 | 23.5161 | 23.7208 | 599.97 | 84.0946 | 0.20005 | 24.8294 | 25.0167 | 999.9 | 10040 | 0 | 5.1025 | 1000.07 | 0.9 | 0.1 | 0 | 2.3139 | 0 | 137.671 | 40.187 | 43.25 | 42.062 | 41.75 | 40.25 | 2.2842 | -0.021535 | -1.1838 | 137.8415 | 15 | 2 | 570.463 | 711.409 | 21.9833 | 27.3152 | 30.0003 | 27.2955 | 27.2615 | 31.1313 | 19.4269 | 0.75899 | 21.9587 | 500 | 23.6507 | 100.63 | 101.375 | -0.016 | 5.393 | 100 | 100
8.4544 | 498.5803 | Empty Chamber Correction for 500 to 0 | 3 | 1475420565.1 | 4 | 20161002 09:02:45 | 1475420564.6 | -0.00021014 | -0.62877 | 500.8678 | 42.2207 | 42.2051 | -4.5127 | 2.0049 | -0.0143 | 3.9709 | 165.2313 | 26.2916 | 25.0182 | 3.1831 | 62.9772 | 1.9824 | 3.1478 | 1.2007 | -0.034881 | 0.039156 | 3.0256 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.00044049 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.73 | 842.7724 | 0.843 | 0.196 | 1475420564.6 | 500.683 | 499.949 | 23.5172 | 23.7224 | 599.982 | 84.095 | 0.20009 | 24.831 | 25.0182 | 999.9 | 10002.5 | 0 | 5.1025 | 1000.08 | 0.89997 | 0.10003 | 0 | 2.446 | 0 | 137.528 | 40.187 | 43.25 | 42 | 41.812 | 40.25 | 2.2934 | 0.16812 | -1.0799 | 137.7848 | 15 | 2 | 567.611 | 711.186 | 21.9753 | 27.3159 | 30.0003 | 27.2955 | 27.2615 | 31.0734 | 19.4269 | 0.75899 | 21.9587 | 495 | 23.6459 | 100.63 | 101.374 | -0.016 | 5.393 | 100 | 100
8.3561 | 498.6306 | Empty Chamber Correction for 500 to 0 | 4 | 1475420567.1 | 6 | 20161002 09:02:47 | 1475420566.6 | -0.00021066 | -0.72754 | 500.9177 | 42.2249 | 42.2072 | -4.5137 | 2.0053 | -0.014334 | 3.9717 | 165.2398 | 26.2947 | 25.0191 | 3.1833 | 62.9655 | 1.9824 | 3.1484 | 1.2009 | -0.034891 | 0.039169 | 3.0264 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.00032318 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.06 | 843.0507 | 0.843 | 0.196 | 1475420566.6 | 500.708 | 499.875 | 23.5172 | 23.7229 | 600.014 | 84.0951 | 0.19995 | 24.8341 | 25.0191 | 999.9 | 10005.6 | 0 | 5.1025 | 1000.06 | 0.9 | 0.1 | 0 | 2.2733 | 0 | 137.628 | 40.187 | 43.25 | 42 | 41.812 | 40.25 | 2.2815 | -0.40965 | -0.55942 | 137.7717 | 15 | 2 | 569.431 | 711.087 | 21.9681 | 27.3171 | 30.0002 | 27.2955 | 27.2615 | 31.0048 | 19.4269 | 0.75899 | 21.9587 | 495 | 23.6416 | 100.63 | 101.374 | -0.016 | 5.393 | 100 | 100
7.6559 | 498.7425 | Empty Chamber Correction for 500 to 0 | 5 | 1475420569.1 | 8 | 20161002 09:02:49 | 1475420568.6 | -0.00021064 | -1.4324 | 501.0313 | 42.2349 | 42.2022 | -4.5127 | 2.0049 | -0.014336 | 3.9709 | 165.2414 | 26.2978 | 25.0173 | 3.183 | 62.954 | 1.9824 | 3.1489 | 1.2006 | -0.034881 | 0.039157 | 3.0256 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0005131 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.73 | 842.7724 | 0.843 | 0.196 | 1475420568.6 | 500.644 | 499.106 | 23.5166 | 23.7223 | 599.954 | 84.096 | 0.19991 | 24.8369 | 25.0173 | 999.9 | 10002.5 | 0 | 5.1025 | 1000.07 | 0.9 | 0.1 | 0 | 2.164 | 0 | 137.589 | 40.187 | 43.25 | 42 | 41.75 | 40.25 | 2.2974 | -0.79526 | -0.4613 | 137.7229 | 15 | 2 | 571.364 | 711.176 | 21.9611 | 27.3176 | 30.0002 | 27.2955 | 27.2626 | 30.8736 | 19.7149 | 0.75899 | 21.9402 | 490 | 23.6319 | 100.63 | 101.373 | -0.016 | 5.393 | 100 | 100
6.4116 | 498.6784 | Empty Chamber Correction for 500 to 0 | 6 | 1475420571.1 | 10 | 20161002 09:02:51 | 1475420570.6 | -0.00020829 | -2.6867 | 500.9656 | 42.2303 | 42.1712 | -4.5206 | 2.0084 | -0.014176 | 3.9779 | 165.2398 | 26.2933 | 25.0182 | 3.1831 | 62.9636 | 1.9825 | 3.1486 | 1.2006 | -0.034967 | 0.039253 | 3.0316 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0020007 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.07 | 843.0591 | 0.843 | 0.196 | 1475420570.6 | 500.264 | 497.473 | 23.5178 | 23.7212 | 599.982 | 84.0976 | 0.20026 | 24.8356 | 25.0182 | 999.9 | 10026.9 | 0 | 5.1025 | 1000.07 | 0.9 | 0.1 | 0 | 2.1695 | 0 | 137.66 | 40.187 | 43.25 | 42 | 41.75 | 40.187 | 2.2894 | -0.40666 | -1.246 | 137.7113 | 15 | 2 | 563.45 | 711.018 | 21.9545 | 27.3176 | 30.0003 | 27.2955 | 27.2638 | 30.8159 | 19.7149 | 0.75899 | 21.9402 | 485 | 23.629 | 100.63 | 101.373 | -0.016 | 5.393 | 100 | 100
5.2688 | 498.1047 | Empty Chamber Correction for 500 to 0 | 7 | 1475420573.1 | 12 | 20161002 09:02:53 | 1475420572.6 | -0.00019399 | -3.8418 | 500.397 | 42.1825 | 42.099 | -4.5157 | 2.0062 | -0.013191 | 3.9745 | 165.1829 | 26.2844 | 25.0242 | 3.1843 | 62.9802 | 1.9825 | 3.1478 | 1.2018 | -0.034915 | 0.039195 | 3.028 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.003372 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.722 | 842.7656 | 0.843 | 0.196 | 1475420572.6 | 499.406 | 495.468 | 23.5176 | 23.707 | 600.102 | 84.098 | 0.20015 | 24.8311 | 25.0242 | 999.9 | 10011.9 | 0 | 5.1025 | 999.722 | 0.9 | 0.1 | 0 | 2.2524 | 0 | 137.325 | 40.187 | 43.25 | 42 | 41.75 | 40.25 | 2.2791 | 0.29451 | -1.9283 | 137.6632 | 15 | 2 | 559.328 | 710.722 | 21.9465 | 27.3176 | 30.0001 | 27.2955 | 27.2638 | 30.6478 | 19.7149 | 0.75899 | 21.9201 | 485 | 23.6255 | 100.631 | 101.374 | -0.016 | 5.393 | 100 | 100
4.1851 | 497.1109 | Empty Chamber Correction for 500 to 0 | 8 | 1475420575.1 | 14 | 20161002 09:02:55 | 1475420574.6 | -0.0001778 | -4.94 | 499.4144 | 42.0994 | 41.9924 | -4.4998 | 1.9992 | -0.012094 | 3.9615 | 165.2398 | 26.28 | 25.0196 | 3.1834 | 62.9806 | 1.9821 | 3.1471 | 1.2013 | -0.034742 | 0.039001 | 3.0159 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0046734 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.07 | 843.0591 | 0.843 | 0.196 | 1475420574.6 | 498.145 | 493.117 | 23.513 | 23.6866 | 600.066 | 84.0979 | 0.19968 | 24.8276 | 25.0196 | 999.9 | 9962.5 | 0 | 5.1025 | 1000.07 | 0.9 | 0.1 | 0 | 2.4005 | 0 | 137.18 | 40.187 | 43.25 | 42 | 41.75 | 40.187 | 2.282 | 0.86155 | -2.1635 | 137.5799 | 15 | 2 | 568.445 | 711.019 | 21.9391 | 27.3176 | 30 | 27.2955 | 27.2638 | 30.4678 | 19.7149 | 0.75899 | 21.9201 | 480 | 23.6316 | 100.631 | 101.374 | -0.016 | 5.393 | 100 | 100
3.2086 | 495.6866 | Empty Chamber Correction for 500 to 0 | 9 | 1475420577.1 | 16 | 20161002 09:02:57 | 1475420576.6 | -0.00017347 | -5.9328 | 497.986 | 41.979 | 41.8514 | -4.5153 | 2.0061 | -0.011794 | 3.9755 | 165.2398 | 26.2718 | 25.0197 | 3.1834 | 62.9678 | 1.9814 | 3.1467 | 1.202 | -0.034912 | 0.039192 | 3.0278 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.005853 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.738 | 842.7791 | 0.843 | 0.196 | 1475420576.6 | 496.472 | 490.453 | 23.5053 | 23.6747 | 599.99 | 84.0977 | 0.19985 | 24.8255 | 25.0197 | 999.9 | 10011.2 | 0 | 5.1025 | 1000.08 | 0.89997 | 0.10003 | 0 | 2.2176 | 0 | 137.623 | 40.125 | 43.25 | 42 | 41.75 | 40.187 | 2.2825 | 0.073685 | -1.4199 | 137.5367 | 15 | 2 | 572.378 | 711.241 | 21.9305 | 27.3176 | 30.0002 | 27.2955 | 27.2638 | 30.3051 | 19.7149 | 0.75899 | 21.9201 | 475 | 23.6312 | 100.631 | 101.373 | -0.016 | 5.393 | 100 | 100
2.4331 | 493.7567 | Empty Chamber Correction for 500 to 0 | 10 | 1475420579.1 | 18 | 20161002 09:02:59 | 1475420578.6 | -0.00017757 | -6.7262 | 496.058 | 41.8167 | 41.6724 | -4.5209 | 2.0086 | -0.012055 | 3.9802 | 165.2414 | 26.2721 | 25.025 | 3.1844 | 62.9445 | 1.9808 | 3.1468 | 1.2037 | -0.034973 | 0.039261 | 3.0321 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0067922 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.07 | 843.0591 | 0.843 | 0.196 | 1475420578.6 | 494.346 | 487.532 | 23.497 | 23.6704 | 599.997 | 84.0978 | 0.2003 | 24.8259 | 25.025 | 999.9 | 10028.8 | 0 | 5.1025 | 1000.07 | 0.9 | 0.1 | 0 | 2.2776 | 0 | 137.575 | 40.125 | 43.25 | 42 | 41.75 | 40.187 | 2.2677 | -0.43663 | -1.0243 | 137.5318 | 15 | 2 | 565.457 | 711.068 | 21.9211 | 27.3176 | 30.0002 | 27.2955 | 27.2638 | 30.1683 | 19.7149 | 0.75899 | 21.899 | 475 | 23.6385 | 100.631 | 101.374 | -0.016 | 5.393 | 100 | 100
1.7392 | 491.5096 | Empty Chamber Correction for 500 to 0 | 11 | 1475420581.1 | 20 | 20161002 09:03:01 | 1475420580.6 | -0.00018423 | -7.4393 | 493.8168 | 41.6282 | 41.4686 | -4.5192 | 2.0078 | -0.012518 | 3.9783 | 165.233 | 26.275 | 25.0171 | 3.1829 | 62.9284 | 1.9802 | 3.1468 | 1.2027 | -0.034954 | 0.039239 | 3.0307 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0076407 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.723 | 842.7665 | 0.843 | 0.196 | 1475420580.6 | 491.924 | 484.394 | 23.4908 | 23.6707 | 599.996 | 84.0988 | 0.20002 | 24.8259 | 25.0171 | 999.9 | 10023.1 | 0 | 5.1025 | 1000.07 | 0.9 | 0.1 | 0 | 2.2701 | 0 | 137.33 | 40.125 | 43.25 | 42 | 41.75 | 40.187 | 2.2732 | -0.46451 | -0.78096 | 137.4736 | 15 | 2 | 563.403 | 710.846 | 21.9131 | 27.3176 | 30.0001 | 27.2956 | 27.2638 | 30.0199 | 19.7149 | 0.75899 | 21.899 | 470 | 23.6369 | 100.631 | 101.374 | -0.016 | 5.393 | 100 | 100
1.3548 | 488.9841 | Empty Chamber Correction for 500 to 0 | 12 | 1475420583.1 | 22 | 20161002 09:03:03 | 1475420582.6 | -0.00018896 | -7.8423 | 491.293 | 41.4163 | 41.2484 | -4.5253 | 2.0105 | -0.012839 | 3.9834 | 165.2398 | 26.275 | 25.0152 | 3.1826 | 62.9173 | 1.9799 | 3.1468 | 1.2027 | -0.03502 | 0.039313 | 3.0354 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0081161 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.07 | 843.0591 | 0.843 | 0.196 | 1475420582.6 | 489.301 | 481.367 | 23.486 | 23.6705 | 600.059 | 84.1007 | 0.1999 | 24.8258 | 25.0152 | 999.9 | 10041.9 | 0 | 5.1025 | 1000.07 | 0.9 | 0.1 | 0 | 2.4571 | 0 | 136.944 | 40.125 | 43.187 | 42 | 41.75 | 40.187 | 2.2883 | -0.24602 | -0.97874 | 137.4152 | 15 | 2 | 566.818 | 710.722 | 21.9042 | 27.3176 | 30.0001 | 27.2968 | 27.2638 | 29.8353 | 19.7149 | 0.75899 | 21.8812 | 465 | 23.64 | 100.631 | 101.374 | -0.016 | 5.393 | 100 | 100
1.0922 | 486.2832 | Empty Chamber Correction for 500 to 0 | 13 | 1475420585.1 | 24 | 20161002 09:03:05 | 1475420584.6 | -0.00019058 | -8.1236 | 488.6053 | 41.1902 | 41.0157 | -4.5088 | 2.0032 | -0.012946 | 3.9687 | 165.233 | 26.2778 | 25.0159 | 3.1827 | 62.9197 | 1.9797 | 3.1463 | 1.203 | -0.03484 | 0.039111 | 3.0228 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0084499 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.05 | 843.0413 | 0.843 | 0.196 | 1475420584.6 | 486.535 | 478.319 | 23.4831 | 23.6692 | 600.025 | 84.1015 | 0.20016 | 24.8233 | 25.0159 | 999.9 | 9990 | 0 | 5.1025 | 999.731 | 0.9 | 0.1 | 0 | 2.0828 | 0 | 137.582 | 40.125 | 43.25 | 42 | 41.75 | 40.187 | 2.2872 | -0.22349 | -1.0686 | 137.3728 | 15 | 2 | 563.619 | 710.846 | 21.8963 | 27.3176 | 30.0001 | 27.2978 | 27.2638 | 29.6823 | 19.7149 | 0.75899 | 21.8812 | 465 | 23.64 | 100.631 | 101.373 | -0.016 | 5.393 | 100 | 100
0.68993 | 483.494 | Empty Chamber Correction for 500 to 0 | 14 | 1475420587.1 | 26 | 20161002 09:03:07 | 1475420586.6 | -0.00018986 | -8.5462 | 485.8261 | 40.9555 | 40.7716 | -4.4994 | 1.999 | -0.0129 | 3.9604 | 165.2414 | 26.2761 | 25.013 | 3.1821 | 62.9271 | 1.9794 | 3.1455 | 1.2028 | -0.034737 | 0.038995 | 3.0156 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0089509 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.08 | 843.0676 | 0.843 | 0.196 | 1475420586.6 | 483.645 | 475.007 | 23.4801 | 23.6655 | 600.002 | 84.1007 | 0.20003 | 24.819 | 25.013 | 999.9 | 9960.62 | 0 | 5.1025 | 1000.09 | 0.9 | 0.1 | 0 | 2.4339 | 0 | 137.01 | 40.125 | 43.25 | 41.937 | 41.75 | 40.187 | 2.2787 | 0.10817 | -1.1188 | 137.327 | 15 | 2 | 562.759 | 710.92 | 21.8883 | 27.3176 | 30.0001 | 27.2978 | 27.2638 | 29.4912 | 19.7149 | 0.75899 | 21.8812 | 460 | 23.64 | 100.632 | 101.372 | -0.016 | 5.393 | 100 | 100
0.41625 | 480.5343 | Empty Chamber Correction for 500 to 0 | 15 | 1475420589.1 | 28 | 20161002 09:03:09 | 1475420588.6 | -0.00019059 | -8.8403 | 482.8634 | 40.7045 | 40.515 | -4.5152 | 2.0061 | -0.012954 | 3.9744 | 165.2414 | 26.2713 | 25.0093 | 3.1814 | 62.9213 | 1.9792 | 3.1455 | 1.2023 | -0.03491 | 0.03919 | 3.0277 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0093029 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.733 | 842.7749 | 0.843 | 0.196 | 1475420588.6 | 480.616 | 471.685 | 23.4781 | 23.6642 | 600.058 | 84.0982 | 0.19989 | 24.8186 | 25.0093 | 999.9 | 10010.6 | 0 | 5.1025 | 1000.08 | 0.89997 | 0.10003 | 0 | 2.0796 | 0 | 137.172 | 40.125 | 43.25 | 41.937 | 41.75 | 40.187 | 2.2591 | 0.43489 | -1.5591 | 137.3202 | 15 | 2 | 566.206 | 710.927 | 21.8797 | 27.3176 | 30.0001 | 27.2978 | 27.2644 | 29.3143 | 19.7149 | 0.75899 | 21.8661 | 455 | 23.64 | 100.631 | 101.372 | -0.016 | 5.393 | 100 | 100
0.41234 | 477.4275 | Empty Chamber Correction for 500 to 0 | 16 | 1475420591.1 | 30 | 20161002 09:03:11 | 1475420590.6 | -0.00019346 | -8.8634 | 479.762 | 40.4419 | 40.2519 | -4.5144 | 2.0057 | -0.013145 | 3.9734 | 165.2414 | 26.2726 | 25.0103 | 3.1816 | 62.9153 | 1.979 | 3.1455 | 1.2027 | -0.0349 | 0.039179 | 3.027 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.009327 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.09 | 843.076 | 0.843 | 0.196 | 1475420590.6 | 477.508 | 468.553 | 23.4765 | 23.6654 | 600.051 | 84.0957 | 0.20004 | 24.8186 | 25.0103 | 999.9 | 10008.1 | 0 | 5.1025 | 1000.09 | 0.9 | 0.1 | 0 | 2.1077 | 0 | 137.164 | 40.125 | 43.187 | 41.937 | 41.75 | 40.187 | 2.2406 | 0.36454 | -1.8188 | 137.2934 | 15 | 2 | 567.692 | 710.868 | 21.8734 | 27.3176 | 30.0001 | 27.2978 | 27.2655 | 29.1932 | 19.7149 | 0.75899 | 21.8661 | 455 | 23.64 | 100.632 | 101.373 | -0.016 | 5.393 | 100 | 100
0.23379 | 474.3356 | Empty Chamber Correction for 500 to 0 | 17 | 1475420593.1 | 32 | 20161002 09:03:13 | 1475420592.6 | -0.00019661 | -9.0624 | 476.6764 | 40.1814 | 39.9871 | -4.5124 | 2.0048 | -0.013362 | 3.9715 | 165.233 | 26.2745 | 25.0086 | 3.1813 | 62.9152 | 1.979 | 3.1455 | 1.2023 | -0.034878 | 0.039154 | 3.0255 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0095664 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.746 | 842.7859 | 0.843 | 0.196 | 1475420592.6 | 474.371 | 465.215 | 23.477 | 23.669 | 599.979 | 84.0949 | 0.2 | 24.8188 | 25.0086 | 999.9 | 10001.9 | 0 | 5.1025 | 1000.08 | 0.9 | 0.1 | 0 | 1.9147 | 0 | 137.571 | 40.125 | 43.25 | 41.937 | 41.75 | 40.187 | 2.2644 | -0.11403 | -1.3199 | 137.244 | 15 | 2 | 568.903 | 710.802 | 21.8667 | 27.3183 | 30.0001 | 27.2978 | 27.2661 | 28.9952 | 19.7149 | 0.75899 | 21.8577 | 450 | 23.64 | 100.633 | 101.374 | -0.016 | 5.393 | 100 | 100
0.093901 | 471.1988 | Empty Chamber Correction for 500 to 0 | 18 | 1475420595.1 | 34 | 20161002 09:03:15 | 1475420594.6 | -0.00019733 | -9.2227 | 473.5423 | 39.9173 | 39.7198 | -4.5172 | 2.0069 | -0.01343 | 3.9757 | 165.2431 | 26.2742 | 25.0001 | 3.1797 | 62.9139 | 1.979 | 3.1456 | 1.2007 | -0.034931 | 0.039213 | 3.0291 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0097533 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.09 | 843.076 | 0.843 | 0.196 | 1475420594.6 | 471.199 | 461.883 | 23.4774 | 23.6701 | 599.982 | 84.0952 | 0.19999 | 24.8195 | 25.0001 | 999.9 | 10016.9 | 0 | 5.1025 | 1000.06 | 0.90003 | 0.09997 | 0 | 2.1796 | 0 | 137.141 | 40.125 | 43.187 | 41.937 | 41.75 | 40.187 | 2.2774 | -0.48462 | -0.41709 | 137.1866 | 15 | 2 | 568.737 | 710.728 | 21.8611 | 27.3194 | 30.0001 | 27.2978 | 27.2661 | 28.8125 | 19.7149 | 0.75899 | 21.8577 | 445 | 23.64 | 100.633 | 101.374 | -0.016 | 5.393 | 100 | 100
0.052211 | 467.9571 | Empty Chamber Correction for 500 to 0 | 19 | 1475420597.1 | 36 | 20161002 09:03:17 | 1475420596.6 | -0.00019705 | -9.2847 | 470.3081 | 39.6455 | 39.4465 | -4.5125 | 2.0048 | -0.013418 | 3.9715 | 165.2464 | 26.2738 | 24.996 | 3.1789 | 62.9167 | 1.9789 | 3.1453 | 1.2 | -0.034879 | 0.039155 | 3.0255 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0098267 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.1 | 843.0844 | 0.843 | 0.196 | 1475420596.6 | 467.947 | 458.571 | 23.4756 | 23.668 | 600.061 | 84.0969 | 0.20002 | 24.8178 | 24.996 | 999.9 | 10001.9 | 0 | 5.1025 | 1000.1 | 0.9 | 0.1 | 0 | 2.3727 | 0 | 136.711 | 40.062 | 43.187 | 41.937 | 41.75 | 40.187 | 2.2927 | 0.019538 | -0.76526 | 137.1425 | 15 | 2 | 568.371 | 710.53 | 21.8577 | 27.3199 | 30.0001 | 27.2978 | 27.2661 | 28.6618 | 19.7149 | 0.75899 | 21.8577 | 445 | 23.64 | 100.633 | 101.375 | -0.016 | 5.393 | 100 | 100
-0.056824 | 464.6631 | Empty Chamber Correction for 500 to 0 | 20 | 1475420599.1 | 38 | 20161002 09:03:19 | 1475420598.6 | -0.00019447 | -9.4149 | 467.0225 | 39.3694 | 39.1674 | -4.5065 | 2.0022 | -0.013235 | 3.9664 | 165.2346 | 26.2735 | 24.9998 | 3.1796 | 62.9185 | 1.9789 | 3.1451 | 1.2008 | -0.034814 | 0.039082 | 3.021 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0099845 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.767 | 842.8036 | 0.843 | 0.196 | 1475420598.6 | 464.626 | 455.121 | 23.4745 | 23.6644 | 600.019 | 84.0988 | 0.19989 | 24.8169 | 24.9998 | 999.9 | 9983.12 | 0 | 5.1025 | 999.763 | 0.9 | 0.1 | 0 | 2.2321 | 0 | 136.762 | 40.125 | 43.187 | 41.937 | 41.75 | 40.125 | 2.2781 | 0.010947 | -1.747 | 137.0713 | 15 | 2 | 571.652 | 710.481 | 21.8545 | 27.3199 | 30.0001 | 27.2978 | 27.2661 | 28.4702 | 19.7149 | 0.75899 | 21.9247 | 440 | 23.64 | 100.633 | 101.374 | -0.016 | 5.393 | 100 | 100
0.044515 | 461.3051 | Empty Chamber Correction for 500 to 0 | 21 | 1475420601.1 | 40 | 20161002 09:03:21 | 1475420600.6 | -0.00019209 | -9.3335 | 463.6675 | 39.0863 | 38.8862 | -4.5103 | 2.0038 | -0.013056 | 3.9699 | 165.2447 | 26.2705 | 25.0081 | 3.1812 | 62.9224 | 1.9789 | 3.1449 | 1.2024 | -0.034856 | 0.039129 | 3.0239 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0098846 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.1 | 843.0844 | 0.843 | 0.196 | 1475420600.6 | 461.294 | 451.871 | 23.4746 | 23.6622 | 599.946 | 84.0981 | 0.19994 | 24.8158 | 25.0081 | 999.9 | 9995 | 0 | 5.1025 | 1000.06 | 0.90003 | 0.099971 | 0 | 2.127 | 0 | 136.657 | 40.062 | 43.187 | 41.937 | 41.75 | 40.125 | 2.2664 | -0.40916 | -1.8189 | 137.0142 | 15 | 2 | 572.146 | 710.777 | 21.8616 | 27.3199 | 30 | 27.2978 | 27.2661 | 28.3745 | 19.7149 | 0.75899 | 21.9247 | 435 | 23.64 | 100.632 | 101.374 | -0.016 | 5.393 | 100 | 100
0.16332 | 457.9886 | Empty Chamber Correction for 500 to 0 | 22 | 1475420603.1 | 42 | 20161002 09:03:23 | 1475420602.6 | -0.00018914 | -9.2343 | 460.3619 | 38.8072 | 38.6088 | -4.4986 | 1.9987 | -0.012847 | 3.9598 | 165.189 | 26.2687 | 25.0126 | 3.1821 | 62.9421 | 1.979 | 3.1442 | 1.203 | -0.034729 | 0.038986 | 3.015 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0097668 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.1 | 843.0844 | 0.843 | 0.196 | 1475420602.6 | 458.008 | 448.687 | 23.477 | 23.6617 | 599.993 | 84.0971 | 0.20014 | 24.8121 | 25.0126 | 999.9 | 9958.75 | 0 | 5.1025 | 1000.1 | 0.9 | 0.1 | 0 | 2.1168 | 0 | 136.887 | 40.062 | 43.187 | 41.937 | 41.687 | 40.125 | 2.2579 | -0.12165 | -1.9917 | 136.9898 | 15 | 2 | 568.681 | 710.703 | 21.8855 | 27.3199 | 29.9999 | 27.2978 | 27.2661 | 28.1722 | 19.7149 | 0.75899 | 21.9088 | 435 | 23.64 | 100.632 | 101.374 | -0.016 | 5.393 | 100 | 100
0.2038 | 454.75 | Empty Chamber Correction for 500 to 0 | 23 | 1475420605.1 | 44 | 20161002 09:03:25 | 1475420604.6 | -0.00018586 | -9.2134 | 457.1188 | 38.5339 | 38.3367 | -4.5164 | 2.0066 | -0.012628 | 3.9757 | 165.2447 | 26.2603 | 25.0116 | 3.1819 | 62.9545 | 1.9791 | 3.1438 | 1.2027 | -0.034923 | 0.039204 | 3.0286 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0097454 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.764 | 842.8011 | 0.843 | 0.196 | 1475420604.6 | 454.78 | 445.482 | 23.4781 | 23.6596 | 599.998 | 84.0972 | 0.20006 | 24.8096 | 25.0116 | 999.9 | 10014.4 | 0 | 5.1025 | 1000.1 | 0.9 | 0.1 | 0 | 2.1507 | 0 | 136.972 | 40.062 | 43.187 | 41.937 | 41.75 | 40.125 | 2.2589 | -0.11712 | -2.0756 | 136.9184 | 15 | 2 | 567.323 | 710.481 | 21.9087 | 27.3199 | 29.9999 | 27.2978 | 27.2661 | 27.9723 | 19.7149 | 0.75899 | 21.9088 | 430 | 23.64 | 100.633 | 101.374 | -0.016 | 5.393 | 100 | 100
0.19477 | 451.5432 | Empty Chamber Correction for 500 to 0 | 24 | 1475420607.1 | 46 | 20161002 09:03:27 | 1475420606.6 | -0.00018597 | -9.2423 | 453.9137 | 38.2636 | 38.0661 | -4.5225 | 2.0093 | -0.012649 | 3.9811 | 165.2464 | 26.257 | 25.0055 | 3.1807 | 62.9654 | 1.9793 | 3.1435 | 1.2014 | -0.03499 | 0.03928 | 3.0333 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0097763 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.11 | 843.0929 | 0.843 | 0.196 | 1475420606.6 | 451.571 | 442.245 | 23.4801 | 23.6617 | 600.02 | 84.097 | 0.20008 | 24.8081 | 25.0055 | 999.9 | 10033.8 | 0 | 5.1025 | 999.76 | 0.9 | 0.1 | 0 | 2.3141 | 0 | 136.566 | 40.062 | 43.187 | 41.875 | 41.687 | 40.125 | 2.2585 | -0.1516 | -2.033 | 136.8802 | 15 | 2 | 566.535 | 710.481 | 21.915 | 27.3199 | 30 | 27.2978 | 27.2661 | 27.7833 | 19.7149 | 0.75899 | 21.9088 | 425 | 23.6387 | 100.633 | 101.374 | -0.016 | 5.393 | 100 | 100
0.050164 | 448.3017 | Empty Chamber Correction for 500 to 0 | 25 | 1475420609.1 | 48 | 20161002 09:03:29 | 1475420608.6 | -0.00018025 | -9.408 | 450.6879 | 37.9919 | 37.7901 | -4.5026 | 2.0005 | -0.012279 | 3.9638 | 165.2464 | 26.2614 | 24.9973 | 3.1792 | 62.9707 | 1.9796 | 3.1436 | 1.1996 | -0.034773 | 0.039035 | 3.0181 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0099728 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.11 | 843.0929 | 0.843 | 0.196 | 1475420608.6 | 448.294 | 438.806 | 23.4829 | 23.6589 | 600.053 | 84.0975 | 0.20017 | 24.8088 | 24.9973 | 999.9 | 9971.25 | 0 | 5.1025 | 1000.11 | 0.9 | 0.1 | 0 | 2.139 | 0 | 136.782 | 40.062 | 43.187 | 41.875 | 41.687 | 40.125 | 2.261 | -0.49818 | -1.4511 | 136.8082 | 15 | 2 | 562.453 | 710.34 | 21.9142 | 27.3199 | 30.0002 | 27.2978 | 27.2667 | 27.6309 | 19.7149 | 0.75899 | 21.9015 | 425 | 23.6399 | 100.634 | 101.374 | -0.016 | 5.393 | 100 | 100
-0.064909 | 445.0242 | Empty Chamber Correction for 500 to 0 | 26 | 1475420611.1 | 50 | 20161002 09:03:31 | 1475420610.6 | -0.00017799 | -9.5442 | 447.405 | 37.7151 | 37.5113 | -4.5225 | 2.0093 | -0.01214 | 3.9816 | 165.1907 | 26.2552 | 24.9899 | 3.1778 | 62.9697 | 1.9796 | 3.1438 | 1.1981 | -0.034991 | 0.03928 | 3.0333 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010134 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.12 | 843.1022 | 0.843 | 0.196 | 1475420610.6 | 444.988 | 435.365 | 23.4839 | 23.6577 | 600.026 | 84.0973 | 0.20003 | 24.8097 | 24.9899 | 999.9 | 10033.8 | 0 | 5.1025 | 1000.12 | 0.9 | 0.1 | 0 | 2.3279 | 0 | 136.714 | 40.062 | 43.187 | 41.875 | 41.75 | 40.125 | 2.2548 | 0.034554 | -0.89583 | 136.7287 | 15 | 2 | 561.505 | 710.281 | 21.9126 | 27.3199 | 30.0002 | 27.2979 | 27.2678 | 27.4782 | 19.7149 | 0.75899 | 21.9015 | 420 | 23.6399 | 100.634 | 101.374 | -0.016 | 5.393 | 100 | 100
0.01607 | 441.6748 | Empty Chamber Correction for 500 to 0 | 27 | 1475420613.1 | 52 | 20161002 09:03:33 | 1475420612.6 | -0.00017942 | -9.4833 | 444.0577 | 37.4325 | 37.2303 | -4.5284 | 2.012 | -0.012234 | 3.9867 | 165.2379 | 26.2542 | 24.9915 | 3.1781 | 62.9691 | 1.9796 | 3.1437 | 1.1985 | -0.035056 | 0.039353 | 3.0378 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010066 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.768 | 842.8044 | 0.843 | 0.196 | 1475420612.6 | 441.659 | 432.097 | 23.4836 | 23.6588 | 600.033 | 84.0966 | 0.19999 | 24.8095 | 24.9915 | 999.9 | 10052.5 | 0 | 5.1025 | 1000.11 | 0.89997 | 0.10003 | 0 | 2.3391 | 0 | 136.464 | 40.062 | 43.187 | 41.875 | 41.687 | 40.125 | 2.2676 | 0.84417 | -1.3428 | 136.6759 | 15 | 2 | 560.954 | 710.265 | 21.91 | 27.3199 | 30.0001 | 27.2991 | 27.2684 | 27.2849 | 19.7149 | 0.75899 | 21.9089 | 415 | 23.6397 | 100.634 | 101.374 | -0.016 | 5.393 | 100 | 100
0.12882 | 438.3645 | Empty Chamber Correction for 500 to 0 | 28 | 1475420615.1 | 54 | 20161002 09:03:35 | 1475420614.6 | -0.00018065 | -9.3901 | 440.7558 | 37.1542 | 36.9537 | -4.5219 | 2.009 | -0.012312 | 3.9809 | 165.2481 | 26.2557 | 24.9945 | 3.1786 | 62.9748 | 1.9796 | 3.1436 | 1.199 | -0.034984 | 0.039272 | 3.0328 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0099549 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.77 | 842.8061 | 0.843 | 0.196 | 1475420614.6 | 438.377 | 428.908 | 23.4843 | 23.6607 | 600.018 | 84.0965 | 0.2001 | 24.8085 | 24.9945 | 999.9 | 10031.9 | 0 | 5.1025 | 999.77 | 0.9 | 0.1 | 0 | 2.1562 | 0 | 136.643 | 40.062 | 43.187 | 41.875 | 41.687 | 40.125 | 2.2735 | 0.56147 | -1.8875 | 136.6395 | 15 | 2 | 557.766 | 710.092 | 21.9074 | 27.3199 | 30.0001 | 27.3001 | 27.2684 | 27.1272 | 19.7149 | 0.75899 | 21.9089 | 415 | 23.6359 | 100.634 | 101.374 | -0.016 | 5.393 | 100 | 100
-0.061488 | 435.1671 | Empty Chamber Correction for 500 to 0 | 29 | 1475420617.1 | 56 | 20161002 09:03:37 | 1475420616.6 | -0.0001791 | -9.6016 | 437.5686 | 36.8853 | 36.6799 | -4.5125 | 2.0049 | -0.012188 | 3.9727 | 165.2481 | 26.2554 | 25.0048 | 3.1806 | 62.9882 | 1.9798 | 3.1431 | 1.2008 | -0.034881 | 0.039157 | 3.0256 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010202 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.11 | 843.0929 | 0.843 | 0.196 | 1475420616.6 | 435.132 | 425.452 | 23.4861 | 23.661 | 599.971 | 84.0962 | 0.19984 | 24.8061 | 25.0048 | 999.9 | 10002.5 | 0 | 5.1025 | 999.771 | 0.9 | 0.1 | 0 | 2.4047 | 0 | 136.357 | 40.062 | 43.187 | 41.875 | 41.687 | 40.125 | 2.2718 | 0.28176 | -1.934 | 136.5765 | 15 | 2 | 560.582 | 710.142 | 21.9077 | 27.3199 | 30.0001 | 27.3001 | 27.2684 | 26.9271 | 19.7149 | 0.75899 | 21.9089 | 410 | 23.6394 | 100.634 | 101.374 | -0.016 | 5.393 | 100 | 100
-0.053953 | 431.8239 | Empty Chamber Correction for 500 to 0 | 30 | 1475420619.1 | 58 | 20161002 09:03:39 | 1475420618.6 | -0.00017708 | -9.6146 | 434.2372 | 36.6047 | 36.3985 | -4.4999 | 1.9993 | -0.012052 | 3.9617 | 165.2514 | 26.2589 | 25.0039 | 3.1804 | 62.9868 | 1.9798 | 3.1432 | 1.2006 | -0.034744 | 0.039003 | 3.0161 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010218 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.12 | 843.1013 | 0.843 | 0.196 | 1475420618.6 | 431.791 | 422.101 | 23.4863 | 23.6592 | 600.067 | 84.0966 | 0.19993 | 24.8067 | 25.0039 | 999.9 | 9963.12 | 0 | 5.1025 | 1000.12 | 0.9 | 0.1 | 0 | 2.3697 | 0 | 136.16 | 40.062 | 43.187 | 41.875 | 41.687 | 40.125 | 2.2985 | 0.29984 | -2.0907 | 136.5152 | 15 | 2 | 562.146 | 710.043 | 21.9096 | 27.3199 | 30 | 27.3001 | 27.2684 | 26.7468 | 19.7149 | 0.75899 | 21.9092 | 405 | 23.6392 | 100.634 | 101.374 | -0.016 | 5.393 | 100 | 100
0.020481 | 428.4839 | Empty Chamber Correction for 500 to 0 | 31 | 1475420621.1 | 60 | 20161002 09:03:41 | 1475420620.6 | -0.00017648 | -9.5602 | 430.9014 | 36.3237 | 36.1188 | -4.5015 | 2 | -0.012034 | 3.9632 | 165.2481 | 26.2579 | 24.9921 | 3.1782 | 62.9862 | 1.9798 | 3.1432 | 1.1984 | -0.034761 | 0.039023 | 3.0173 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010153 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.12 | 843.1013 | 0.843 | 0.196 | 1475420620.6 | 428.47 | 418.836 | 23.4855 | 23.6578 | 600.114 | 84.0971 | 0.19995 | 24.8064 | 24.9921 | 999.9 | 9968.12 | 0 | 5.1025 | 1000.12 | 0.9 | 0.1 | 0 | 2.4599 | 0 | 136.235 | 40 | 43.187 | 41.875 | 41.687 | 40.125 | 2.3127 | 0.56589 | -2.252 | 136.4502 | 15 | 2 | 565.516 | 709.944 | 21.9105 | 27.3206 | 30.0002 | 27.3001 | 27.2684 | 26.6236 | 19.7149 | 0.75899 | 21.9092 | 405 | 23.6367 | 100.633 | 101.373 | -0.016 | 5.393 | 100 | 100
0.081217 | 425.1871 | Empty Chamber Correction for 500 to 0 | 32 | 1475420623.1 | 62 | 20161002 09:03:43 | 1475420622.6 | -0.00017325 | -9.5194 | 427.6001 | 36.0452 | 35.8419 | -4.519 | 2.0078 | -0.011809 | 3.9788 | 165.1915 | 26.2503 | 24.9951 | 3.1787 | 62.9918 | 1.9798 | 3.143 | 1.1989 | -0.034953 | 0.039238 | 3.0307 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010105 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.12 | 843.1022 | 0.843 | 0.196 | 1475420622.6 | 425.189 | 415.595 | 23.4865 | 23.6557 | 599.939 | 84.0965 | 0.19991 | 24.8055 | 24.9951 | 999.9 | 10023.1 | 0 | 5.1025 | 1000.12 | 0.9 | 0.1 | 0 | 2.5 | 0 | 136.1 | 40 | 43.187 | 41.875 | 41.687 | 40.125 | 2.3289 | 0.52291 | -2.3515 | 136.3615 | 15 | 2 | 569.169 | 710.191 | 21.9104 | 27.3217 | 30.0002 | 27.3001 | 27.2684 | 26.4352 | 19.7149 | 0.75899 | 21.9124 | 400 | 23.6387 | 100.633 | 101.373 | -0.016 | 5.393 | 100 | 100
0.21366 | 421.8558 | Empty Chamber Correction for 500 to 0 | 33 | 1475420625.1 | 64 | 20161002 09:03:45 | 1475420624.6 | -0.00017479 | -9.4065 | 424.2777 | 35.7659 | 35.5647 | -4.5115 | 2.0045 | -0.011909 | 3.9721 | 165.2481 | 26.2544 | 24.9978 | 3.1793 | 62.9882 | 1.9798 | 3.1432 | 1.1994 | -0.034871 | 0.039146 | 3.0249 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0099744 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.771 | 842.807 | 0.843 | 0.196 | 1475420624.6 | 421.891 | 412.41 | 23.4859 | 23.6566 | 599.949 | 84.0981 | 0.2002 | 24.8064 | 24.9978 | 999.9 | 9999.38 | 0 | 5.1025 | 1000.12 | 0.89997 | 0.10003 | 0 | 2.1141 | 0 | 136.367 | 40 | 43.125 | 41.875 | 41.687 | 40.062 | 2.3174 | -0.28027 | -1.6014 | 136.3197 | 15 | 2 | 563.072 | 709.919 | 21.9104 | 27.3222 | 30 | 27.3001 | 27.2684 | 26.2424 | 19.7149 | 0.75899 | 21.9124 | 395 | 23.6371 | 100.633 | 101.373 | -0.016 | 5.393 | 100 | 100
0.13688 | 418.587 | Empty Chamber Correction for 500 to 0 | 34 | 1475420627.1 | 66 | 20161002 09:03:47 | 1475420626.6 | -0.00017493 | -9.504 | 421.0119 | 35.4918 | 35.2888 | -4.5158 | 2.0064 | -0.011916 | 3.9759 | 165.2396 | 26.2545 | 24.9993 | 3.1795 | 62.9833 | 1.9798 | 3.1434 | 1.1997 | -0.034918 | 0.039198 | 3.0282 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.01009 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.78 | 842.8145 | 0.843 | 0.196 | 1475420626.6 | 418.603 | 409.027 | 23.4852 | 23.656 | 600.076 | 84.1012 | 0.20009 | 24.8078 | 24.9993 | 999.9 | 10012.5 | 0 | 5.1025 | 999.78 | 0.9 | 0.1 | 0 | 2.1395 | 0 | 136.346 | 40 | 43.125 | 41.812 | 41.687 | 40.062 | 2.3254 | -0.20692 | -0.83337 | 136.2815 | 15 | 2 | 560.67 | 709.697 | 21.911 | 27.3222 | 30 | 27.3001 | 27.2684 | 26.085 | 19.7149 | 0.75899 | 21.9124 | 395 | 23.6366 | 100.633 | 101.373 | -0.016 | 5.393 | 100 | 100
-0.011944 | 415.3819 | Empty Chamber Correction for 500 to 0 | 35 | 1475420629.1 | 68 | 20161002 09:03:49 | 1475420628.6 | -0.00017173 | -9.6737 | 417.8107 | 35.222 | 35.0155 | -4.5182 | 2.0074 | -0.011689 | 3.9782 | 165.2464 | 26.2546 | 25.0046 | 3.1805 | 62.9799 | 1.98 | 3.1438 | 1.2006 | -0.034944 | 0.039228 | 3.0301 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010288 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.12 | 843.1013 | 0.843 | 0.196 | 1475420628.6 | 415.361 | 405.616 | 23.4866 | 23.6543 | 600.004 | 84.1014 | 0.19988 | 24.8097 | 25.0046 | 999.9 | 10020 | 0 | 5.1025 | 999.789 | 0.9 | 0.1 | 0 | 2.329 | 0 | 136.412 | 40 | 43.125 | 41.812 | 41.687 | 40.062 | 2.3142 | 0.29454 | -0.29983 | 136.2753 | 15 | 2 | 563.596 | 710.018 | 21.912 | 27.3222 | 30.0001 | 27.3001 | 27.2684 | 25.8839 | 19.7149 | 0.75899 | 21.9127 | 390 | 23.6351 | 100.633 | 101.373 | -0.016 | 5.393 | 100 | 100
0.037618 | 412.0802 | Empty Chamber Correction for 500 to 0 | 36 | 1475420631.1 | 70 | 20161002 09:03:51 | 1475420630.6 | -0.0001702 | -9.6442 | 414.5134 | 34.9434 | 34.7376 | -4.5191 | 2.0079 | -0.011598 | 3.9791 | 165.2497 | 26.2536 | 24.9978 | 3.1793 | 62.9818 | 1.98 | 3.1437 | 1.1993 | -0.034954 | 0.039239 | 3.0308 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010253 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.12 | 843.1013 | 0.843 | 0.196 | 1475420630.6 | 412.072 | 402.358 | 23.4874 | 23.6536 | 600.018 | 84.0999 | 0.19998 | 24.8095 | 24.9978 | 999.9 | 10023.1 | 0 | 5.1025 | 1000.08 | 0.90003 | 0.09997 | 0 | 2.0839 | 0 | 136.199 | 40 | 43.125 | 41.812 | 41.687 | 40.062 | 2.3256 | -0.2534 | -0.11832 | 136.2477 | 15 | 2 | 564.747 | 710.166 | 21.9128 | 27.3222 | 30.0002 | 27.3001 | 27.2684 | 25.7259 | 19.7149 | 0.75899 | 21.9127 | 385 | 23.6377 | 100.632 | 101.373 | -0.016 | 5.393 | 100 | 100
0.12899 | 408.7561 | Empty Chamber Correction for 500 to 0 | 37 | 1475420633.1 | 72 | 20161002 09:03:53 | 1475420632.6 | -0.00017072 | -9.5726 | 411.1966 | 34.663 | 34.4585 | -4.5149 | 2.006 | -0.011646 | 3.9753 | 165.2481 | 26.253 | 24.9899 | 3.1778 | 62.9867 | 1.9799 | 3.1434 | 1.1979 | -0.034908 | 0.039187 | 3.0275 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010168 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.12 | 843.1013 | 0.843 | 0.196 | 1475420632.6 | 408.771 | 399.129 | 23.487 | 23.6537 | 600.023 | 84.0977 | 0.20017 | 24.8075 | 24.9899 | 999.9 | 10010 | 0 | 5.1025 | 1000.12 | 0.9 | 0.1 | 0 | 2.425 | 0 | 136.297 | 40 | 43.125 | 41.812 | 41.687 | 40.062 | 2.3329 | -0.29024 | 0.50116 | 136.24 | 15 | 2 | 561.932 | 710.074 | 21.9132 | 27.3222 | 30 | 27.3001 | 27.269 | 25.5592 | 19.7149 | 0.75899 | 21.9151 | 385 | 23.6353 | 100.632 | 101.372 | -0.016 | 5.393 | 100 | 100
-0.042523 | 405.5464 | Empty Chamber Correction for 500 to 0 | 38 | 1475420635.1 | 74 | 20161002 09:03:55 | 1475420634.6 | -0.00017214 | -9.7652 | 408.0011 | 34.3931 | 34.1838 | -4.4987 | 1.9988 | -0.011732 | 3.9609 | 165.1923 | 26.2558 | 24.9962 | 3.179 | 62.9952 | 1.9799 | 3.143 | 1.199 | -0.034731 | 0.038988 | 3.0151 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010396 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.12 | 843.1022 | 0.843 | 0.196 | 1475420634.6 | 405.518 | 395.683 | 23.4877 | 23.6558 | 600.003 | 84.0966 | 0.20004 | 24.8055 | 24.9962 | 999.9 | 9959.38 | 0 | 5.1025 | 1000.12 | 0.9 | 0.1 | 0 | 2.3391 | 0 | 136.085 | 40 | 43.125 | 41.812 | 41.687 | 40.062 | 2.3295 | 0.029262 | -0.22865 | 136.2136 | 15 | 2 | 560.163 | 709.941 | 21.9133 | 27.3222 | 30 | 27.3001 | 27.2701 | 25.356 | 19.7149 | 0.75899 | 21.9151 | 380 | 23.6337 | 100.633 | 101.372 | -0.016 | 5.393 | 100 | 100
-0.12028 | 402.2427 | Empty Chamber Correction for 500 to 0 | 39 | 1475420637.1 | 76 | 20161002 09:03:57 | 1475420636.6 | -0.00017225 | -9.864 | 404.6938 | 34.1149 | 33.9043 | -4.5149 | 2.006 | -0.011732 | 3.9752 | 165.2396 | 26.251 | 25.0008 | 3.1798 | 63.0003 | 1.98 | 3.1429 | 1.1998 | -0.034908 | 0.039187 | 3.0275 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010517 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.775 | 842.8103 | 0.843 | 0.196 | 1475420636.6 | 402.195 | 392.262 | 23.4885 | 23.6567 | 600.016 | 84.0982 | 0.19989 | 24.805 | 25.0008 | 999.9 | 10010 | 0 | 5.1025 | 1000.12 | 0.89997 | 0.10003 | 0 | 2.475 | 0 | 136.042 | 40 | 43.125 | 41.812 | 41.687 | 40.062 | 2.3184 | 0.51157 | -1.0359 | 136.1977 | 15 | 2 | 563.729 | 709.826 | 21.9139 | 27.3222 | 30.0002 | 27.3001 | 27.2707 | 25.1667 | 19.7149 | 0.75899 | 21.9151 | 375 | 23.6349 | 100.633 | 101.372 | -0.016 | 5.393 | 100 | 100
-0.064709 | 398.8384 | Empty Chamber Correction for 500 to 0 | 40 | 1475420639.1 | 78 | 20161002 09:03:59 | 1475420638.6 | -0.00017133 | -9.829 | 401.2882 | 33.8293 | 33.62 | -4.5268 | 2.0113 | -0.01168 | 3.9858 | 165.2481 | 26.2486 | 24.995 | 3.1787 | 62.995 | 1.98 | 3.1431 | 1.1987 | -0.035038 | 0.039333 | 3.0366 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010476 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.784 | 842.8179 | 0.843 | 0.196 | 1475420638.6 | 398.805 | 388.908 | 23.4873 | 23.6546 | 600.019 | 84.1018 | 0.2 | 24.8063 | 24.995 | 999.9 | 10046.9 | 0 | 5.1025 | 999.784 | 0.9 | 0.1 | 0 | 2.4559 | 0 | 135.657 | 40 | 43.125 | 41.812 | 41.687 | 40.062 | 2.3274 | 0.76852 | -2.0707 | 136.159 | 15 | 2 | 565.241 | 709.654 | 21.9147 | 27.3222 | 30.0002 | 27.3001 | 27.2707 | 25.0178 | 19.7149 | 0.75899 | 21.9165 | 375 | 23.6344 | 100.633 | 101.372 | -0.016 | 5.393 | 100 | 100
-0.095267 | 395.4709 | Empty Chamber Correction for 500 to 0 | 41 | 1475420641.1 | 80 | 20161002 09:04:01 | 1475420640.6 | -0.0001703 | -9.8806 | 397.9291 | 33.547 | 33.3363 | -4.5209 | 2.0086 | -0.011628 | 3.9806 | 165.2497 | 26.2496 | 24.986 | 3.177 | 62.9988 | 1.9801 | 3.1431 | 1.1969 | -0.034974 | 0.039261 | 3.0321 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010534 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.08 | 843.0666 | 0.843 | 0.196 | 1475420640.6 | 395.43 | 385.482 | 23.4877 | 23.654 | 599.994 | 84.104 | 0.19995 | 24.806 | 24.986 | 999.9 | 10028.1 | 0 | 5.1025 | 999.784 | 0.9 | 0.1 | 0 | 2.3999 | 0 | 135.705 | 40 | 43.125 | 41.812 | 41.687 | 40.062 | 2.3435 | 0.20624 | -2.6126 | 136.083 | 15 | 2 | 566.272 | 709.851 | 21.9154 | 27.3222 | 30 | 27.3008 | 27.2707 | 24.8606 | 19.7149 | 0.75899 | 21.9165 | 370 | 23.6273 | 100.633 | 101.371 | -0.016 | 5.393 | 100 | 100
0.013613 | 392.1158 | Empty Chamber Correction for 500 to 0 | 42 | 1475420643.1 | 82 | 20161002 09:04:03 | 1475420642.6 | -0.00016496 | -9.7915 | 394.5756 | 33.2639 | 33.0554 | -4.527 | 2.0114 | -0.011272 | 3.9863 | 165.2497 | 26.2437 | 24.9819 | 3.1762 | 63.0111 | 1.9802 | 3.1427 | 1.196 | -0.035041 | 0.039336 | 3.0368 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010428 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.13 | 843.1097 | 0.843 | 0.196 | 1475420642.6 | 392.103 | 382.246 | 23.4895 | 23.6506 | 599.952 | 84.1028 | 0.20013 | 24.8038 | 24.9819 | 999.9 | 10047.5 | 0 | 5.1025 | 1000.09 | 0.90003 | 0.09997 | 0 | 2.2512 | 0 | 135.807 | 40 | 43.125 | 41.812 | 41.687 | 40.062 | 2.3318 | -0.16886 | -2.4762 | 136.0264 | 15 | 2 | 563.429 | 709.975 | 21.9161 | 27.3222 | 30 | 27.3019 | 27.2707 | 24.6656 | 19.7149 | 0.75899 | 21.9259 | 365 | 23.6325 | 100.634 | 101.371 | -0.016 | 5.393 | 100 | 100
0.14503 | 388.8303 | Empty Chamber Correction for 500 to 0 | 43 | 1475420645.1 | 84 | 20161002 09:04:05 | 1475420644.6 | -0.00016131 | -9.6793 | 391.2986 | 32.9865 | 32.7801 | -4.5198 | 2.0082 | -0.011014 | 3.9803 | 165.2531 | 26.2436 | 24.9866 | 3.1771 | 63.0154 | 1.9803 | 3.1425 | 1.1969 | -0.034963 | 0.039249 | 3.0314 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010294 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.13 | 843.1097 | 0.843 | 0.196 | 1475420644.6 | 388.851 | 379.11 | 23.4908 | 23.6483 | 600.066 | 84.0999 | 0.20011 | 24.803 | 24.9866 | 999.9 | 10025.6 | 0 | 5.1025 | 1000.13 | 0.9 | 0.1 | 0 | 2.1446 | 0 | 135.769 | 40 | 43.125 | 41.812 | 41.625 | 40.062 | 2.3188 | -0.27421 | -2.2654 | 135.9661 | 15 | 2 | 560.659 | 709.703 | 21.917 | 27.3223 | 30 | 27.3024 | 27.2707 | 24.5029 | 19.7149 | 0.75899 | 21.9259 | 365 | 23.6269 | 100.634 | 101.371 | -0.016 | 5.393 | 100 | 100
-0.0083678 | 385.6401 | Empty Chamber Correction for 500 to 0 | 44 | 1475420647.1 | 86 | 20161002 09:04:07 | 1475420646.6 | -0.00016376 | -9.8536 | 388.1252 | 32.7185 | 32.5074 | -4.4989 | 1.9989 | -0.01116 | 3.9617 | 165.193 | 26.2504 | 24.9978 | 3.1793 | 63.0104 | 1.9802 | 3.1426 | 1.1991 | -0.034734 | 0.038992 | 3.0154 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010501 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.13 | 843.1106 | 0.843 | 0.196 | 1475420646.6 | 385.622 | 375.706 | 23.4897 | 23.6496 | 600.047 | 84.0989 | 0.19982 | 24.8033 | 24.9978 | 999.9 | 9960 | 0 | 5.1025 | 1000.12 | 0.9 | 0.1 | 0 | 2.4702 | 0 | 135.848 | 40 | 43.125 | 41.812 | 41.625 | 40.062 | 2.3387 | -0.43508 | -1.3359 | 135.9153 | 15 | 2 | 564.868 | 709.777 | 21.9198 | 27.3235 | 30.0001 | 27.3024 | 27.2707 | 24.2993 | 19.7149 | 0.75899 | 21.9259 | 360 | 23.6278 | 100.634 | 101.371 | -0.016 | 5.393 | 100 | 100
-0.024008 | 382.3478 | Empty Chamber Correction for 500 to 0 | 45 | 1475420649.1 | 88 | 20161002 09:04:09 | 1475420648.6 | -0.00016324 | -9.8897 | 384.8361 | 32.4412 | 32.2296 | -4.5025 | 2.0005 | -0.011114 | 3.9649 | 165.2412 | 26.2479 | 25.0036 | 3.1804 | 63.0176 | 1.9802 | 3.1423 | 1.2002 | -0.034773 | 0.039036 | 3.0181 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010548 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.783 | 842.8171 | 0.843 | 0.196 | 1475420648.6 | 382.326 | 372.374 | 23.49 | 23.6494 | 600.007 | 84.0987 | 0.19998 | 24.8016 | 25.0036 | 999.9 | 9971.25 | 0 | 5.1025 | 1000.13 | 0.89997 | 0.10003 | 0 | 2.213 | 0 | 136 | 39.937 | 43.125 | 41.75 | 41.625 | 40 | 2.3251 | -0.79467 | -0.16014 | 135.8725 | 15 | 2 | 565.905 | 709.826 | 21.9236 | 27.3245 | 30.0002 | 27.3024 | 27.2707 | 24.1112 | 19.7149 | 0.75899 | 21.9308 | 355 | 23.6265 | 100.633 | 101.371 | -0.016 | 5.393 | 100 | 100
0.029008 | 379.0532 | Empty Chamber Correction for 500 to 0 | 46 | 1475420651.1 | 90 | 20161002 09:04:11 | 1475420650.6 | -0.00016271 | -9.8566 | 381.5379 | 32.1625 | 31.9524 | -4.5181 | 2.0074 | -0.011082 | 3.9787 | 165.2429 | 26.241 | 25.0013 | 3.1799 | 63.0254 | 1.9802 | 3.1418 | 1.1998 | -0.034944 | 0.039228 | 3.0301 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010508 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.802 | 842.8331 | 0.843 | 0.196 | 1475420650.6 | 379.045 | 369.126 | 23.4903 | 23.6492 | 599.957 | 84.0968 | 0.20024 | 24.7994 | 25.0013 | 999.9 | 10020.6 | 0 | 5.1025 | 999.802 | 0.9 | 0.1 | 0 | 2.3489 | 0 | 135.7 | 39.937 | 43.125 | 41.75 | 41.687 | 40 | 2.3086 | -0.36815 | -0.17747 | 135.8553 | 15 | 2 | 560.63 | 709.802 | 21.9265 | 27.3245 | 30.0001 | 27.3024 | 27.2707 | 23.9865 | 19.7149 | 0.75899 | 21.9308 | 355 | 23.6271 | 100.633 | 101.371 | -0.016 | 5.393 | 100 | 100
-0.0315 | 375.8143 | Empty Chamber Correction for 500 to 0 | 47 | 1475420653.1 | 92 | 20161002 09:04:13 | 1475420652.6 | -0.00016229 | -9.9376 | 378.3079 | 31.8901 | 31.6779 | -4.5112 | 2.0043 | -0.011064 | 3.9726 | 165.253 | 26.2446 | 24.9954 | 3.1788 | 63.019 | 1.9802 | 3.1422 | 1.1986 | -0.034868 | 0.039142 | 3.0247 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010604 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.797 | 842.8289 | 0.843 | 0.196 | 1475420652.6 | 375.791 | 365.791 | 23.4904 | 23.6489 | 599.914 | 84.0964 | 0.20025 | 24.8011 | 24.9954 | 999.9 | 9998.75 | 0 | 5.1025 | 1000.15 | 0.9 | 0.1 | 0 | 2.1066 | 0 | 135.937 | 39.937 | 43.125 | 41.75 | 41.625 | 40 | 2.2865 | -0.13546 | 0.36164 | 135.8382 | 15 | 2 | 552.013 | 709.579 | 21.9291 | 27.3245 | 30 | 27.3024 | 27.2707 | 23.7943 | 19.7149 | 0.75899 | 21.9194 | 350 | 23.6268 | 100.633 | 101.371 | -0.016 | 5.393 | 100 | 100
0.099014 | 372.4912 | Empty Chamber Correction for 500 to 0 | 48 | 1475420655.1 | 94 | 20161002 09:04:15 | 1475420654.6 | -0.00016162 | -9.8265 | 374.9965 | 31.6109 | 31.4006 | -4.499 | 1.9989 | -0.011015 | 3.9619 | 165.243 | 26.2491 | 24.9976 | 3.1792 | 63.0159 | 1.9802 | 3.1424 | 1.199 | -0.034735 | 0.038993 | 3.0154 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010469 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.1 | 843.0835 | 0.843 | 0.196 | 1475420654.6 | 372.501 | 362.616 | 23.4911 | 23.6489 | 600.106 | 84.0967 | 0.1999 | 24.8024 | 24.9976 | 999.9 | 9960.62 | 0 | 5.1025 | 999.807 | 0.9 | 0.1 | 0 | 2.4407 | 0 | 135.663 | 39.937 | 43.062 | 41.75 | 41.5 | 40 | 2.2761 | -0.45604 | 0.44807 | 135.8357 | 15 | 2 | 551.709 | 709.135 | 21.9309 | 27.3245 | 30.0001 | 27.3024 | 27.2707 | 23.5945 | 19.7149 | 0.75899 | 21.9194 | 345 | 23.6242 | 100.633 | 101.371 | -0.016 | 5.393 | 100 | 100
0.14633 | 369.2491 | Empty Chamber Correction for 500 to 0 | 49 | 1475420657.1 | 96 | 20161002 09:04:17 | 1475420656.6 | -0.00015987 | -9.7989 | 371.7485 | 31.337 | 31.1282 | -4.5183 | 2.0075 | -0.01089 | 3.979 | 165.195 | 26.2427 | 25.0012 | 3.1799 | 63.0178 | 1.9803 | 3.1425 | 1.1996 | -0.034946 | 0.03923 | 3.0302 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.01044 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.8 | 842.8314 | 0.843 | 0.196 | 1475420656.6 | 369.271 | 359.414 | 23.4923 | 23.6484 | 600.055 | 84.0964 | 0.19985 | 24.8027 | 25.0012 | 999.9 | 10021.2 | 0 | 5.1025 | 999.8 | 0.9 | 0.1 | 0 | 2.444 | 0 | 135.525 | 39.937 | 43.062 | 41.75 | 41.562 | 40 | 2.281 | 0.26437 | -1.0679 | 135.798 | 15 | 2 | 558.094 | 709.372 | 21.9285 | 27.3245 | 30.0002 | 27.3024 | 27.2719 | 23.4336 | 19.7149 | 0.75899 | 21.9194 | 345 | 23.6226 | 100.633 | 101.371 | -0.016 | 5.393 | 100 | 100
0.0034378 | 366.0468 | Empty Chamber Correction for 500 to 0 | 50 | 1475420659.1 | 98 | 20161002 09:04:19 | 1475420658.6 | -0.00015922 | -9.9626 | 368.5484 | 31.0676 | 30.8556 | -4.5237 | 2.0099 | -0.010857 | 3.9839 | 165.2514 | 26.2431 | 24.996 | 3.1789 | 63.0157 | 1.9804 | 3.1428 | 1.1985 | -0.035006 | 0.039297 | 3.0343 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010634 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.804 | 842.8348 | 0.843 | 0.196 | 1475420658.6 | 366.033 | 356.011 | 23.4936 | 23.6491 | 599.935 | 84.0972 | 0.20013 | 24.8044 | 24.996 | 999.9 | 10038.1 | 0 | 5.1025 | 999.796 | 0.9 | 0.1 | 0 | 1.9208 | 0 | 135.488 | 39.937 | 43.062 | 41.75 | 41.437 | 40 | 2.296 | 0.33936 | -2.1444 | 135.7445 | 15 | 2 | 556.105 | 709.511 | 21.9241 | 27.3245 | 30.0001 | 27.3024 | 27.273 | 23.225 | 19.7149 | 0.75899 | 21.9226 | 340 | 23.6218 | 100.633 | 101.372 | -0.016 | 5.393 | 100 | 100
-0.057119 | 362.7841 | Empty Chamber Correction for 500 to 0 | 51 | 1475420661.1 | 100 | 20161002 09:04:21 | 1475420660.6 | -0.00015968 | -10.0438 | 365.3061 | 30.7941 | 30.5791 | -4.4966 | 1.9979 | -0.010891 | 3.9599 | 165.2514 | 26.2491 | 24.9936 | 3.1785 | 63.0212 | 1.9804 | 3.1424 | 1.1981 | -0.034709 | 0.038964 | 3.0136 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.01073 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.797 | 842.8289 | 0.843 | 0.196 | 1475420660.6 | 362.755 | 352.655 | 23.493 | 23.6489 | 600.101 | 84.0968 | 0.19996 | 24.8024 | 24.9936 | 999.9 | 9953.12 | 0 | 5.1025 | 1000.14 | 0.89997 | 0.10003 | 0 | 2.2559 | 0 | 135.333 | 39.875 | 43.062 | 41.75 | 41.5 | 40 | 2.284 | 0.16772 | -2.8357 | 135.7105 | 15 | 2 | 555.169 | 709.042 | 21.9219 | 27.3245 | 30.0001 | 27.3024 | 27.273 | 23.1254 | 19.7149 | 0.75899 | 21.9226 | 335 | 23.6207 | 100.633 | 101.373 | -0.016 | 5.393 | 100 | 100
0.02633 | 359.504 | Empty Chamber Correction for 500 to 0 | 52 | 1475420663.1 | 102 | 20161002 09:04:23 | 1475420662.6 | -0.0001613 | -9.9799 | 362.0285 | 30.5178 | 30.3043 | -4.501 | 1.9998 | -0.011006 | 3.9637 | 165.2429 | 26.245 | 24.9912 | 3.178 | 63.033 | 1.9804 | 3.1418 | 1.1977 | -0.034757 | 0.039018 | 3.017 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010654 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.799 | 842.8306 | 0.843 | 0.196 | 1475420662.6 | 359.496 | 349.459 | 23.4927 | 23.6502 | 600.055 | 84.0967 | 0.19985 | 24.799 | 24.9912 | 999.9 | 9966.88 | 0 | 5.1025 | 999.799 | 0.9 | 0.1 | 0 | 2.1346 | 0 | 135.605 | 39.937 | 43.062 | 41.75 | 41.5 | 40 | 2.2787 | 0.063552 | -2.4581 | 135.6436 | 15 | 2 | 559.922 | 709.092 | 21.922 | 27.3245 | 30.0001 | 27.3024 | 27.273 | 22.9131 | 19.7149 | 0.75899 | 21.9268 | 335 | 23.6198 | 100.633 | 101.372 | -0.016 | 5.393 | 100 | 100
-0.013779 | 356.2428 | Empty Chamber Correction for 500 to 0 | 53 | 1475420665.1 | 104 | 20161002 09:04:25 | 1475420664.6 | -0.00015973 | -10.0405 | 358.7609 | 30.2427 | 30.0289 | -4.5215 | 2.009 | -0.010886 | 3.9819 | 165.2514 | 26.2402 | 24.9984 | 3.1794 | 63.0276 | 1.9804 | 3.1421 | 1.199 | -0.034982 | 0.03927 | 3.0327 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010726 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.796 | 842.828 | 0.843 | 0.196 | 1475420664.6 | 356.225 | 346.126 | 23.4928 | 23.6488 | 599.904 | 84.0975 | 0.20006 | 24.8007 | 24.9984 | 999.9 | 10031.2 | 0 | 5.1025 | 1000.14 | 0.9 | 0.1 | 0 | 2.3467 | 0 | 135.141 | 39.875 | 43.062 | 41.687 | 41.375 | 40 | 2.3201 | 0.1289 | -2.5322 | 135.5192 | 15 | 2 | 560.421 | 709.314 | 21.9224 | 27.3245 | 30 | 27.3024 | 27.273 | 22.6974 | 19.7149 | 0.75899 | 21.9268 | 330 | 23.619 | 100.633 | 101.372 | -0.016 | 5.393 | 100 | 100
0.0064979 | 352.9415 | Empty Chamber Correction for 500 to 0 | 54 | 1475420667.1 | 106 | 20161002 09:04:27 | 1475420666.6 | -0.00015923 | -10.0404 | 355.4741 | 29.9656 | 29.751 | -4.5048 | 2.0015 | -0.010834 | 3.9672 | 165.253 | 26.2471 | 25.0096 | 3.1815 | 63.0204 | 1.9804 | 3.1425 | 1.2011 | -0.034799 | 0.039065 | 3.0199 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010722 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.15 | 843.1266 | 0.843 | 0.196 | 1475420666.6 | 352.929 | 342.832 | 23.4932 | 23.6487 | 599.977 | 84.0974 | 0.20011 | 24.8029 | 25.0096 | 999.9 | 9978.75 | 0 | 5.1025 | 999.795 | 0.9 | 0.1 | 0 | 2.0522 | 0 | 135.374 | 39.875 | 43.062 | 41.687 | 41.5 | 40 | 2.2743 | -0.18201 | -2.2739 | 135.4944 | 15 | 2 | 557.484 | 709.289 | 21.9236 | 27.3245 | 30.0001 | 27.3024 | 27.273 | 22.5031 | 19.7149 | 0.75899 | 21.9268 | 325 | 23.6162 | 100.632 | 101.371 | -0.016 | 5.393 | 100 | 100
-0.048044 | 349.6934 | Empty Chamber Correction for 500 to 0 | 55 | 1475420669.1 | 108 | 20161002 09:04:29 | 1475420668.6 | -0.0001617 | -10.1154 | 352.2376 | 29.6928 | 29.4761 | -4.4936 | 1.9966 | -0.011004 | 3.9572 | 165.1953 | 26.2502 | 25.008 | 3.1812 | 63.0222 | 1.9804 | 3.1424 | 1.2008 | -0.034677 | 0.038928 | 3.0113 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010815 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.802 | 842.8331 | 0.843 | 0.196 | 1475420668.6 | 349.667 | 339.495 | 23.493 | 23.6509 | 599.998 | 84.0974 | 0.20026 | 24.8023 | 25.008 | 999.9 | 9943.75 | 0 | 5.1025 | 999.802 | 0.9 | 0.1 | 0 | 2.507 | 0 | 135.243 | 39.875 | 43.125 | 41.75 | 41.5 | 40 | 2.2947 | -0.10716 | -1.4461 | 135.4296 | 15 | 2 | 550.75 | 709.141 | 21.9252 | 27.3246 | 30.0002 | 27.3031 | 27.273 | 22.3401 | 19.7149 | 0.75899 | 21.9237 | 325 | 23.6137 | 100.632 | 101.371 | -0.016 | 5.393 | 100 | 100
-0.20183 | 346.4448 | Empty Chamber Correction for 500 to 0 | 56 | 1475420671.1 | 110 | 20161002 09:04:31 | 1475420670.6 | -0.00016088 | -10.2904 | 348.9833 | 29.4186 | 29.1992 | -4.513 | 2.0052 | -0.010961 | 3.9743 | 165.2445 | 26.2425 | 25.0013 | 3.1799 | 63.033 | 1.9805 | 3.142 | 1.1994 | -0.034888 | 0.039165 | 3.0262 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011023 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.802 | 842.8331 | 0.843 | 0.196 | 1475420670.6 | 346.38 | 336.034 | 23.4938 | 23.6509 | 600.01 | 84.098 | 0.20008 | 24.8001 | 25.0013 | 999.9 | 10004.4 | 0 | 5.1025 | 999.798 | 0.9 | 0.1 | 0 | 2.3775 | 0 | 135.074 | 39.875 | 43.062 | 41.687 | 41.562 | 39.937 | 2.2909 | -0.91344 | -0.11197 | 135.3812 | 15 | 2 | 548.422 | 708.993 | 21.9257 | 27.3251 | 30.0001 | 27.3042 | 27.273 | 22.1756 | 19.7149 | 0.75899 | 21.9237 | 320 | 23.6131 | 100.632 | 101.371 | -0.016 | 5.393 | 100 | 100
-0.063267 | 343.0625 | Empty Chamber Correction for 500 to 0 | 57 | 1475420673.1 | 112 | 20161002 09:04:33 | 1475420672.6 | -0.00015751 | -10.1717 | 345.6057 | 29.1338 | 28.9169 | -4.5134 | 2.0053 | -0.010747 | 3.9748 | 165.2497 | 26.2392 | 24.9926 | 3.1783 | 63.045 | 1.9806 | 3.1416 | 1.1977 | -0.034892 | 0.03917 | 3.0264 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010882 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.799 | 842.8306 | 0.843 | 0.196 | 1475420672.6 | 343.033 | 332.808 | 23.4956 | 23.6494 | 600.041 | 84.0978 | 0.19989 | 24.7981 | 24.9926 | 999.9 | 10005.6 | 0 | 5.1025 | 1000.15 | 0.89997 | 0.10003 | 0 | 2.4396 | 0 | 135.036 | 39.875 | 43.062 | 41.75 | 41.562 | 40 | 2.2487 | -0.53368 | -0.32872 | 135.3793 | 15 | 2 | 552.048 | 708.993 | 21.9249 | 27.3256 | 30.0001 | 27.3047 | 27.273 | 21.971 | 19.7149 | 0.75899 | 21.9242 | 315 | 23.6087 | 100.633 | 101.372 | -0.016 | 5.393 | 100 | 100
0.13728 | 339.697 | Empty Chamber Correction for 500 to 0 | 58 | 1475420675.1 | 114 | 20161002 09:04:35 | 1475420674.6 | -0.00015863 | -9.9903 | 342.2435 | 28.8503 | 28.6374 | -4.5162 | 2.0066 | -0.010829 | 3.9772 | 165.2445 | 26.2383 | 24.9898 | 3.1777 | 63.047 | 1.9806 | 3.1415 | 1.1971 | -0.034923 | 0.039204 | 3.0286 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010667 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.808 | 842.8381 | 0.843 | 0.196 | 1475420674.6 | 339.718 | 329.674 | 23.4958 | 23.6507 | 600.014 | 84.0975 | 0.20007 | 24.7977 | 24.9898 | 999.9 | 10014.4 | 0 | 5.1025 | 999.808 | 0.9 | 0.1 | 0 | 2.3179 | 0 | 135.275 | 39.875 | 43.062 | 41.687 | 41.437 | 40 | 2.2565 | -0.026841 | -0.54561 | 135.3502 | 15 | 2 | 551.062 | 708.845 | 21.9241 | 27.3263 | 30.0001 | 27.3047 | 27.273 | 21.7983 | 19.7149 | 0.75899 | 21.9242 | 315 | 23.6104 | 100.633 | 101.372 | -0.016 | 5.393 | 100 | 100
-0.038374 | 336.4769 | Empty Chamber Correction for 500 to 0 | 59 | 1475420677.1 | 116 | 20161002 09:04:37 | 1475420676.6 | -0.00015718 | -10.1872 | 339.0202 | 28.5786 | 28.3623 | -4.531 | 2.0132 | -0.010736 | 3.9905 | 165.2514 | 26.233 | 24.986 | 3.177 | 63.0479 | 1.9806 | 3.1414 | 1.1964 | -0.035086 | 0.039387 | 3.04 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0109 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.808 | 842.8381 | 0.843 | 0.196 | 1475420676.6 | 336.454 | 326.213 | 23.4953 | 23.6488 | 599.948 | 84.0975 | 0.20007 | 24.7971 | 24.986 | 999.9 | 10061.2 | 0 | 5.1025 | 1000.14 | 0.9 | 0.1 | 0 | 2.5296 | 0 | 135.121 | 39.875 | 43.062 | 41.687 | 41.437 | 39.937 | 2.2903 | 0.5009 | -0.84332 | 135.3195 | 15 | 2 | 549.587 | 708.796 | 21.9238 | 27.3268 | 30.0001 | 27.3047 | 27.273 | 21.5895 | 19.7149 | 0.75899 | 21.9242 | 310 | 23.6053 | 100.633 | 101.371 | -0.016 | 5.393 | 100 | 100
-0.059259 | 333.1557 | Empty Chamber Correction for 500 to 0 | 60 | 1475420679.1 | 118 | 20161002 09:04:39 | 1475420678.6 | -0.00015331 | -10.2287 | 335.7109 | 28.2996 | 28.0819 | -4.5189 | 2.0078 | -0.010468 | 3.98 | 165.2463 | 26.2332 | 24.9881 | 3.1774 | 63.0541 | 1.9806 | 3.1411 | 1.1968 | -0.034953 | 0.039238 | 3.0307 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010946 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.12 | 843.1003 | 0.843 | 0.196 | 1475420678.6 | 333.128 | 322.849 | 23.4951 | 23.6448 | 600.048 | 84.0977 | 0.19988 | 24.7953 | 24.9881 | 999.9 | 10023.1 | 0 | 5.1025 | 999.812 | 0.9 | 0.1 | 0 | 2.3407 | 0 | 135.125 | 39.875 | 43.062 | 41.75 | 41.5 | 40 | 2.2814 | 0.31047 | -0.032342 | 135.2938 | 15 | 2 | 550.223 | 708.647 | 21.9238 | 27.3268 | 30.0001 | 27.3047 | 27.273 | 21.4 | 19.7149 | 0.75899 | 21.9327 | 305 | 23.6063 | 100.633 | 101.371 | -0.016 | 5.393 | 100 | 100
-0.033001 | 329.826 | Empty Chamber Correction for 500 to 0 | 61 | 1475420681.1 | 120 | 20161002 09:04:41 | 1475420680.6 | -0.00015344 | -10.2228 | 332.3897 | 28.0196 | 27.8017 | -4.5129 | 2.0052 | -0.010468 | 3.9747 | 165.1986 | 26.2351 | 24.9938 | 3.1785 | 63.0532 | 1.9806 | 3.1412 | 1.1979 | -0.034888 | 0.039165 | 3.0261 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010942 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.821 | 842.8491 | 0.843 | 0.196 | 1475420680.6 | 329.805 | 319.534 | 23.4956 | 23.6454 | 600.142 | 84.0976 | 0.19993 | 24.7959 | 24.9938 | 999.9 | 10004.4 | 0 | 5.1025 | 999.821 | 0.9 | 0.1 | 0 | 2.1552 | 0 | 135.099 | 39.875 | 43.062 | 41.687 | 41.437 | 40 | 2.2666 | 0.87974 | -1.2763 | 135.3013 | 15 | 2 | 554.167 | 708.746 | 21.9251 | 27.3268 | 30 | 27.3047 | 27.273 | 21.2687 | 19.7149 | 0.75899 | 21.9327 | 305 | 23.6028 | 100.633 | 101.371 | -0.016 | 5.393 | 100 | 100
-0.062076 | 326.5442 | Empty Chamber Correction for 500 to 0 | 62 | 1475420683.1 | 122 | 20161002 09:04:43 | 1475420682.6 | -0.00015309 | -10.2724 | 329.1153 | 27.7435 | 27.5244 | -4.5091 | 2.0035 | -0.010454 | 3.9714 | 165.1974 | 26.2392 | 24.9877 | 3.1773 | 63.0412 | 1.9806 | 3.1418 | 1.1967 | -0.034847 | 0.039118 | 3.0232 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010998 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.15 | 843.1274 | 0.843 | 0.196 | 1475420682.6 | 326.516 | 316.193 | 23.4956 | 23.6451 | 599.964 | 84.0972 | 0.19997 | 24.799 | 24.9877 | 999.9 | 9992.5 | 0 | 5.1025 | 1000.16 | 0.9 | 0.1 | 0 | 2.2391 | 0 | 134.879 | 39.875 | 43.125 | 41.687 | 41.437 | 39.937 | 2.3019 | 0.73314 | -1.7494 | 135.2338 | 15 | 2 | 557.472 | 709.018 | 21.9283 | 27.3268 | 30 | 27.3047 | 27.273 | 21.0705 | 19.7149 | 0.75899 | 21.9392 | 300 | 23.6021 | 100.633 | 101.373 | -0.016 | 5.393 | 100 | 100
0.056355 | 323.2152 | Empty Chamber Correction for 500 to 0 | 63 | 1475420685.1 | 124 | 20161002 09:04:45 | 1475420684.6 | -0.0001526 | -10.1736 | 325.7923 | 27.4632 | 27.2461 | -4.5071 | 2.0026 | -0.010422 | 3.9696 | 165.1969 | 26.2402 | 24.987 | 3.1772 | 63.0412 | 1.9807 | 3.1419 | 1.1965 | -0.034825 | 0.039094 | 3.0217 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.01088 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.16 | 843.135 | 0.843 | 0.196 | 1475420684.6 | 323.217 | 312.995 | 23.4966 | 23.6456 | 600.053 | 84.0967 | 0.1999 | 24.7996 | 24.987 | 999.9 | 9986.25 | 0 | 5.1025 | 1000.12 | 0.90003 | 0.099971 | 0 | 2.2726 | 0 | 135.001 | 39.812 | 43.062 | 41.687 | 41.5 | 39.937 | 2.3197 | 1.0546 | -1.8206 | 135.1881 | 15 | 2 | 558.018 | 708.845 | 21.9314 | 27.3268 | 30 | 27.3047 | 27.273 | 20.8677 | 19.7149 | 0.75899 | 21.9392 | 295 | 23.5984 | 100.633 | 101.372 | -0.016 | 5.393 | 100 | 100
0.13504 | 319.9093 | Empty Chamber Correction for 500 to 0 | 64 | 1475420687.1 | 126 | 20161002 09:04:47 | 1475420686.6 | -0.0001533 | -10.1147 | 322.4935 | 27.1853 | 26.9693 | -4.5035 | 2.001 | -0.010463 | 3.9664 | 165.1968 | 26.2388 | 24.9909 | 3.178 | 63.0511 | 1.9807 | 3.1414 | 1.1973 | -0.034786 | 0.03905 | 3.019 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010811 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.15 | 843.1266 | 0.843 | 0.196 | 1475420686.6 | 319.931 | 309.767 | 23.4963 | 23.646 | 599.985 | 84.0974 | 0.19991 | 24.7969 | 24.9909 | 999.9 | 9975 | 0 | 5.1025 | 1000.15 | 0.9 | 0.1 | 0 | 2.6944 | 0 | 134.581 | 39.875 | 43 | 41.687 | 41.5 | 39.937 | 2.3508 | 0.68291 | -1.975 | 135.1232 | 15 | 2 | 560.925 | 708.943 | 21.9343 | 27.3268 | 30.0001 | 27.3047 | 27.273 | 20.6964 | 19.7149 | 0.75899 | 21.9392 | 295 | 23.5942 | 100.632 | 101.372 | -0.016 | 5.393 | 100 | 100
0.0068158 | 316.7054 | Empty Chamber Correction for 500 to 0 | 65 | 1475420689.1 | 128 | 20161002 09:04:49 | 1475420688.6 | -0.00015245 | -10.2637 | 319.2929 | 26.9164 | 26.6974 | -4.5068 | 2.0024 | -0.0104 | 3.9694 | 165.1976 | 26.2359 | 24.995 | 3.1787 | 63.062 | 1.9808 | 3.1411 | 1.1979 | -0.034821 | 0.03909 | 3.0215 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010987 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.15 | 843.1266 | 0.843 | 0.196 | 1475420688.6 | 316.695 | 306.381 | 23.4974 | 23.6463 | 599.882 | 84.0996 | 0.20031 | 24.7953 | 24.995 | 999.9 | 9985 | 0 | 5.1025 | 999.812 | 0.9 | 0.1 | 0 | 2.3224 | 0 | 134.794 | 39.875 | 43.062 | 41.687 | 41.5 | 39.937 | 2.3704 | -0.17186 | -2.2552 | 135.0399 | 15 | 2 | 553.368 | 708.803 | 21.9375 | 27.3268 | 30.0001 | 27.3047 | 27.2736 | 20.4852 | 19.7149 | 0.75899 | 21.946 | 290 | 23.5892 | 100.632 | 101.371 | -0.016 | 5.393 | 100 | 100
-0.046173 | 313.4415 | Empty Chamber Correction for 500 to 0 | 66 | 1475420691.1 | 130 | 20161002 09:04:51 | 1475420690.6 | -0.00015068 | -10.3372 | 316.0321 | 26.6415 | 26.4211 | -4.5102 | 2.004 | -0.010283 | 3.9725 | 165.258 | 26.2355 | 24.9934 | 3.1784 | 63.0632 | 1.9809 | 3.1412 | 1.1975 | -0.034859 | 0.039132 | 3.0241 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011074 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.15 | 843.1266 | 0.843 | 0.196 | 1475420690.6 | 313.418 | 303.036 | 23.4988 | 23.6459 | 600.14 | 84.0999 | 0.20008 | 24.796 | 24.9934 | 999.9 | 9995.62 | 0 | 5.1025 | 999.817 | 0.9 | 0.1 | 0 | 2.1623 | 0 | 134.714 | 39.875 | 43.062 | 41.687 | 41.5 | 39.937 | 2.3437 | -0.24272 | -2.8313 | 134.9779 | 15 | 2 | 550.286 | 708.645 | 21.9403 | 27.3268 | 30.0001 | 27.3047 | 27.2747 | 20.3783 | 19.7149 | 0.75899 | 21.946 | 285 | 23.5914 | 100.633 | 101.371 | -0.016 | 5.393 | 100 | 100
0.12483 | 310.0673 | Empty Chamber Correction for 500 to 0 | 67 | 1475420693.1 | 132 | 20161002 09:04:53 | 1475420692.6 | -0.00014985 | -10.1857 | 312.6586 | 26.3563 | 26.1395 | -4.5175 | 2.0072 | -0.010225 | 3.979 | 165.1991 | 26.2317 | 24.9939 | 3.1785 | 63.0652 | 1.9809 | 3.141 | 1.1976 | -0.034939 | 0.039222 | 3.0297 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010898 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.821 | 842.8491 | 0.843 | 0.196 | 1475420692.6 | 310.087 | 299.857 | 23.499 | 23.6453 | 600.127 | 84.0978 | 0.19964 | 24.7951 | 24.9939 | 999.9 | 10018.8 | 0 | 5.1025 | 999.821 | 0.9 | 0.1 | 0 | 2.2484 | 0 | 134.599 | 39.875 | 43 | 41.687 | 41.5 | 39.937 | 2.3334 | 0.0070086 | -3.1323 | 134.9009 | 15 | 2 | 561.534 | 709.048 | 21.943 | 27.3268 | 30.0001 | 27.3047 | 27.2753 | 20.1575 | 19.7149 | 0.75899 | 21.9505 | 285 | 23.5887 | 100.634 | 101.372 | -0.016 | 5.393 | 100 | 100
0.14977 | 306.7757 | Empty Chamber Correction for 500 to 0 | 68 | 1475420695.1 | 134 | 20161002 09:04:55 | 1475420694.6 | -0.00014754 | -10.1809 | 309.3732 | 26.0791 | 25.8623 | -4.5154 | 2.0063 | -0.010059 | 3.9773 | 165.1986 | 26.2304 | 24.9993 | 3.1795 | 63.0701 | 1.9809 | 3.1409 | 1.1986 | -0.034916 | 0.039196 | 3.0281 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010889 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.15 | 843.1274 | 0.843 | 0.196 | 1475420694.6 | 306.802 | 296.574 | 23.4997 | 23.6438 | 599.892 | 84.0965 | 0.19995 | 24.7941 | 24.9993 | 999.9 | 10012.5 | 0 | 5.1025 | 1000.16 | 0.9 | 0.1 | 0 | 2.245 | 0 | 134.607 | 39.875 | 43.062 | 41.687 | 41.437 | 39.937 | 2.3313 | -0.55199 | -2.688 | 134.788 | 15 | 2 | 564.102 | 709.098 | 21.9455 | 27.3268 | 30.0001 | 27.3047 | 27.2753 | 19.9382 | 19.7149 | 0.75899 | 21.9505 | 280 | 23.586 | 100.634 | 101.373 | -0.016 | 5.393 | 100 | 100
0.11573 | 303.5212 | Empty Chamber Correction for 500 to 0 | 69 | 1475420697.1 | 136 | 20161002 09:04:57 | 1475420696.6 | -0.00014776 | -10.2353 | 306.1225 | 25.8052 | 25.5874 | -4.5176 | 2.0073 | -0.010086 | 3.9792 | 165.2478 | 26.2296 | 24.9925 | 3.1783 | 63.0771 | 1.9811 | 3.1407 | 1.1972 | -0.03494 | 0.039223 | 3.0298 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010957 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.826 | 842.8533 | 0.843 | 0.196 | 1475420696.6 | 303.539 | 293.258 | 23.5012 | 23.6455 | 599.949 | 84.0966 | 0.2003 | 24.7934 | 24.9925 | 999.9 | 10019.4 | 0 | 5.1025 | 1000.17 | 0.89997 | 0.10003 | 0 | 2.404 | 0 | 134.438 | 39.875 | 43 | 41.625 | 41.437 | 39.937 | 2.3329 | -0.46166 | -2.1684 | 134.7328 | 15 | 2 | 555.664 | 708.727 | 21.9482 | 27.3268 | 30 | 27.3047 | 27.2753 | 19.7388 | 19.7149 | 0.75899 | 21.9505 | 275 | 23.5807 | 100.634 | 101.373 | -0.016 | 5.393 | 100 | 100
-0.010047 | 300.2826 | Empty Chamber Correction for 500 to 0 | 70 | 1475420699.1 | 138 | 20161002 09:04:59 | 1475420698.6 | -0.00014706 | -10.382 | 302.8889 | 25.5327 | 25.3118 | -4.518 | 2.0074 | -0.010044 | 3.9796 | 165.1991 | 26.2274 | 24.9888 | 3.1776 | 63.0843 | 1.9811 | 3.1405 | 1.1964 | -0.034944 | 0.039228 | 3.0301 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011127 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.18 | 843.1519 | 0.843 | 0.196 | 1475420698.6 | 300.269 | 289.843 | 23.5018 | 23.6454 | 600.009 | 84.0968 | 0.2004 | 24.792 | 24.9888 | 999.9 | 10020.6 | 0 | 5.1025 | 1000.18 | 0.9 | 0.1 | 0 | 2.5193 | 0 | 134.378 | 39.812 | 43 | 41.625 | 41.5 | 39.875 | 2.3352 | -0.27668 | -2.0791 | 134.6739 | 15 | 2 | 543.868 | 708.234 | 21.9508 | 27.3268 | 30.0001 | 27.3047 | 27.2753 | 19.57 | 19.7149 | 0.75899 | 21.9541 | 275 | 23.5795 | 100.634 | 101.373 | -0.016 | 5.393 | 100 | 100
-0.16653 | 297.0227 | Empty Chamber Correction for 500 to 0 | 71 | 1475420701.1 | 140 | 20161002 09:05:01 | 1475420700.6 | -0.00014552 | -10.5598 | 299.6334 | 25.2582 | 25.0337 | -4.5196 | 2.0082 | -0.0099278 | 3.9811 | 165.1999 | 26.2265 | 24.9958 | 3.1789 | 63.0824 | 1.9811 | 3.1405 | 1.1978 | -0.034962 | 0.039248 | 3.0313 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011338 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.17 | 843.1434 | 0.843 | 0.196 | 1475420700.6 | 296.97 | 286.367 | 23.5013 | 23.6434 | 600.001 | 84.097 | 0.19999 | 24.7921 | 24.9958 | 999.9 | 10025.6 | 0 | 5.1025 | 999.831 | 0.9 | 0.1 | 0 | 2.1172 | 0 | 134.56 | 39.812 | 42.937 | 41.687 | 41.437 | 39.937 | 2.2958 | -0.095727 | -1.1987 | 134.6088 | 15 | 2 | 542.276 | 708.259 | 21.9526 | 27.3268 | 30.0002 | 27.3047 | 27.2753 | 19.4044 | 19.7149 | 0.75899 | 21.9541 | 270 | 23.5736 | 100.634 | 101.373 | -0.016 | 5.393 | 100 | 100
-0.05653 | 293.6531 | Empty Chamber Correction for 500 to 0 | 72 | 1475420703.1 | 142 | 20161002 09:05:03 | 1475420702.6 | -0.00014635 | -10.4697 | 296.2781 | 24.9753 | 24.7519 | -4.5037 | 2.0011 | -0.0099887 | 3.967 | 165.2597 | 26.2317 | 24.9929 | 3.1783 | 63.0821 | 1.9811 | 3.1404 | 1.1973 | -0.034788 | 0.039052 | 3.0191 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011231 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.17 | 843.1434 | 0.843 | 0.196 | 1475420702.6 | 293.628 | 283.116 | 23.501 | 23.6439 | 600.041 | 84.0969 | 0.19977 | 24.7919 | 24.9929 | 999.9 | 9975.62 | 0 | 5.1025 | 999.834 | 0.9 | 0.1 | 0 | 2.5972 | 0 | 134.356 | 39.812 | 43 | 41.625 | 41.437 | 39.875 | 2.3176 | 0.237 | -0.65962 | 134.5655 | 15 | 2 | 548.114 | 708.382 | 21.9543 | 27.3268 | 30.0001 | 27.3047 | 27.2753 | 19.1951 | 19.7149 | 0.75899 | 21.9594 | 265 | 23.5737 | 100.633 | 101.372 | -0.016 | 5.393 | 100 | 100
0.061723 | 290.313 | Empty Chamber Correction for 500 to 0 | 73 | 1475420705.1 | 144 | 20161002 09:05:05 | 1475420704.6 | -0.00014472 | -10.3712 | 292.9495 | 24.695 | 24.4731 | -4.4925 | 1.9961 | -0.0098887 | 3.9572 | 165.2026 | 26.2342 | 24.9865 | 3.1771 | 63.0855 | 1.9812 | 3.1405 | 1.1959 | -0.034666 | 0.038915 | 3.0106 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011118 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.842 | 842.8668 | 0.843 | 0.196 | 1475420704.6 | 290.318 | 279.906 | 23.5024 | 23.6437 | 600.069 | 84.0977 | 0.20002 | 24.7922 | 24.9865 | 999.9 | 9940.62 | 0 | 5.1025 | 999.842 | 0.9 | 0.1 | 0 | 2.0956 | 0 | 134.556 | 39.812 | 43 | 41.625 | 41.437 | 39.875 | 2.3054 | 0.22862 | -0.82308 | 134.5226 | 15 | 2 | 549.486 | 708.234 | 21.9558 | 27.3268 | 30 | 27.3047 | 27.2753 | 19.0186 | 19.7149 | 0.75899 | 21.9594 | 265 | 23.5684 | 100.634 | 101.371 | -0.016 | 5.393 | 100 | 100
-0.049823 | 287.0549 | Empty Chamber Correction for 500 to 0 | 74 | 1475420707.1 | 146 | 20161002 09:05:07 | 1475420706.6 | -0.00014459 | -10.5037 | 289.6811 | 24.4196 | 24.1963 | -4.5191 | 2.0079 | -0.009883 | 3.9807 | 165.2021 | 26.229 | 24.9848 | 3.1768 | 63.0772 | 1.9812 | 3.141 | 1.1956 | -0.034956 | 0.039241 | 3.0309 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011272 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.19 | 843.1612 | 0.843 | 0.196 | 1475420706.6 | 287.032 | 276.486 | 23.5028 | 23.644 | 599.955 | 84.0984 | 0.19988 | 24.7948 | 24.9848 | 999.9 | 10023.8 | 0 | 5.1025 | 1000.19 | 0.9 | 0.1 | 0 | 2.4309 | 0 | 134.266 | 39.812 | 43 | 41.625 | 41.437 | 39.875 | 2.3286 | 0.26455 | -0.64369 | 134.4967 | 15 | 2 | 553.321 | 708.333 | 21.9576 | 27.3268 | 30.0001 | 27.3047 | 27.2753 | 18.8031 | 19.7149 | 0.75899 | 21.9594 | 260 | 23.5677 | 100.635 | 101.371 | -0.016 | 5.393 | 100 | 100
-0.063315 | 283.7823 | Empty Chamber Correction for 500 to 0 | 75 | 1475420709.1 | 148 | 20161002 09:05:09 | 1475420708.6 | -0.00014583 | -10.5375 | 286.4123 | 24.1439 | 23.92 | -4.5214 | 2.009 | -0.0099699 | 3.9827 | 165.2021 | 26.2291 | 24.9833 | 3.1765 | 63.076 | 1.9812 | 3.141 | 1.1953 | -0.034982 | 0.03927 | 3.0327 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011312 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.15 | 843.1256 | 0.843 | 0.196 | 1475420708.6 | 283.756 | 273.177 | 23.5029 | 23.6453 | 599.993 | 84.0978 | 0.20004 | 24.7951 | 24.9833 | 999.9 | 10031.2 | 0 | 5.1025 | 999.848 | 0.9 | 0.1 | 0 | 2.463 | 0 | 134.111 | 39.812 | 43 | 41.625 | 41.437 | 39.875 | 2.3532 | 0.53034 | -1.072 | 134.4661 | 15 | 2 | 553.609 | 708.333 | 21.9597 | 27.3268 | 30.0001 | 27.3047 | 27.2753 | 18.6025 | 20.0062 | 0.75899 | 21.9688 | 255 | 23.5608 | 100.635 | 101.37 | -0.016 | 5.393 | 100 | 100
-0.0033087 | 280.4921 | Empty Chamber Correction for 500 to 0 | 76 | 1475420711.1 | 150 | 20161002 09:05:11 | 1475420710.6 | -0.00014532 | -10.4973 | 283.1301 | 23.867 | 23.6437 | -4.5162 | 2.0067 | -0.0099326 | 3.9781 | 165.2511 | 26.2298 | 24.9846 | 3.1768 | 63.0784 | 1.9812 | 3.1409 | 1.1956 | -0.034925 | 0.039206 | 3.0287 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011268 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.847 | 842.871 | 0.843 | 0.196 | 1475420710.6 | 280.481 | 269.943 | 23.5026 | 23.6445 | 600.005 | 84.0968 | 0.20027 | 24.7941 | 24.9846 | 999.9 | 10015 | 0 | 5.1025 | 999.847 | 0.9 | 0.1 | 0 | 2.3196 | 0 | 134.215 | 39.812 | 43 | 41.625 | 41.437 | 39.875 | 2.3655 | 0.53193 | -1.4925 | 134.4331 | 15 | 2 | 548.596 | 708.135 | 21.9622 | 27.3268 | 30.0001 | 27.3047 | 27.2753 | 18.4709 | 20.0062 | 0.75899 | 21.9688 | 255 | 23.5604 | 100.635 | 101.37 | -0.016 | 5.393 | 100 | 100
-0.094124 | 277.2165 | Empty Chamber Correction for 500 to 0 | 77 | 1475420713.1 | 152 | 20161002 09:05:13 | 1475420712.6 | -0.00011214 | -10.609 | 279.8622 | 23.5914 | 23.3658 | -4.5095 | 2.0038 | -0.0076535 | 3.9744 | 165.2596 | 26.2194 | 24.9939 | 3.1785 | 63.0801 | 1.9812 | 3.1408 | 1.1973 | -0.034855 | 0.039128 | 3.0238 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0114 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.835 | 842.8609 | 0.843 | 0.196 | 1475420712.6 | 277.185 | 266.545 | 23.5032 | 23.6127 | 600.004 | 84.0967 | 0.19993 | 24.794 | 24.9939 | 999.9 | 9995 | 0 | 5.1025 | 1000.15 | 0.90003 | 0.09997 | 0 | 2.2108 | 0 | 134.506 | 39.812 | 43 | 41.625 | 41.5 | 39.875 | 2.3453 | -0.39743 | -0.13422 | 134.4292 | 15 | 2 | 550.186 | 708.037 | 21.966 | 27.3268 | 30 | 27.3047 | 27.2753 | 18.2654 | 20.0062 | 0.75899 | 21.9777 | 250 | 23.5626 | 100.635 | 101.371 | -0.016 | 5.393 | 100 | 100
0.0047524 | 273.8589 | Empty Chamber Correction for 500 to 0 | 78 | 1475420715.1 | 154 | 20161002 09:05:15 | 1475420714.6 | -7.3946e-05 | -10.53 | 276.5006 | 23.3083 | 23.0853 | -4.522 | 2.0095 | -0.00504 | 3.9879 | 165.2034 | 26.2003 | 24.9988 | 3.1795 | 63.0604 | 1.9805 | 3.1407 | 1.1989 | -0.034995 | 0.039285 | 3.0336 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011307 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.839 | 842.8643 | 0.843 | 0.196 | 1475420714.6 | 273.855 | 263.306 | 23.4944 | 23.5666 | 600.07 | 84.0974 | 0.19995 | 24.7931 | 24.9988 | 999.9 | 10035 | 0 | 5.1025 | 1000.15 | 0.90003 | 0.09997 | 0 | 1.983 | 0 | 134.578 | 39.812 | 43 | 41.625 | 41.437 | 39.875 | 2.3119 | -0.33204 | -0.18639 | 134.426 | 15 | 2 | 551.793 | 707.938 | 21.9694 | 27.3268 | 30.0001 | 27.3047 | 27.2753 | 18.0582 | 20.0062 | 0.75899 | 21.9777 | 245 | 23.5706 | 100.635 | 101.371 | -0.016 | 5.393 | 100 | 100
0.13431 | 270.535 | Empty Chamber Correction for 500 to 0 | 79 | 1475420717.1 | 156 | 20161002 09:05:17 | 1475420716.6 | -6.9633e-05 | -10.4199 | 273.1839 | 23.0289 | 22.8081 | -4.5178 | 2.0077 | -0.0047498 | 3.9845 | 165.2528 | 26.2014 | 24.9865 | 3.1771 | 63.0102 | 1.9791 | 3.1409 | 1.198 | -0.034949 | 0.039234 | 3.0304 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011173 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.15 | 843.1256 | 0.843 | 0.196 | 1475420716.6 | 270.564 | 260.125 | 23.477 | 23.545 | 599.987 | 84.0982 | 0.20009 | 24.7942 | 24.9865 | 999.9 | 10021.9 | 0 | 5.1025 | 999.854 | 0.9 | 0.1 | 0 | 2.2184 | 0 | 134.301 | 39.75 | 43 | 41.562 | 41.312 | 39.875 | 2.3309 | -0.85263 | 0.083521 | 134.4079 | 15 | 2 | 551.442 | 708.061 | 21.973 | 27.3268 | 30.0001 | 27.3047 | 27.2753 | 17.8809 | 20.0062 | 0.75899 | 21.9777 | 245 | 23.5706 | 100.635 | 101.371 | -0.016 | 5.393 | 100 | 100
0.031252 | 267.3038 | Empty Chamber Correction for 500 to 0 | 80 | 1475420719.1 | 158 | 20161002 09:05:19 | 1475420718.6 | -8.3964e-05 | -10.5438 | 269.9678 | 22.7579 | 22.5335 | -4.502 | 2.0006 | -0.0057227 | 3.9696 | 165.2511 | 26.2126 | 24.9846 | 3.1768 | 62.9625 | 1.9777 | 3.1411 | 1.199 | -0.034775 | 0.039038 | 3.0182 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011323 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.851 | 842.8744 | 0.843 | 0.196 | 1475420718.6 | 267.306 | 256.739 | 23.4611 | 23.5431 | 599.955 | 84.0985 | 0.20007 | 24.7956 | 24.9846 | 999.9 | 9971.88 | 0 | 5.1025 | 999.833 | 0.9 | 0.1 | 0 | 1.8882 | 0 | 134.761 | 39.812 | 43 | 41.625 | 41.437 | 39.875 | 2.2946 | -1.0322 | 0.45559 | 134.4164 | 15 | 2 | 548.046 | 708.012 | 21.9771 | 27.3268 | 30.0001 | 27.3059 | 27.2753 | 17.6612 | 20.0062 | 0.75899 | 21.9839 | 240 | 23.5706 | 100.634 | 101.371 | -0.016 | 5.393 | 100 | 100
0.031293 | 263.9884 | Empty Chamber Correction for 500 to 0 | 81 | 1475420721.1 | 160 | 20161002 09:05:21 | 1475420720.6 | -9.4955e-05 | -10.5642 | 266.6621 | 22.4793 | 22.254 | -4.4953 | 1.9976 | -0.0064611 | 3.963 | 165.258 | 26.2164 | 24.9895 | 3.1777 | 62.9371 | 1.9767 | 3.1407 | 1.201 | -0.034701 | 0.038955 | 3.0131 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011347 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.843 | 842.8676 | 0.843 | 0.196 | 1475420720.6 | 263.99 | 253.404 | 23.4485 | 23.5412 | 600.184 | 84.0989 | 0.19985 | 24.7934 | 24.9895 | 999.9 | 9950.62 | 0 | 5.1025 | 1000.15 | 0.9 | 0.1 | 0 | 2.0853 | 0 | 134.688 | 39.812 | 43 | 41.562 | 41.375 | 39.875 | 2.2957 | -1.0592 | 0.56718 | 134.4205 | 15 | 2 | 549.053 | 707.642 | 21.9802 | 27.3268 | 30.0001 | 27.307 | 27.2753 | 17.5553 | 20.0062 | 0.75899 | 21.9839 | 235 | 23.5706 | 100.635 | 101.371 | -0.016 | 5.393 | 100 | 100
0.066875 | 260.688 | Empty Chamber Correction for 500 to 0 | 82 | 1475420723.1 | 162 | 20161002 09:05:23 | 1475420722.6 | -0.00010343 | -10.5487 | 263.3606 | 22.2009 | 21.9765 | -4.5062 | 2.0023 | -0.0070288 | 3.972 | 165.2026 | 26.2147 | 24.9935 | 3.1784 | 62.9169 | 1.9759 | 3.1405 | 1.2026 | -0.034819 | 0.039087 | 3.0213 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011325 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.15 | 843.1256 | 0.843 | 0.196 | 1475420722.6 | 260.698 | 250.123 | 23.4392 | 23.5402 | 600.039 | 84.0989 | 0.19971 | 24.7921 | 24.9935 | 999.9 | 9984.38 | 0 | 5.1025 | 1000.15 | 0.90003 | 0.09997 | 0 | 2.3026 | 0 | 134.735 | 39.75 | 43 | 41.625 | 41.375 | 39.875 | 2.3098 | -0.2144 | 0.88168 | 134.425 | 15 | 2 | 560.061 | 708.16 | 21.9832 | 27.3268 | 30.0001 | 27.307 | 27.2753 | 17.3284 | 20.0062 | 0.75899 | 21.9911 | 235 | 23.5706 | 100.635 | 101.371 | -0.016 | 5.393 | 100 | 100
0.013837 | 257.4596 | Empty Chamber Correction for 500 to 0 | 83 | 1475420725.1 | 164 | 20161002 09:05:25 | 1475420724.6 | -0.00011005 | -10.6221 | 260.1359 | 21.929 | 21.7031 | -4.5092 | 2.0037 | -0.0074651 | 3.9743 | 165.2017 | 26.2179 | 25.0024 | 3.1801 | 62.8943 | 1.9754 | 3.1408 | 1.2048 | -0.034851 | 0.039124 | 3.0236 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011413 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.14 | 843.1172 | 0.843 | 0.196 | 1475420724.6 | 257.456 | 246.803 | 23.4331 | 23.5406 | 599.857 | 84.0984 | 0.19997 | 24.7937 | 25.0024 | 999.9 | 9993.75 | 0 | 5.1025 | 999.841 | 0.9 | 0.1 | 0 | 2.2152 | 0 | 134.507 | 39.75 | 43 | 41.562 | 41.312 | 39.875 | 2.2799 | 0.33878 | 0.33203 | 134.4581 | 15 | 2 | 562.337 | 708.414 | 21.9857 | 27.3268 | 30.0001 | 27.307 | 27.2759 | 17.124 | 20.0062 | 0.75899 | 21.9911 | 230 | 23.5706 | 100.635 | 101.37 | -0.016 | 5.393 | 100 | 100
0.16328 | 254.1712 | Empty Chamber Correction for 500 to 0 | 84 | 1475420727.1 | 166 | 20161002 09:05:27 | 1475420726.6 | -0.00011376 | -10.4918 | 256.8505 | 21.6521 | 21.4291 | -4.5124 | 2.0051 | -0.0077165 | 3.9769 | 165.2613 | 26.2204 | 25.0003 | 3.1797 | 62.8748 | 1.9749 | 3.1411 | 1.2048 | -0.034886 | 0.039163 | 3.026 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011257 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.19 | 843.1603 | 0.843 | 0.196 | 1475420726.6 | 254.205 | 243.684 | 23.4279 | 23.539 | 599.985 | 84.0982 | 0.2003 | 24.7952 | 25.0003 | 999.9 | 10003.8 | 0 | 5.1025 | 999.842 | 0.9 | 0.1 | 0 | 2.2883 | 0 | 134.106 | 39.75 | 43 | 41.625 | 41.437 | 39.812 | 2.2847 | 1.1607 | -1.0659 | 134.439 | 15 | 2 | 553.881 | 707.935 | 21.9879 | 27.3268 | 30.0001 | 27.307 | 27.277 | 16.9109 | 20.0062 | 0.75899 | 21.9911 | 225 | 23.572 | 100.636 | 101.37 | -0.016 | 5.393 | 100 | 100
0.12738 | 250.9945 | Empty Chamber Correction for 500 to 0 | 85 | 1475420729.1 | 168 | 20161002 09:05:29 | 1475420728.6 | -0.00011745 | -10.5476 | 253.6712 | 21.3844 | 21.1608 | -4.5254 | 2.0109 | -0.0079768 | 3.9881 | 165.2004 | 26.2176 | 24.9905 | 3.1779 | 62.8627 | 1.9746 | 3.1411 | 1.2033 | -0.035028 | 0.039322 | 3.0359 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011328 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.838 | 842.8634 | 0.843 | 0.196 | 1475420728.6 | 251.019 | 240.442 | 23.4232 | 23.5379 | 600.005 | 84.0994 | 0.20024 | 24.7953 | 24.9905 | 999.9 | 10044.4 | 0 | 5.1025 | 999.838 | 0.9 | 0.1 | 0 | 2.3982 | 0 | 134.245 | 39.75 | 42.937 | 41.562 | 41.375 | 39.875 | 2.3184 | 0.71728 | -0.7532 | 134.3853 | 15 | 2 | 547.745 | 707.524 | 21.9908 | 27.3268 | 30 | 27.307 | 27.2775 | 16.7392 | 20.0062 | 0.75899 | 21.9912 | 225 | 23.5746 | 100.636 | 101.369 | -0.016 | 5.393 | 100 | 100
-0.12955 | 247.847 | Empty Chamber Correction for 500 to 0 | 86 | 1475420731.1 | 170 | 20161002 09:05:31 | 1475420730.6 | -0.00012033 | -10.8259 | 250.5343 | 21.1203 | 20.8904 | -4.517 | 2.0071 | -0.0081705 | 3.9805 | 165.2019 | 26.2243 | 24.9909 | 3.178 | 62.845 | 1.9744 | 3.1417 | 1.2036 | -0.034935 | 0.039218 | 3.0294 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011654 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.19 | 843.1612 | 0.843 | 0.196 | 1475420730.6 | 247.807 | 236.952 | 23.4207 | 23.5382 | 600.042 | 84.1012 | 0.19984 | 24.7985 | 24.9909 | 999.9 | 10017.5 | 0 | 5.1025 | 1000.18 | 0.9 | 0.1 | 0 | 2.4641 | 0 | 134.323 | 39.75 | 42.937 | 41.562 | 41.437 | 39.812 | 2.3284 | 0.34536 | -0.31395 | 134.3998 | 15 | 2 | 552.428 | 707.525 | 21.9928 | 27.3268 | 30.0002 | 27.307 | 27.2776 | 16.5625 | 20.0062 | 0.75899 | 21.9912 | 220 | 23.5739 | 100.636 | 101.37 | -0.016 | 5.393 | 100 | 100
-0.073665 | 244.533 | Empty Chamber Correction for 500 to 0 | 87 | 1475420733.1 | 172 | 20161002 09:05:33 | 1475420732.6 | -0.00012043 | -10.7901 | 247.2368 | 20.842 | 20.6119 | -4.4978 | 1.9986 | -0.0081695 | 3.9635 | 165.2022 | 26.2328 | 24.996 | 3.1789 | 62.8297 | 1.9743 | 3.1422 | 1.2047 | -0.034726 | 0.038983 | 3.0148 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011612 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.15 | 843.1256 | 0.843 | 0.196 | 1475420732.6 | 244.507 | 233.688 | 23.4194 | 23.537 | 600.029 | 84.1 | 0.19979 | 24.8014 | 24.996 | 999.9 | 9957.5 | 0 | 5.1025 | 999.844 | 0.9 | 0.1 | 0 | 2.0406 | 0 | 134.112 | 39.75 | 43 | 41.562 | 41.375 | 39.812 | 2.3231 | -0.10779 | -0.79477 | 134.3922 | 15 | 2 | 559.522 | 707.92 | 21.9935 | 27.3268 | 30.0002 | 27.307 | 27.2776 | 16.3433 | 20.0062 | 0.75899 | 21.9971 | 215 | 23.5779 | 100.636 | 101.371 | -0.016 | 5.393 | 100 | 100
0.051484 | 241.2275 | Empty Chamber Correction for 500 to 0 | 88 | 1475420735.1 | 174 | 20161002 09:05:35 | 1475420734.6 | -0.00012215 | -10.6844 | 243.9296 | 20.5627 | 20.3354 | -4.5088 | 2.0035 | -0.0082865 | 3.9732 | 165.2511 | 26.2283 | 24.9945 | 3.1786 | 62.8308 | 1.974 | 3.1418 | 1.2046 | -0.034847 | 0.039118 | 3.0232 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.01149 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.855 | 842.8778 | 0.843 | 0.196 | 1475420734.6 | 241.233 | 230.518 | 23.4172 | 23.5365 | 599.935 | 84.0973 | 0.20028 | 24.7991 | 24.9945 | 999.9 | 9992.5 | 0 | 5.1025 | 999.855 | 0.9 | 0.1 | 0 | 2.0984 | 0 | 134.369 | 39.687 | 43 | 41.562 | 41.375 | 39.812 | 2.3385 | -0.62653 | -0.25316 | 134.383 | 15 | 2 | 554.662 | 708.043 | 21.9939 | 27.3268 | 30.0001 | 27.307 | 27.2776 | 16.1589 | 20.0062 | 0.75899 | 21.9971 | 215 | 23.5818 | 100.636 | 101.371 | -0.016 | 5.393 | 100 | 100
-0.2268 | 238.0594 | Empty Chamber Correction for 500 to 0 | 89 | 1475420737.1 | 176 | 20161002 09:05:37 | 1475420736.6 | -0.0001242 | -10.9844 | 240.7631 | 20.2957 | 20.0624 | -4.5154 | 2.0064 | -0.0084263 | 3.9789 | 165.2528 | 26.2242 | 24.9927 | 3.1783 | 62.8353 | 1.9738 | 3.1412 | 1.2045 | -0.034918 | 0.039199 | 3.0283 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011845 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.861 | 842.8828 | 0.843 | 0.196 | 1475420736.6 | 237.995 | 226.98 | 23.4147 | 23.536 | 599.941 | 84.0973 | 0.20019 | 24.7961 | 24.9927 | 999.9 | 10013.1 | 0 | 5.1025 | 999.858 | 0.9 | 0.1 | 0 | 1.8759 | 0 | 134.418 | 39.75 | 43 | 41.562 | 41.437 | 39.812 | 2.3091 | -0.71663 | -0.8852 | 134.3569 | 15 | 2 | 546.487 | 707.772 | 21.9954 | 27.3268 | 30 | 27.307 | 27.2776 | 15.9383 | 20.0062 | 0.75899 | 21.9971 | 210 | 23.5847 | 100.637 | 101.371 | -0.016 | 5.393 | 100 | 100
-0.31807 | 234.7495 | Empty Chamber Correction for 500 to 0 | 90 | 1475420739.1 | 178 | 20161002 09:05:39 | 1475420738.6 | -0.00012372 | -11.0967 | 237.4673 | 20.018 | 19.7816 | -4.5009 | 1.9999 | -0.0083958 | 3.9661 | 165.263 | 26.2268 | 24.9911 | 3.178 | 62.84 | 1.9738 | 3.1409 | 1.2042 | -0.03476 | 0.039021 | 3.0172 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011975 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.18 | 843.1519 | 0.843 | 0.196 | 1475420738.6 | 234.662 | 223.538 | 23.4142 | 23.535 | 600.096 | 84.0981 | 0.2 | 24.7946 | 24.9911 | 999.9 | 9967.5 | 0 | 5.1025 | 999.843 | 0.9 | 0.1 | 0 | 2.2808 | 0 | 133.954 | 39.75 | 42.937 | 41.562 | 41.375 | 39.812 | 2.295 | -0.68067 | -0.38926 | 134.3082 | 15 | 2 | 545.275 | 707.476 | 21.9975 | 27.3268 | 30 | 27.307 | 27.2776 | 15.7349 | 20.0062 | 0.75899 | 22.0012 | 205 | 23.5882 | 100.637 | 101.371 | -0.016 | 5.393 | 100 | 100
-0.26308 | 231.3796 | Empty Chamber Correction for 500 to 0 | 91 | 1475420741.1 | 180 | 20161002 09:05:41 | 1475420740.6 | -0.00012617 | -11.0621 | 234.0981 | 19.734 | 19.4986 | -4.5083 | 2.0032 | -0.0085602 | 3.9725 | 165.2049 | 26.223 | 24.9925 | 3.1783 | 62.8476 | 1.9738 | 3.1406 | 1.2045 | -0.03484 | 0.039111 | 3.0228 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011938 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.855 | 842.8778 | 0.843 | 0.196 | 1475420740.6 | 231.306 | 220.216 | 23.4141 | 23.5373 | 600.071 | 84.0982 | 0.19984 | 24.7925 | 24.9925 | 999.9 | 9990.62 | 0 | 5.1025 | 999.855 | 0.9 | 0.1 | 0 | 2.4138 | 0 | 134.157 | 39.75 | 42.937 | 41.562 | 41.437 | 39.812 | 2.3047 | -0.56782 | -0.9908 | 134.2425 | 15 | 2 | 552.003 | 707.624 | 21.999 | 27.3268 | 30.0002 | 27.307 | 27.2776 | 15.5961 | 20.0062 | 0.75899 | 22.0012 | 205 | 23.5862 | 100.636 | 101.371 | -0.016 | 5.393 | 100 | 100
-0.22828 | 228.0396 | Empty Chamber Correction for 500 to 0 | 92 | 1475420743.1 | 182 | 20161002 09:05:43 | 1475420742.6 | -0.00012461 | -11.0477 | 230.7627 | 19.4528 | 19.2178 | -4.5091 | 2.0036 | -0.0084472 | 3.9733 | 165.2047 | 26.2207 | 24.9979 | 3.1793 | 62.8537 | 1.9738 | 3.1403 | 1.2055 | -0.034849 | 0.039121 | 3.0234 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011917 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.15 | 843.1264 | 0.843 | 0.196 | 1475420742.6 | 227.975 | 216.898 | 23.4143 | 23.536 | 599.951 | 84.0981 | 0.19985 | 24.791 | 24.9979 | 999.9 | 9993.12 | 0 | 5.1025 | 1000.15 | 0.90003 | 0.09997 | 0 | 2.105 | 0 | 134.143 | 39.75 | 42.937 | 41.562 | 41.375 | 39.812 | 2.2657 | -0.099921 | -2.3557 | 134.2263 | 15 | 2 | 556.969 | 707.796 | 22.0006 | 27.3274 | 30.0002 | 27.307 | 27.2776 | 15.3893 | 20.0062 | 0.75899 | 22.0053 | 200 | 23.5892 | 100.636 | 101.37 | -0.016 | 5.393 | 100 | 100
-0.03089 | 224.6818 | Empty Chamber Correction for 500 to 0 | 93 | 1475420745.1 | 184 | 20161002 09:05:45 | 1475420744.6 | -0.0001247 | -10.8695 | 227.4132 | 19.1706 | 18.9391 | -4.5035 | 2.0011 | -0.0084583 | 3.9683 | 165.2029 | 26.2239 | 24.9942 | 3.1786 | 62.8457 | 1.9737 | 3.1406 | 1.2049 | -0.034788 | 0.039053 | 3.0192 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011706 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.15 | 843.1256 | 0.843 | 0.196 | 1475420744.6 | 224.667 | 213.768 | 23.4134 | 23.5352 | 599.918 | 84.0983 | 0.20023 | 24.7926 | 24.9942 | 999.9 | 9975.62 | 0 | 5.1025 | 999.856 | 0.9 | 0.1 | 0 | 2.4647 | 0 | 133.978 | 39.75 | 42.937 | 41.562 | 41.375 | 39.75 | 2.2881 | 0.3115 | -2.3203 | 134.1753 | 15 | 2 | 551.36 | 707.599 | 22.0021 | 27.3286 | 30 | 27.307 | 27.2776 | 15.1694 | 20.0062 | 0.75899 | 22.0053 | 195 | 23.5902 | 100.636 | 101.371 | -0.016 | 5.393 | 100 | 100
0.075112 | 221.3971 | Empty Chamber Correction for 500 to 0 | 94 | 1475420747.1 | 186 | 20161002 09:05:47 | 1475420746.6 | -0.00012566 | -10.7829 | 224.133 | 18.8941 | 18.6645 | -4.5041 | 2.0014 | -0.0085275 | 3.9688 | 165.2528 | 26.2254 | 24.9907 | 3.1779 | 62.84 | 1.9736 | 3.1407 | 1.2043 | -0.034795 | 0.03906 | 3.0196 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011607 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.855 | 842.8778 | 0.843 | 0.196 | 1475420746.6 | 221.409 | 210.6 | 23.4125 | 23.5352 | 600.097 | 84.0984 | 0.20026 | 24.7935 | 24.9907 | 999.9 | 9977.5 | 0 | 5.1025 | 999.855 | 0.9 | 0.1 | 0 | 2.4643 | 0 | 133.849 | 39.75 | 43 | 41.5 | 41.375 | 39.812 | 2.2776 | 0.31583 | -2.0639 | 134.1126 | 15 | 2 | 542.683 | 707.155 | 22.0038 | 27.3291 | 30 | 27.307 | 27.2776 | 14.9861 | 20.0062 | 0.75899 | 22.0053 | 195 | 23.5983 | 100.637 | 101.371 | -0.016 | 5.393 | 100 | 100
-0.093417 | 218.1754 | Empty Chamber Correction for 500 to 0 | 95 | 1475420749.1 | 188 | 20161002 09:05:49 | 1475420748.6 | -0.00012504 | -10.9726 | 220.9151 | 18.6228 | 18.3892 | -4.5065 | 2.0024 | -0.0084828 | 3.971 | 165.204 | 26.2243 | 24.9922 | 3.1782 | 62.8379 | 1.9736 | 3.1408 | 1.2046 | -0.034821 | 0.03909 | 3.0214 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011828 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.2 | 843.1687 | 0.843 | 0.196 | 1475420748.6 | 218.145 | 207.146 | 23.4121 | 23.5342 | 600.05 | 84.0986 | 0.19967 | 24.7937 | 24.9922 | 999.9 | 9985 | 0 | 5.1025 | 999.865 | 0.9 | 0.1 | 0 | 2.487 | 0 | 133.658 | 39.75 | 42.937 | 41.562 | 41.375 | 39.812 | 2.2895 | 0.13779 | -1.7881 | 134.0412 | 15 | 2 | 550.979 | 707.378 | 22.0056 | 27.3291 | 30.0001 | 27.307 | 27.2776 | 14.7642 | 20.0062 | 0.75899 | 22.0105 | 190 | 23.5982 | 100.637 | 101.37 | -0.016 | 5.393 | 100 | 100
-0.16173 | 214.8954 | Empty Chamber Correction for 500 to 0 | 96 | 1475420751.1 | 190 | 20161002 09:05:51 | 1475420750.6 | -0.00012502 | -11.0617 | 217.637 | 18.3467 | 18.1116 | -4.5122 | 2.0049 | -0.0084744 | 3.9759 | 165.263 | 26.2247 | 24.9978 | 3.1793 | 62.8319 | 1.9736 | 3.1411 | 1.2057 | -0.034883 | 0.039159 | 3.0258 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011933 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.19 | 843.1603 | 0.843 | 0.196 | 1475420750.6 | 214.848 | 203.759 | 23.4117 | 23.5338 | 599.978 | 84.0999 | 0.19979 | 24.7953 | 24.9978 | 999.9 | 10002.5 | 0 | 5.1025 | 999.851 | 0.9 | 0.1 | 0 | 2.1786 | 0 | 133.951 | 39.687 | 42.937 | 41.5 | 41.437 | 39.812 | 2.2954 | 0.32099 | -1.5954 | 133.9673 | 15 | 2 | 559.386 | 707.649 | 22.0071 | 27.3291 | 30.0002 | 27.307 | 27.2776 | 14.6492 | 20.0062 | 0.75899 | 22.0105 | 185 | 23.599 | 100.637 | 101.369 | -0.016 | 5.393 | 100 | 100
-0.10926 | 211.602 | Empty Chamber Correction for 500 to 0 | 97 | 1475420753.1 | 192 | 20161002 09:05:53 | 1475420752.6 | -0.00012339 | -11.0292 | 214.3423 | 18.069 | 17.8351 | -4.5223 | 2.0094 | -0.0083612 | 3.985 | 165.2054 | 26.2238 | 25.0003 | 3.1797 | 62.8236 | 1.9737 | 3.1416 | 1.206 | -0.034994 | 0.039283 | 3.0335 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011899 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.869 | 842.8896 | 0.843 | 0.196 | 1475420752.6 | 211.568 | 200.513 | 23.4128 | 23.5333 | 600.015 | 84.0993 | 0.2004 | 24.7983 | 25.0003 | 999.9 | 10034.4 | 0 | 5.1025 | 999.869 | 0.9 | 0.1 | 0 | 2.4342 | 0 | 133.577 | 39.75 | 42.937 | 41.562 | 41.375 | 39.812 | 2.2996 | 0.048626 | -1.595 | 133.9156 | 15 | 2 | 550.651 | 707.328 | 22.0092 | 27.3291 | 30.0002 | 27.3076 | 27.2776 | 14.4195 | 20.0062 | 0.75899 | 22.0129 | 185 | 23.6014 | 100.637 | 101.369 | -0.016 | 5.393 | 100 | 100
-0.13152 | 208.3287 | Empty Chamber Correction for 500 to 0 | 98 | 1475420755.1 | 194 | 20161002 09:05:55 | 1475420754.6 | -0.00012564 | -11.0718 | 211.0817 | 17.7941 | 17.5586 | -4.5099 | 2.0039 | -0.0085163 | 3.974 | 165.2613 | 26.2291 | 24.9986 | 3.1794 | 62.8236 | 1.9738 | 3.1417 | 1.2057 | -0.034858 | 0.039131 | 3.0241 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011949 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.86 | 842.882 | 0.843 | 0.196 | 1475420754.6 | 208.289 | 197.191 | 23.4136 | 23.5363 | 599.999 | 84.0993 | 0.20005 | 24.7988 | 24.9986 | 999.9 | 9995.62 | 0 | 5.1025 | 999.866 | 0.9 | 0.1 | 0 | 2.2815 | 0 | 133.9 | 39.687 | 42.937 | 41.562 | 41.375 | 39.75 | 2.3049 | -0.00050941 | -1.1331 | 133.8886 | 15 | 2 | 547.994 | 707.303 | 22.0111 | 27.3291 | 30 | 27.3081 | 27.2776 | 14.1876 | 20.0062 | 0.75899 | 22.0129 | 180 | 23.602 | 100.637 | 101.37 | -0.016 | 5.393 | 100 | 100
-0.095297 | 205.0383 | Empty Chamber Correction for 500 to 0 | 99 | 1475420757.1 | 196 | 20161002 09:05:57 | 1475420756.6 | -0.00012431 | -11.0556 | 207.7841 | 17.5163 | 17.2823 | -4.5299 | 2.0128 | -0.008429 | 3.9916 | 165.2613 | 26.2223 | 24.9967 | 3.1791 | 62.8265 | 1.9738 | 3.1416 | 1.2053 | -0.035077 | 0.039377 | 3.0393 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.01193 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.86 | 842.882 | 0.843 | 0.196 | 1475420756.6 | 205.008 | 193.927 | 23.4135 | 23.5349 | 600.005 | 84.1006 | 0.19999 | 24.7982 | 24.9967 | 999.9 | 10058.1 | 0 | 5.1025 | 1000.17 | 0.9 | 0.1 | 0 | 1.9955 | 0 | 134.173 | 39.687 | 42.875 | 41.5 | 41.375 | 39.812 | 2.2848 | 0.04694 | -0.49846 | 133.8889 | 15 | 2 | 547.625 | 707.303 | 22.0124 | 27.3291 | 30 | 27.3087 | 27.2776 | 13.9751 | 20.0062 | 0.75899 | 22.0129 | 175 | 23.6052 | 100.637 | 101.369 | -0.016 | 5.393 | 100 | 100
-0.08951 | 201.7437 | Empty Chamber Correction for 500 to 0 | 100 | 1475420759.1 | 198 | 20161002 09:05:59 | 1475420758.6 | -0.00012371 | -11.0701 | 204.5009 | 17.2398 | 17.0049 | -4.5196 | 2.0082 | -0.008401 | 3.9826 | 165.2059 | 26.2236 | 24.9869 | 3.1772 | 62.8292 | 1.9737 | 3.1414 | 1.2035 | -0.034964 | 0.03925 | 3.0314 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011947 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.876 | 842.8955 | 0.843 | 0.196 | 1475420758.6 | 201.715 | 190.621 | 23.4128 | 23.5336 | 600.057 | 84.1018 | 0.19996 | 24.7972 | 24.9869 | 999.9 | 10025.6 | 0 | 5.1025 | 999.876 | 0.9 | 0.1 | 0 | 2.3345 | 0 | 133.667 | 39.75 | 42.937 | 41.5 | 41.375 | 39.75 | 2.3129 | -0.16735 | -0.70301 | 133.8474 | 15 | 2 | 548.692 | 707.219 | 22.0136 | 27.3291 | 30 | 27.3093 | 27.2787 | 13.7963 | 20.0062 | 0.75899 | 22.0152 | 175 | 23.6063 | 100.638 | 101.369 | -0.016 | 5.393 | 100 | 100
-0.21642 | 198.4458 | Empty Chamber Correction for 500 to 0 | 101 | 1475420761.1 | 200 | 20161002 09:06:01 | 1475420760.6 | -0.00012369 | -11.2183 | 201.221 | 16.9633 | 16.7242 | -4.4991 | 1.9991 | -0.0083891 | 3.9645 | 165.2629 | 26.2294 | 24.9948 | 3.1787 | 62.8292 | 1.9737 | 3.1413 | 1.205 | -0.03474 | 0.038998 | 3.0157 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012123 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.868 | 842.8887 | 0.843 | 0.196 | 1475420760.6 | 198.385 | 187.142 | 23.4121 | 23.5329 | 599.994 | 84.1019 | 0.19981 | 24.7967 | 24.9948 | 999.9 | 9961.25 | 0 | 5.1025 | 1000.16 | 0.90003 | 0.09997 | 0 | 2.3649 | 0 | 133.986 | 39.687 | 42.875 | 41.5 | 41.312 | 39.75 | 2.293 | -0.17702 | 0.34687 | 133.8539 | 15 | 2 | 555.228 | 707.407 | 22.0145 | 27.3291 | 30.0001 | 27.3093 | 27.2798 | 13.6189 | 20.0062 | 0.75899 | 22.0152 | 170 | 23.6092 | 100.638 | 101.369 | -0.016 | 5.393 | 100 | 100
-0.10092 | 195.0913 | Empty Chamber Correction for 500 to 0 | 102 | 1475420763.1 | 202 | 20161002 09:06:03 | 1475420762.6 | -0.00012041 | -11.1227 | 197.8626 | 16.6804 | 16.4441 | -4.5132 | 2.0054 | -0.0081535 | 3.9772 | 165.2047 | 26.2241 | 25.0048 | 3.1806 | 62.8273 | 1.9737 | 3.1414 | 1.2069 | -0.034894 | 0.039172 | 3.0266 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.01201 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.865 | 842.8862 | 0.843 | 0.196 | 1475420762.6 | 195.06 | 183.913 | 23.4118 | 23.5294 | 599.954 | 84.1024 | 0.20033 | 24.7972 | 25.0048 | 999.9 | 10005.6 | 0 | 5.1025 | 999.866 | 0.9 | 0.1 | 0 | 2.2934 | 0 | 134.051 | 39.687 | 42.937 | 41.5 | 41.375 | 39.75 | 2.2871 | -0.30829 | 0.9388 | 133.8632 | 15 | 2 | 549.527 | 707.261 | 22.0155 | 27.3291 | 30.0002 | 27.3093 | 27.2799 | 13.3981 | 19.6928 | 0.75899 | 22.0181 | 165 | 23.6098 | 100.638 | 101.369 | -0.016 | 5.393 | 100 | 100
-0.027374 | 191.7954 | Empty Chamber Correction for 500 to 0 | 103 | 1475420765.1 | 204 | 20161002 09:06:05 | 1475420764.6 | -0.00013641 | -11.0689 | 194.5682 | 16.4023 | 16.1674 | -4.5201 | 2.0084 | -0.0092397 | 3.9822 | 165.2067 | 26.2277 | 25.003 | 3.1802 | 62.8312 | 1.9738 | 3.1414 | 1.2065 | -0.034968 | 0.039255 | 3.0317 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011946 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.864 | 842.8854 | 0.843 | 0.196 | 1475420764.6 | 191.782 | 180.688 | 23.4132 | 23.5464 | 600.057 | 84.1008 | 0.20017 | 24.7968 | 25.003 | 999.9 | 10026.9 | 0 | 5.1025 | 999.864 | 0.9 | 0.1 | 0 | 2.4652 | 0 | 133.581 | 39.687 | 42.937 | 41.5 | 41.375 | 39.75 | 2.2977 | 0.0071043 | 0.57125 | 133.8446 | 15 | 2 | 544.183 | 706.99 | 22.0163 | 27.3291 | 30.0001 | 27.3093 | 27.2799 | 13.2094 | 19.6928 | 0.75899 | 22.0181 | 165 | 23.6078 | 100.637 | 101.369 | -0.016 | 5.393 | 100 | 100
-0.14525 | 188.5485 | Empty Chamber Correction for 500 to 0 | 104 | 1475420767.1 | 206 | 20161002 09:06:07 | 1475420766.6 | -0.00016876 | -11.2078 | 191.3281 | 16.1288 | 15.8907 | -4.5199 | 2.0082 | -0.011431 | 3.9799 | 165.2072 | 26.2419 | 25.0053 | 3.1807 | 62.8373 | 1.9742 | 3.1418 | 1.2065 | -0.034963 | 0.039249 | 3.0314 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.01211 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.869 | 842.8896 | 0.843 | 0.196 | 1475420766.6 | 188.504 | 177.265 | 23.4189 | 23.5837 | 600.031 | 84.0995 | 0.19968 | 24.7989 | 25.0053 | 999.9 | 10025.6 | 0 | 5.1025 | 999.873 | 0.9 | 0.1 | 0 | 2.4016 | 0 | 133.367 | 39.75 | 42.937 | 41.5 | 41.375 | 39.812 | 2.2908 | 0.11294 | 0.039829 | 133.8615 | 15 | 2 | 550.726 | 707.31 | 22.0171 | 27.3291 | 30 | 27.3093 | 27.2799 | 12.9804 | 19.6928 | 0.75899 | 22.0181 | 160 | 23.6034 | 100.637 | 101.37 | -0.016 | 5.393 | 100 | 100
-0.076254 | 185.2297 | Empty Chamber Correction for 500 to 0 | 105 | 1475420769.1 | 208 | 20161002 09:06:09 | 1475420768.6 | -0.00018368 | -11.1587 | 188.0161 | 15.8498 | 15.6126 | -4.5182 | 2.0074 | -0.01245 | 3.9775 | 165.2629 | 26.2492 | 25.0072 | 3.181 | 62.8709 | 1.9753 | 3.1419 | 1.2057 | -0.034944 | 0.039227 | 3.03 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012052 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.87 | 842.8904 | 0.843 | 0.196 | 1475420768.6 | 185.202 | 174.008 | 23.4322 | 23.6116 | 599.934 | 84.1003 | 0.2001 | 24.7997 | 25.0072 | 999.9 | 10020 | 0 | 5.1025 | 1000.19 | 0.9 | 0.1 | 0 | 2.5514 | 0 | 133.456 | 39.625 | 42.875 | 41.437 | 41.125 | 39.75 | 2.2993 | 0.017214 | -0.81231 | 133.8638 | 15 | 2 | 550.447 | 707.408 | 22.0184 | 27.3291 | 30.0001 | 27.3093 | 27.2799 | 12.7688 | 19.6928 | 0.75899 | 21.9711 | 155 | 23.6034 | 100.637 | 101.37 | -0.016 | 5.393 | 100 | 100
-0.061816 | 181.9266 | Empty Chamber Correction for 500 to 0 | 106 | 1475420771.1 | 210 | 20161002 09:06:11 | 1475420770.6 | -0.0001735 | -11.1645 | 184.7284 | 15.5724 | 15.3343 | -4.5008 | 1.9997 | -0.011767 | 3.9627 | 165.2083 | 26.2485 | 25.0101 | 3.1816 | 62.9188 | 1.9767 | 3.1416 | 1.2049 | -0.034754 | 0.039014 | 3.0167 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012059 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.877 | 842.8963 | 0.843 | 0.196 | 1475420770.6 | 181.903 | 170.709 | 23.4483 | 23.6177 | 600.113 | 84.0989 | 0.20017 | 24.7982 | 25.0101 | 999.9 | 9965.62 | 0 | 5.1025 | 1000.22 | 0.89997 | 0.10003 | 0 | 2.3374 | 0 | 133.62 | 39.625 | 42.875 | 41.437 | 41.125 | 39.75 | 2.3115 | 0.55805 | -1.6705 | 133.8124 | 15 | 2 | 544.318 | 706.916 | 22.0122 | 27.3291 | 30.0002 | 27.3093 | 27.2799 | 12.6274 | 19.6928 | 0.75899 | 21.9711 | 155 | 23.6034 | 100.638 | 101.37 | -0.016 | 5.393 | 100 | 100
-0.15851 | 178.6403 | Empty Chamber Correction for 500 to 0 | 107 | 1475420773.1 | 212 | 20161002 09:06:13 | 1475420772.6 | -0.00016253 | -11.2823 | 181.446 | 15.2956 | 15.0551 | -4.5023 | 2.0004 | -0.011033 | 3.9648 | 165.2629 | 26.2435 | 25.0095 | 3.1815 | 62.9526 | 1.9776 | 3.1414 | 1.2039 | -0.034771 | 0.039033 | 3.0179 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012199 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.877 | 842.8963 | 0.843 | 0.196 | 1475420772.6 | 178.593 | 167.283 | 23.4597 | 23.6184 | 600.068 | 84.0987 | 0.1997 | 24.7972 | 25.0095 | 999.9 | 9970.62 | 0 | 5.1025 | 999.867 | 0.9 | 0.1 | 0 | 2.3565 | 0 | 133.532 | 39.625 | 42.937 | 41.5 | 41.312 | 39.75 | 2.2978 | 0.37817 | -1.9266 | 133.7478 | 15 | 2 | 551.591 | 707.113 | 21.9951 | 27.3291 | 30.0004 | 27.3093 | 27.2799 | 12.4146 | 19.6928 | 0.75899 | 21.9616 | 150 | 23.6034 | 100.638 | 101.37 | -0.016 | 5.393 | 100 | 100
-0.052772 | 175.3062 | Empty Chamber Correction for 500 to 0 | 108 | 1475420775.1 | 214 | 20161002 09:06:15 | 1475420774.6 | -0.00015625 | -11.1963 | 178.1093 | 15.0145 | 14.7765 | -4.514 | 2.0056 | -0.010624 | 3.9755 | 165.2663 | 26.2367 | 25.0027 | 3.1802 | 62.9788 | 1.9783 | 3.1412 | 1.2019 | -0.0349 | 0.039178 | 3.027 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012093 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.22 | 843.1856 | 0.843 | 0.196 | 1475420774.6 | 175.286 | 164.061 | 23.4678 | 23.6204 | 599.929 | 84.0994 | 0.19994 | 24.7962 | 25.0027 | 999.9 | 10007.5 | 0 | 5.1025 | 999.879 | 0.9 | 0.1 | 0 | 2.3809 | 0 | 133.051 | 39.625 | 42.937 | 41.5 | 41.25 | 39.75 | 2.2901 | 0.092328 | -3.2754 | 133.6912 | 15 | 2 | 556.124 | 707.359 | 21.9788 | 27.3291 | 30.0004 | 27.3093 | 27.2799 | 12.1899 | 19.6928 | 0.75899 | 21.9616 | 145 | 23.6034 | 100.638 | 101.369 | -0.016 | 5.393 | 100 | 100
0.053492 | 172.0141 | Empty Chamber Correction for 500 to 0 | 109 | 1475420777.1 | 216 | 20161002 09:06:17 | 1475420776.6 | -0.00015074 | -11.1095 | 174.8228 | 14.7376 | 14.5014 | -4.5124 | 2.0049 | -0.010262 | 3.9744 | 165.2108 | 26.2331 | 24.9978 | 3.1793 | 63.0012 | 1.9788 | 3.1409 | 1.2004 | -0.034883 | 0.039159 | 3.0258 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011994 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.881 | 842.8997 | 0.843 | 0.196 | 1475420776.6 | 172.021 | 160.886 | 23.4736 | 23.6208 | 600.023 | 84.1002 | 0.20019 | 24.7946 | 24.9978 | 999.9 | 10002.5 | 0 | 5.1025 | 999.881 | 0.9 | 0.1 | 0 | 2.3769 | 0 | 133.453 | 39.687 | 42.875 | 41.437 | 41.437 | 39.75 | 2.3153 | 0.16104 | -3.2091 | 133.6055 | 15 | 2 | 551.575 | 707.039 | 21.9692 | 27.3291 | 30.0003 | 27.3093 | 27.2799 | 12.0022 | 19.6928 | 0.75899 | 21.9616 | 145 | 23.6034 | 100.638 | 101.37 | -0.016 | 5.393 | 100 | 100
-0.060234 | 168.7754 | Empty Chamber Correction for 500 to 0 | 110 | 1475420779.1 | 218 | 20161002 09:06:19 | 1475420778.6 | -0.00014654 | -11.2441 | 171.5883 | 14.4651 | 14.2262 | -4.5138 | 2.0056 | -0.0099836 | 3.976 | 165.2679 | 26.2305 | 24.9948 | 3.1787 | 63.0163 | 1.9792 | 3.1407 | 1.1995 | -0.034898 | 0.039177 | 3.0269 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012153 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.894 | 842.9107 | 0.843 | 0.196 | 1475420778.6 | 168.754 | 157.485 | 23.4775 | 23.6206 | 599.99 | 84.1012 | 0.19995 | 24.7935 | 24.9948 | 999.9 | 10006.9 | 0 | 5.1025 | 999.884 | 0.9 | 0.1 | 0 | 2.0846 | 0 | 133.389 | 39.625 | 42.875 | 41.5 | 41.375 | 39.75 | 2.2948 | -0.35822 | -2.4915 | 133.5096 | 15 | 2 | 552.581 | 707.137 | 21.9618 | 27.3291 | 30.0002 | 27.3093 | 27.2799 | 11.7722 | 19.6928 | 0.75899 | 21.9677 | 140 | 23.6034 | 100.638 | 101.37 | -0.016 | 5.393 | 100 | 100
-0.12407 | 165.505 | Empty Chamber Correction for 500 to 0 | 111 | 1475420781.1 | 220 | 20161002 09:06:21 | 1475420780.6 | -0.00014152 | -11.3286 | 168.3195 | 14.1894 | 13.949 | -4.5193 | 2.008 | -0.0096503 | 3.9811 | 165.2093 | 26.2256 | 24.991 | 3.178 | 63.0306 | 1.9795 | 3.1406 | 1.1985 | -0.034959 | 0.039245 | 3.0311 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.01225 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.2 | 843.1678 | 0.843 | 0.196 | 1475420780.6 | 165.468 | 154.116 | 23.4816 | 23.6198 | 600 | 84.1004 | 0.20012 | 24.7925 | 24.991 | 999.9 | 10024.4 | 0 | 5.1025 | 999.897 | 0.9 | 0.1 | 0 | 2.4153 | 0 | 133.397 | 39.625 | 42.875 | 41.5 | 41.375 | 39.75 | 2.2955 | -0.37555 | -1.8012 | 133.4598 | 15 | 2 | 549.528 | 707.088 | 21.9587 | 27.3291 | 30.0001 | 27.3093 | 27.2799 | 11.6529 | 19.6928 | 0.75899 | 21.9677 | 135 | 23.6034 | 100.638 | 101.369 | -0.016 | 5.393 | 100 | 100
-0.042428 | 162.1803 | Empty Chamber Correction for 500 to 0 | 112 | 1475420783.1 | 222 | 20161002 09:06:23 | 1475420782.6 | -0.00013815 | -11.2668 | 164.9977 | 13.9092 | 13.6703 | -4.5222 | 2.0094 | -0.0094255 | 3.9839 | 165.2093 | 26.2231 | 24.9884 | 3.1775 | 63.0386 | 1.9797 | 3.1405 | 1.1978 | -0.034991 | 0.039281 | 3.0334 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012176 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.23 | 843.194 | 0.843 | 0.196 | 1475420782.6 | 162.164 | 150.875 | 23.4843 | 23.6192 | 600.012 | 84.0993 | 0.20017 | 24.7921 | 24.9884 | 999.9 | 10033.8 | 0 | 5.1025 | 1000.23 | 0.9 | 0.1 | 0 | 2.3775 | 0 | 133.588 | 39.625 | 42.875 | 41.5 | 41.187 | 39.75 | 2.3024 | -0.10725 | -1.3236 | 133.413 | 15 | 2 | 545.061 | 706.891 | 21.9595 | 27.3291 | 30.0001 | 27.3093 | 27.2799 | 11.4188 | 19.6928 | 0.75899 | 21.9741 | 135 | 23.6013 | 100.638 | 101.368 | -0.016 | 5.393 | 100 | 100
-0.10198 | 158.9061 | Empty Chamber Correction for 500 to 0 | 113 | 1475420785.1 | 224 | 20161002 09:06:25 | 1475420784.6 | -0.00013845 | -11.347 | 161.7306 | 13.6338 | 13.3931 | -4.5194 | 2.0081 | -0.009442 | 3.9815 | 165.2103 | 26.2244 | 24.9921 | 3.1782 | 63.0423 | 1.9799 | 3.1406 | 1.1983 | -0.034961 | 0.039247 | 3.0312 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012275 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.889 | 842.9064 | 0.843 | 0.196 | 1475420784.6 | 158.875 | 147.506 | 23.4862 | 23.6214 | 600.002 | 84.0997 | 0.19995 | 24.7925 | 24.9921 | 999.9 | 10025 | 0 | 5.1025 | 1000.23 | 0.9 | 0.1 | 0 | 2.5346 | 0 | 133.557 | 39.625 | 42.875 | 41.437 | 41.312 | 39.687 | 2.2932 | -0.082574 | 0.46568 | 133.3927 | 15 | 2 | 545.629 | 706.867 | 21.9614 | 27.3291 | 30.0001 | 27.3093 | 27.2799 | 11.1855 | 19.6928 | 0.75899 | 21.9741 | 130 | 23.6015 | 100.638 | 101.368 | -0.016 | 5.393 | 100 | 100
-0.10716 | 155.6282 | Empty Chamber Correction for 500 to 0 | 114 | 1475420787.1 | 226 | 20161002 09:06:27 | 1475420786.6 | -0.00013518 | -11.3725 | 158.4599 | 13.3582 | 13.1168 | -4.5156 | 2.0064 | -0.0092119 | 3.9783 | 165.2104 | 26.2237 | 24.9981 | 3.1793 | 63.0516 | 1.9801 | 3.1404 | 1.1992 | -0.03492 | 0.0392 | 3.0284 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012306 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.89 | 842.9073 | 0.843 | 0.196 | 1475420786.6 | 155.596 | 144.203 | 23.4886 | 23.6206 | 600.027 | 84.1004 | 0.19995 | 24.7919 | 24.9981 | 999.9 | 10013.1 | 0 | 5.1025 | 999.885 | 0.9 | 0.1 | 0 | 2.2623 | 0 | 133.435 | 39.687 | 42.937 | 41.437 | 41.312 | 39.75 | 2.2752 | 0.33734 | 0.64581 | 133.3925 | 15 | 2 | 546.658 | 706.891 | 21.9642 | 27.3291 | 30.0001 | 27.3093 | 27.2799 | 10.9662 | 19.6928 | 0.75899 | 21.9741 | 125 | 23.5974 | 100.638 | 101.368 | -0.016 | 5.393 | 100 | 100
-0.10127 | 152.3425 | Empty Chamber Correction for 500 to 0 | 115 | 1475420789.1 | 228 | 20161002 09:06:29 | 1475420788.6 | -0.00013395 | -11.3868 | 155.1728 | 13.0813 | 12.8401 | -4.526 | 2.0111 | -0.0091198 | 3.9876 | 165.2097 | 26.2188 | 25.0047 | 3.1806 | 63.063 | 1.9803 | 3.1402 | 1.2003 | -0.035033 | 0.039328 | 3.0363 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012323 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.881 | 842.8997 | 0.843 | 0.196 | 1475420788.6 | 152.312 | 140.905 | 23.4905 | 23.6213 | 600.012 | 84.1011 | 0.20013 | 24.7904 | 25.0047 | 999.9 | 10045.6 | 0 | 5.1025 | 999.881 | 0.9 | 0.1 | 0 | 2.0527 | 0 | 133.537 | 39.625 | 42.875 | 41.5 | 41.312 | 39.687 | 2.2884 | 0.18218 | 1.0449 | 133.3822 | 15 | 2 | 543.222 | 706.768 | 21.9678 | 27.3291 | 30.0001 | 27.3093 | 27.2799 | 10.7835 | 19.6928 | 0.75899 | 21.9763 | 125 | 23.5984 | 100.639 | 101.368 | -0.016 | 5.393 | 100 | 100
-0.2418 | 149.0775 | Empty Chamber Correction for 500 to 0 | 116 | 1475420791.1 | 230 | 20161002 09:06:31 | 1475420790.6 | -0.00013281 | -11.5486 | 151.9094 | 12.8065 | 12.5622 | -4.5319 | 2.0137 | -0.0090562 | 3.9928 | 165.2646 | 26.2177 | 24.9963 | 3.179 | 63.0668 | 1.9805 | 3.1403 | 1.1985 | -0.035098 | 0.0394 | 3.0408 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012514 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.893 | 842.9098 | 0.843 | 0.196 | 1475420790.6 | 149.012 | 137.443 | 23.4922 | 23.6219 | 599.968 | 84.1033 | 0.20001 | 24.791 | 24.9963 | 999.9 | 10063.8 | 0 | 5.1025 | 999.882 | 0.9 | 0.1 | 0 | 2.4074 | 0 | 133.236 | 39.625 | 42.875 | 41.437 | 41 | 39.75 | 2.2868 | 0.10175 | 0.32 | 133.4204 | 15 | 2 | 543.701 | 706.571 | 21.9706 | 27.3291 | 30 | 27.3093 | 27.2799 | 10.6034 | 19.6928 | 0.75899 | 21.9763 | 120 | 23.5934 | 100.639 | 101.369 | -0.016 | 5.393 | 100 | 100
-0.22866 | 145.7444 | Empty Chamber Correction for 500 to 0 | 117 | 1475420793.1 | 232 | 20161002 09:06:33 | 1475420792.6 | -0.0001315 | -11.5559 | 148.5965 | 12.5275 | 12.2818 | -4.5078 | 2.003 | -0.0089756 | 3.9717 | 165.2092 | 26.2237 | 24.9913 | 3.178 | 63.0728 | 1.9807 | 3.1403 | 1.1974 | -0.034835 | 0.039105 | 3.0224 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012519 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.23 | 843.194 | 0.843 | 0.196 | 1475420792.6 | 145.682 | 134.108 | 23.4939 | 23.6223 | 600.055 | 84.1053 | 0.19994 | 24.791 | 24.9913 | 999.9 | 9988.12 | 0 | 5.1025 | 999.889 | 0.9 | 0.1 | 0 | 2.392 | 0 | 133.174 | 39.625 | 42.875 | 41.437 | 41.25 | 39.75 | 2.304 | 0.38574 | -0.16632 | 133.3863 | 15 | 2 | 544.637 | 706.472 | 21.9729 | 27.3291 | 30.0001 | 27.3093 | 27.2799 | 10.3779 | 19.6928 | 0.75899 | 21.9791 | 115 | 23.5918 | 100.64 | 101.369 | -0.016 | 5.393 | 100 | 100
0.019705 | 142.3546 | Empty Chamber Correction for 500 to 0 | 118 | 1475420795.1 | 234 | 20161002 09:06:35 | 1475420794.6 | -0.0001271 | -11.3265 | 145.2183 | 12.2425 | 12.0011 | -4.4968 | 1.9981 | -0.0086561 | 3.9622 | 165.2097 | 26.2236 | 25.006 | 3.1808 | 63.0839 | 1.9808 | 3.14 | 1.2 | -0.034714 | 0.03897 | 3.014 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012247 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.24 | 843.2024 | 0.843 | 0.196 | 1475420794.6 | 142.355 | 131.012 | 23.4962 | 23.6203 | 600.086 | 84.1039 | 0.19989 | 24.7894 | 25.006 | 999.9 | 9953.75 | 0 | 5.1025 | 1000.24 | 0.9 | 0.1 | 0 | 2.4667 | 0 | 132.944 | 39.625 | 42.875 | 41.437 | 41.312 | 39.75 | 2.3281 | 0.38262 | -1.328 | 133.3555 | 15 | 2 | 547.233 | 706.719 | 21.9749 | 27.3291 | 30 | 27.3093 | 27.2799 | 10.1803 | 19.6928 | 0.75899 | 21.9791 | 115 | 23.5886 | 100.64 | 101.369 | -0.016 | 5.393 | 100 | 100
-0.058513 | 139.1081 | Empty Chamber Correction for 500 to 0 | 119 | 1475420797.1 | 236 | 20161002 09:06:37 | 1475420796.6 | -0.00012369 | -11.4254 | 141.9722 | 11.9685 | 11.7254 | -4.504 | 2.0013 | -0.0084348 | 3.9688 | 165.2138 | 26.222 | 24.998 | 3.1793 | 63.0799 | 1.9809 | 3.1403 | 1.1984 | -0.034794 | 0.039059 | 3.0196 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012368 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.888 | 842.9056 | 0.843 | 0.196 | 1475420796.6 | 139.089 | 127.645 | 23.4977 | 23.6185 | 599.927 | 84.1016 | 0.20001 | 24.7911 | 24.998 | 999.9 | 9976.88 | 0 | 5.1025 | 1000.2 | 0.90003 | 0.09997 | 0 | 2.6385 | 0 | 133.084 | 39.625 | 42.875 | 41.437 | 41.187 | 39.75 | 2.3247 | 0.535 | -2.1407 | 133.3326 | 15 | 2 | 548.871 | 706.965 | 21.9767 | 27.3291 | 30 | 27.3093 | 27.2799 | 9.9488 | 19.6928 | 0.75899 | 21.9791 | 110 | 23.5862 | 100.639 | 101.369 | -0.016 | 5.393 | 100 | 100
-0.17457 | 135.8572 | Empty Chamber Correction for 500 to 0 | 120 | 1475420799.1 | 238 | 20161002 09:06:39 | 1475420798.6 | -0.00012402 | -11.5624 | 138.7271 | 11.6952 | 11.4492 | -4.5033 | 2.001 | -0.0084695 | 3.9681 | 165.2102 | 26.2254 | 24.9891 | 3.1776 | 63.07 | 1.9809 | 3.1409 | 1.1967 | -0.034786 | 0.03905 | 3.019 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012531 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.885 | 842.9031 | 0.843 | 0.196 | 1475420798.6 | 135.809 | 124.23 | 23.4978 | 23.6189 | 600.013 | 84.1034 | 0.20017 | 24.7942 | 24.9891 | 999.9 | 9974.38 | 0 | 5.1025 | 999.892 | 0.9 | 0.1 | 0 | 2.2917 | 0 | 133.27 | 39.562 | 42.812 | 41.437 | 41.125 | 39.75 | 2.3352 | 0.71287 | -2.3921 | 133.2816 | 15 | 2 | 541.367 | 706.35 | 21.9785 | 27.3291 | 30 | 27.3093 | 27.2799 | 9.736 | 19.6928 | 0.75899 | 21.9792 | 105 | 23.5821 | 100.639 | 101.37 | -0.016 | 5.393 | 100 | 100
-0.22738 | 132.5373 | Empty Chamber Correction for 500 to 0 | 121 | 1475420801.1 | 240 | 20161002 09:06:41 | 1475420800.6 | -0.00012403 | -11.6361 | 135.4045 | 11.4154 | 11.1685 | -4.5158 | 2.0065 | -0.0084693 | 3.9791 | 165.21 | 26.221 | 24.9908 | 3.1779 | 63.077 | 1.9811 | 3.1407 | 1.1969 | -0.034922 | 0.039203 | 3.0285 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012618 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.889 | 842.9064 | 0.843 | 0.196 | 1475420800.6 | 132.476 | 120.825 | 23.4986 | 23.6197 | 600.079 | 84.106 | 0.19999 | 24.7934 | 24.9908 | 999.9 | 10013.1 | 0 | 5.1025 | 999.889 | 0.9 | 0.1 | 0 | 2.5812 | 0 | 133.091 | 39.562 | 42.875 | 41.437 | 41.312 | 39.687 | 2.3642 | 0.76113 | -1.6953 | 133.2274 | 15 | 2 | 542.337 | 706.226 | 21.9796 | 27.3291 | 30 | 27.3093 | 27.2799 | 9.5931 | 19.6928 | 0.75899 | 21.9792 | 105 | 23.5774 | 100.639 | 101.369 | -0.016 | 5.393 | 100 | 100
-0.24826 | 129.2124 | Empty Chamber Correction for 500 to 0 | 122 | 1475420803.1 | 242 | 20161002 09:06:43 | 1475420802.6 | -0.00012236 | -11.6777 | 132.0844 | 11.1354 | 10.8877 | -4.5161 | 2.0067 | -0.0083552 | 3.9796 | 165.2561 | 26.2182 | 24.9918 | 3.1781 | 63.0925 | 1.9813 | 3.1403 | 1.1969 | -0.034926 | 0.039208 | 3.0288 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012668 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.891 | 842.9081 | 0.843 | 0.196 | 1475420802.6 | 129.146 | 117.451 | 23.5011 | 23.6206 | 599.922 | 84.1052 | 0.20001 | 24.7909 | 24.9918 | 999.9 | 10014.4 | 0 | 5.1025 | 999.879 | 0.9 | 0.1 | 0 | 2.3881 | 0 | 133.195 | 39.625 | 42.875 | 41.375 | 41.125 | 39.687 | 2.3526 | 0.37701 | -0.75238 | 133.2067 | 15 | 2 | 544.692 | 706.628 | 21.9801 | 27.3291 | 30 | 27.3093 | 27.2805 | 9.3713 | 19.6928 | 0.75899 | 21.9856 | 100 | 23.5739 | 100.639 | 101.369 | -0.016 | 5.393 | 100 | 100
-0.027761 | 125.8418 | Empty Chamber Correction for 500 to 0 | 123 | 1475420805.1 | 244 | 20161002 09:06:45 | 1475420804.6 | -0.00012137 | -11.4763 | 128.7213 | 10.8515 | 10.6079 | -4.512 | 2.0049 | -0.008283 | 3.976 | 165.2092 | 26.2187 | 24.9949 | 3.1787 | 63.092 | 1.9813 | 3.1403 | 1.1975 | -0.034881 | 0.039158 | 3.0257 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012424 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.24 | 843.2024 | 0.843 | 0.196 | 1475420804.6 | 125.831 | 114.341 | 23.5018 | 23.6203 | 600.08 | 84.1024 | 0.20011 | 24.791 | 24.9949 | 999.9 | 10001.9 | 0 | 5.1025 | 999.89 | 0.9 | 0.1 | 0 | 2.3992 | 0 | 133.053 | 39.562 | 42.812 | 41.375 | 41.187 | 39.687 | 2.3923 | 0.328 | -0.42468 | 133.1654 | 15 | 2 | 540.271 | 706.494 | 21.9805 | 27.3291 | 30 | 27.3093 | 27.2816 | 9.1446 | 19.6928 | 0.75899 | 21.9856 | 95 | 23.5744 | 100.639 | 101.369 | -0.016 | 5.393 | 100 | 100
0.080003 | 122.5535 | Empty Chamber Correction for 500 to 0 | 124 | 1475420807.1 | 246 | 20161002 09:06:47 | 1475420806.6 | -0.00011951 | -11.388 | 125.4411 | 10.5749 | 10.3328 | -4.5068 | 2.0026 | -0.0081601 | 3.9715 | 165.2087 | 26.2201 | 24.9915 | 3.1781 | 63.0861 | 1.9811 | 3.1404 | 1.1969 | -0.034824 | 0.039093 | 3.0217 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.01232 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.23 | 843.194 | 0.843 | 0.196 | 1475420806.6 | 122.57 | 111.168 | 23.5007 | 23.6174 | 600.032 | 84.1015 | 0.20005 | 24.7916 | 24.9915 | 999.9 | 9985.62 | 0 | 5.1025 | 999.894 | 0.9 | 0.1 | 0 | 2.091 | 0 | 133.176 | 39.562 | 42.812 | 41.375 | 41.062 | 39.687 | 2.3825 | -0.25508 | -0.15624 | 133.1126 | 15 | 2 | 540.461 | 706.502 | 21.9823 | 27.3291 | 30 | 27.3093 | 27.2821 | 8.9554 | 19.6928 | 0.75899 | 21.9856 | 95 | 23.5691 | 100.639 | 101.369 | -0.016 | 5.393 | 100 | 100
-0.091463 | 119.3527 | Empty Chamber Correction for 500 to 0 | 125 | 1475420809.1 | 248 | 20161002 09:06:49 | 1475420808.6 | -0.00011867 | -11.5805 | 122.2521 | 10.3062 | 10.0596 | -4.4967 | 1.9981 | -0.0080964 | 3.9626 | 165.2118 | 26.2231 | 24.9965 | 3.179 | 63.0874 | 1.9812 | 3.1405 | 1.1978 | -0.034714 | 0.038969 | 3.014 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012552 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.9 | 842.9157 | 0.843 | 0.196 | 1475420808.6 | 119.326 | 107.729 | 23.5013 | 23.6172 | 599.88 | 84.1031 | 0.20007 | 24.792 | 24.9965 | 999.9 | 9953.75 | 0 | 5.1025 | 1000.2 | 0.90003 | 0.09997 | 0 | 2.1315 | 0 | 133.083 | 39.625 | 42.812 | 41.375 | 41 | 39.687 | 2.3373 | -0.67537 | -0.25371 | 133.1071 | 15 | 2 | 537.971 | 706.429 | 21.9847 | 27.3291 | 30 | 27.3093 | 27.2822 | 8.7186 | 19.6928 | 0.75899 | 21.9903 | 90 | 23.5633 | 100.64 | 101.37 | -0.016 | 5.393 | 100 | 100
-0.12705 | 116.0735 | Empty Chamber Correction for 500 to 0 | 126 | 1475420811.1 | 250 | 20161002 09:06:51 | 1475420810.6 | -0.00012076 | -11.6366 | 118.9719 | 10.0296 | 9.7823 | -4.5064 | 2.0024 | -0.0082356 | 3.9711 | 165.2098 | 26.2206 | 25 | 3.1797 | 63.0923 | 1.9813 | 3.1404 | 1.1983 | -0.03482 | 0.039089 | 3.0214 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012619 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.88 | 842.8988 | 0.843 | 0.196 | 1475420810.6 | 116.038 | 104.39 | 23.5026 | 23.6205 | 600.134 | 84.1025 | 0.20001 | 24.7915 | 25 | 999.9 | 9984.38 | 0 | 5.1025 | 999.89 | 0.9 | 0.1 | 0 | 2.2885 | 0 | 132.917 | 39.562 | 42.812 | 41.375 | 41.187 | 39.625 | 2.3398 | -0.77249 | -0.6706 | 133.091 | 15 | 2 | 533.645 | 705.986 | 21.9865 | 27.3291 | 30 | 27.3099 | 27.2822 | 8.5978 | 19.6928 | 0.75899 | 21.9903 | 85 | 23.5626 | 100.641 | 101.37 | -0.016 | 5.393 | 100 | 100
0.035913 | 112.7104 | Empty Chamber Correction for 500 to 0 | 127 | 1475420813.1 | 252 | 20161002 09:06:53 | 1475420812.6 | -0.00012014 | -11.4931 | 115.593 | 9.7447 | 9.5022 | -4.5387 | 2.0167 | -0.0082021 | 3.9996 | 165.2109 | 26.209 | 24.9931 | 3.1784 | 63.1008 | 1.9813 | 3.14 | 1.197 | -0.035173 | 0.039485 | 3.0461 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012449 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.896 | 842.9123 | 0.843 | 0.196 | 1475420812.6 | 112.716 | 101.211 | 23.5029 | 23.6202 | 600.087 | 84.1021 | 0.1997 | 24.7893 | 24.9931 | 999.9 | 10085.6 | 0 | 5.1025 | 999.896 | 0.9 | 0.1 | 0 | 2.4799 | 0 | 132.896 | 39.625 | 42.875 | 41.437 | 41.125 | 39.687 | 2.3452 | -0.50602 | -0.84065 | 133.0696 | 15 | 2 | 544.587 | 706.208 | 21.9882 | 27.3292 | 30.0001 | 27.3099 | 27.2822 | 8.3565 | 19.6928 | 0.75899 | 21.9932 | 85 | 23.5597 | 100.642 | 101.371 | -0.016 | 5.393 | 100 | 100
0.039102 | 109.4135 | Empty Chamber Correction for 500 to 0 | 128 | 1475420815.1 | 254 | 20161002 09:06:55 | 1475420814.6 | -0.00011989 | -11.5103 | 112.3062 | 9.4677 | 9.2244 | -4.5307 | 2.0132 | -0.0081927 | 3.9926 | 165.2662 | 26.2107 | 24.9878 | 3.1774 | 63.1071 | 1.9814 | 3.1398 | 1.1959 | -0.035086 | 0.039387 | 3.04 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012469 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.907 | 842.9216 | 0.843 | 0.196 | 1475420814.6 | 109.42 | 97.8941 | 23.5038 | 23.6209 | 599.87 | 84.1025 | 0.19999 | 24.7884 | 24.9878 | 999.9 | 10060.6 | 0 | 5.1025 | 999.895 | 0.9 | 0.1 | 0 | 2.5298 | 0 | 132.777 | 39.562 | 42.812 | 41.375 | 41.062 | 39.687 | 2.3474 | 0.094803 | -1.5909 | 133.0375 | 15 | 2 | 548.295 | 706.577 | 21.99 | 27.3298 | 30.0001 | 27.3105 | 27.2822 | 8.1129 | 19.6928 | 0.75899 | 21.9932 | 80 | 23.5554 | 100.641 | 101.371 | -0.016 | 5.393 | 100 | 100
0.056833 | 106.1139 | Empty Chamber Correction for 500 to 0 | 129 | 1475420817.1 | 256 | 20161002 09:06:57 | 1475420816.6 | -0.00012156 | -11.5128 | 109.0253 | 9.1912 | 8.9467 | -4.5099 | 2.0039 | -0.0083009 | 3.9741 | 165.2121 | 26.2184 | 24.9927 | 3.1783 | 63.1034 | 1.9815 | 3.1401 | 1.1968 | -0.034858 | 0.039131 | 3.024 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012468 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.2 | 843.1678 | 0.843 | 0.196 | 1475420816.6 | 106.125 | 94.5998 | 23.5045 | 23.6232 | 600.024 | 84.1028 | 0.20018 | 24.79 | 24.9927 | 999.9 | 9995 | 0 | 5.1025 | 999.894 | 0.9 | 0.1 | 0 | 2.4507 | 0 | 132.801 | 39.625 | 42.812 | 41.375 | 41.062 | 39.625 | 2.3338 | 0.17594 | -1.8093 | 132.9962 | 15 | 2 | 542.275 | 706.134 | 21.9919 | 27.3302 | 30 | 27.3116 | 27.2822 | 7.893 | 19.6928 | 0.75899 | 21.9932 | 75 | 23.5513 | 100.64 | 101.37 | -0.016 | 5.393 | 100 | 100
0.011056 | 102.8344 | Empty Chamber Correction for 500 to 0 | 130 | 1475420819.1 | 258 | 20161002 09:06:59 | 1475420818.6 | -0.00012116 | -11.5791 | 105.7507 | 8.9152 | 8.6693 | -4.5101 | 2.004 | -0.0082648 | 3.9743 | 165.2111 | 26.2191 | 24.9994 | 3.1796 | 63.1012 | 1.9815 | 3.1403 | 1.198 | -0.03486 | 0.039134 | 3.0242 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012547 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.2 | 843.1678 | 0.843 | 0.196 | 1475420818.6 | 102.834 | 91.2434 | 23.5047 | 23.623 | 600.051 | 84.1039 | 0.19994 | 24.7909 | 24.9994 | 999.9 | 9995.62 | 0 | 5.1025 | 1000.2 | 0.90003 | 0.09997 | 0 | 2.0749 | 0 | 133.113 | 39.562 | 42.812 | 41.375 | 41.187 | 39.625 | 2.3306 | -0.16091 | -0.8854 | 132.9512 | 15 | 2 | 544.469 | 706.109 | 21.9936 | 27.3297 | 30 | 27.3116 | 27.2822 | 7.7013 | 19.6928 | 0.75899 | 21.9981 | 75 | 23.5503 | 100.641 | 101.37 | -0.016 | 5.393 | 100 | 100
-0.11141 | 99.548 | Empty Chamber Correction for 500 to 0 | 131 | 1475420821.1 | 260 | 20161002 09:07:01 | 1475420820.6 | -0.00011755 | -11.7228 | 102.4745 | 8.6392 | 8.3898 | -4.5023 | 2.0006 | -0.0080098 | 3.9677 | 165.2111 | 26.218 | 25.0067 | 3.1809 | 63.1102 | 1.9816 | 3.1399 | 1.1994 | -0.034776 | 0.039039 | 3.0183 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012721 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.9 | 842.9157 | 0.843 | 0.196 | 1475420820.6 | 99.517 | 87.7812 | 23.5048 | 23.6196 | 599.935 | 84.1053 | 0.20016 | 24.7889 | 25.0067 | 999.9 | 9971.25 | 0 | 5.1025 | 1000.2 | 0.90003 | 0.09997 | 0 | 2.4624 | 0 | 132.95 | 39.562 | 42.875 | 41.375 | 41.187 | 39.687 | 2.3125 | 0.45672 | -0.87241 | 132.9169 | 15 | 2 | 541.605 | 706.208 | 21.9951 | 27.3303 | 30 | 27.3116 | 27.2822 | 7.5171 | 19.6928 | 0.75899 | 21.9981 | 70 | 23.5469 | 100.641 | 101.37 | -0.016 | 5.393 | 100 | 100
-0.0096891 | 96.1584 | Empty Chamber Correction for 500 to 0 | 132 | 1475420823.1 | 262 | 20161002 09:07:03 | 1475420822.6 | -0.00011645 | -11.6412 | 99.0869 | 8.3536 | 8.1062 | -4.5069 | 2.0026 | -0.0079318 | 3.9718 | 165.2111 | 26.2147 | 25.0096 | 3.1815 | 63.1189 | 1.9817 | 3.1396 | 1.1998 | -0.034826 | 0.039095 | 3.0218 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012624 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.9 | 842.9157 | 0.843 | 0.196 | 1475420822.6 | 96.1531 | 84.502 | 23.5059 | 23.6196 | 600.068 | 84.1055 | 0.20006 | 24.7874 | 25.0096 | 999.9 | 9985.62 | 0 | 5.1025 | 999.891 | 0.9 | 0.1 | 0 | 2.4131 | 0 | 133.08 | 39.562 | 42.812 | 41.375 | 41.125 | 39.625 | 2.337 | 0.31698 | -0.44034 | 132.9095 | 15 | 2 | 536.449 | 706.011 | 21.9971 | 27.3309 | 30 | 27.3116 | 27.2822 | 7.2877 | 19.6928 | 0.38755 | 21.9827 | 65 | 23.5421 | 100.641 | 101.37 | -0.016 | 5.393 | 100 | 100
0.16224 | 92.7717 | Empty Chamber Correction for 500 to 0 | 133 | 1475420825.1 | 264 | 20161002 09:07:05 | 1475420824.6 | -0.00011767 | -11.4889 | 95.6999 | 8.0681 | 7.8244 | -4.5151 | 2.0063 | -0.0080243 | 3.979 | 165.211 | 26.2144 | 25.0022 | 3.1801 | 63.1128 | 1.9817 | 3.1399 | 1.1984 | -0.034916 | 0.039196 | 3.0281 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012444 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.9 | 842.9157 | 0.843 | 0.196 | 1475420824.6 | 92.8097 | 81.3101 | 23.5057 | 23.6206 | 600.01 | 84.1064 | 0.19987 | 24.789 | 25.0022 | 999.9 | 10011.2 | 0 | 5.1025 | 999.9 | 0.9 | 0.1 | 0 | 2.0582 | 0 | 132.63 | 39.5 | 42.812 | 41.437 | 41.062 | 39.687 | 2.35 | -0.29235 | -0.60362 | 132.8669 | 15 | 2 | 540.671 | 706.06 | 21.9984 | 27.331 | 30.0001 | 27.3116 | 27.2822 | 7.0894 | 19.9743 | 0.38755 | 21.9827 | 65 | 23.5401 | 100.641 | 101.37 | -0.016 | 5.393 | 100 | 100
0.013548 | 89.5049 | Empty Chamber Correction for 500 to 0 | 134 | 1475420827.1 | 266 | 20161002 09:07:07 | 1475420826.6 | -0.00011365 | -11.6589 | 92.4352 | 7.7929 | 7.546 | -4.5197 | 2.0083 | -0.0077562 | 3.9833 | 165.2594 | 26.2157 | 24.998 | 3.1793 | 63.1013 | 1.9818 | 3.1406 | 1.1975 | -0.034966 | 0.039252 | 3.0316 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012645 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.917 | 842.93 | 0.843 | 0.196 | 1475420826.6 | 89.5057 | 77.8346 | 23.5064 | 23.6174 | 599.895 | 84.107 | 0.19999 | 24.7927 | 24.998 | 999.9 | 10025.6 | 0 | 5.1025 | 999.912 | 0.9 | 0.1 | 0 | 2.0687 | 0 | 132.813 | 39.562 | 42.812 | 41.375 | 41.062 | 39.625 | 2.3311 | -0.86259 | -0.52896 | 132.8475 | 15 | 2 | 543.099 | 706.208 | 21.9948 | 27.3315 | 30.0002 | 27.3116 | 27.2822 | 6.8515 | 19.9743 | 0.38755 | 21.9827 | 60 | 23.5341 | 100.641 | 101.369 | -0.016 | 5.393 | 100 | 100
0.023773 | 86.1677 | Empty Chamber Correction for 500 to 0 | 135 | 1475420829.1 | 268 | 20161002 09:07:09 | 1475420828.6 | -9.8622e-05 | -11.6691 | 89.1115 | 7.5124 | 7.2645 | -4.5059 | 2.0022 | -0.0067364 | 3.9721 | 165.2126 | 26.214 | 24.9911 | 3.178 | 63.0939 | 1.9816 | 3.1406 | 1.1964 | -0.034816 | 0.039084 | 3.0211 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012657 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.905 | 842.9199 | 0.843 | 0.196 | 1475420828.6 | 86.1715 | 74.4943 | 23.5051 | 23.6014 | 600.023 | 84.1031 | 0.20012 | 24.793 | 24.9911 | 999.9 | 9983.12 | 0 | 5.1025 | 1000.25 | 0.89997 | 0.10003 | 0 | 2.4031 | 0 | 132.711 | 39.5 | 42.812 | 41.375 | 41.125 | 39.625 | 2.3195 | -0.40316 | -0.44615 | 132.8248 | 15 | 2 | 537.389 | 706.109 | 21.9893 | 27.3315 | 30.0003 | 27.3116 | 27.2822 | 6.6331 | 19.9743 | 0.38755 | 21.9867 | 55 | 23.5373 | 100.641 | 101.368 | -0.016 | 5.393 | 100 | 100
0.17319 | 82.7954 | Empty Chamber Correction for 500 to 0 | 136 | 1475420831.1 | 270 | 20161002 09:07:11 | 1475420830.6 | -9.2184e-05 | -11.5394 | 85.7418 | 7.2281 | 6.9832 | -4.5089 | 2.0036 | -0.0062922 | 3.9752 | 165.2125 | 26.2085 | 24.993 | 3.1784 | 63.0871 | 1.9811 | 3.1402 | 1.1973 | -0.03485 | 0.039122 | 3.0235 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012503 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.904 | 842.9191 | 0.843 | 0.196 | 1475420830.6 | 82.837 | 71.2923 | 23.5001 | 23.5901 | 600.12 | 84.101 | 0.20007 | 24.7908 | 24.993 | 999.9 | 9993.12 | 0 | 5.1025 | 999.904 | 0.9 | 0.1 | 0 | 2.4532 | 0 | 132.835 | 39.5 | 42.812 | 41.375 | 41.062 | 39.687 | 2.3009 | -0.077217 | -0.54007 | 132.8142 | 15 | 2 | 534.763 | 705.765 | 21.9863 | 27.3315 | 30.0002 | 27.3116 | 27.2822 | 6.4817 | 19.9743 | 0.38755 | 21.9867 | 55 | 23.5399 | 100.641 | 101.368 | -0.016 | 5.393 | 100 | 100
0.074747 | 79.5291 | Empty Chamber Correction for 500 to 0 | 137 | 1475420833.1 | 272 | 20161002 09:07:13 | 1475420832.6 | -9.5129e-05 | -11.6588 | 82.4794 | 6.9532 | 6.7059 | -4.5114 | 2.0047 | -0.0064876 | 3.9772 | 165.261 | 26.2074 | 24.9958 | 3.1789 | 63.0787 | 1.9806 | 3.1399 | 1.1983 | -0.034877 | 0.039152 | 3.0254 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012645 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.908 | 842.9224 | 0.843 | 0.196 | 1475420832.6 | 79.546 | 67.8789 | 23.4939 | 23.5868 | 599.96 | 84.1021 | 0.19985 | 24.7888 | 24.9958 | 999.9 | 10000.6 | 0 | 5.1025 | 999.908 | 0.9 | 0.1 | 0 | 2.6878 | 0 | 132.43 | 39.562 | 42.812 | 41.312 | 41.125 | 39.625 | 2.3182 | 0.69555 | -1.2367 | 132.7918 | 15 | 2 | 539.855 | 705.863 | 21.9858 | 27.3315 | 30.0001 | 27.3116 | 27.2822 | 6.2611 | 19.9743 | 0.38755 | 21.991 | 50 | 23.5423 | 100.641 | 101.367 | -0.016 | 5.393 | 100 | 100
0.10413 | 76.1983 | Empty Chamber Correction for 500 to 0 | 138 | 1475420835.1 | 274 | 20161002 09:07:15 | 1475420834.6 | -9.862e-05 | -11.6497 | 79.1518 | 6.6726 | 6.4257 | -4.5146 | 2.0061 | -0.0067309 | 3.9798 | 165.2627 | 26.2052 | 24.9885 | 3.1775 | 63.0734 | 1.9801 | 3.1394 | 1.1974 | -0.034911 | 0.039191 | 3.0278 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.01263 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.21 | 843.1762 | 0.843 | 0.196 | 1475420834.6 | 76.2226 | 64.5658 | 23.4884 | 23.5847 | 600.021 | 84.1016 | 0.19998 | 24.7862 | 24.9885 | 999.9 | 10010.6 | 0 | 5.1025 | 999.908 | 0.9 | 0.1 | 0 | 2.4224 | 0 | 132.644 | 39.562 | 42.812 | 41.375 | 41.125 | 39.687 | 2.3443 | -0.047221 | -1.0394 | 132.7366 | 15 | 2 | 538.881 | 705.765 | 21.986 | 27.3315 | 30.0001 | 27.3116 | 27.2822 | 6.0337 | 19.9743 | 0.38755 | 21.991 | 45 | 23.545 | 100.64 | 101.366 | -0.016 | 5.393 | 100 | 100
0.16224 | 72.8571 | Empty Chamber Correction for 500 to 0 | 139 | 1475420837.1 | 276 | 20161002 09:07:17 | 1475420836.6 | -0.00010314 | -11.6117 | 75.8196 | 6.3917 | 6.1452 | -4.5088 | 2.0035 | -0.0070405 | 3.9744 | 165.2131 | 26.2067 | 24.9861 | 3.177 | 63.0694 | 1.9798 | 3.1391 | 1.1972 | -0.034848 | 0.03912 | 3.0233 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012589 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.917 | 842.93 | 0.843 | 0.196 | 1475420836.6 | 72.896 | 61.2791 | 23.4851 | 23.5858 | 600.122 | 84.1011 | 0.19993 | 24.7848 | 24.9861 | 999.9 | 9992.5 | 0 | 5.1025 | 999.917 | 0.9 | 0.1 | 0 | 1.8115 | 0 | 132.685 | 39.562 | 42.812 | 41.312 | 41 | 39.625 | 2.326 | 0.055928 | -1.6808 | 132.6983 | 15 | 2 | 541.915 | 705.691 | 21.9867 | 27.3315 | 30.0001 | 27.3116 | 27.2822 | 5.8341 | 19.9743 | 0.38755 | 21.991 | 45 | 23.545 | 100.639 | 101.367 | -0.016 | 5.393 | 100 | 100
0.13971 | 69.5755 | Empty Chamber Correction for 500 to 0 | 140 | 1475420839.1 | 278 | 20161002 09:07:19 | 1475420838.6 | -0.00010936 | -11.6547 | 72.5367 | 6.115 | 5.8682 | -4.5191 | 2.0081 | -0.0074619 | 3.9831 | 165.2135 | 26.2071 | 24.9883 | 3.1775 | 63.0615 | 1.9797 | 3.1393 | 1.1978 | -0.03496 | 0.039246 | 3.0312 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.01264 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.92 | 842.9326 | 0.843 | 0.196 | 1475420838.6 | 69.6087 | 57.9458 | 23.4833 | 23.5901 | 599.971 | 84.1028 | 0.19966 | 24.7859 | 24.9883 | 999.9 | 10024.4 | 0 | 5.1025 | 999.911 | 0.9 | 0.1 | 0 | 2.3683 | 0 | 132.429 | 39.562 | 42.812 | 41.375 | 41 | 39.625 | 2.3274 | 1.0535 | -1.539 | 132.6216 | 15 | 2 | 553.244 | 706.159 | 21.988 | 27.3315 | 30.0001 | 27.3116 | 27.2822 | 5.6001 | 19.9743 | 0.38755 | 21.9994 | 40 | 23.545 | 100.64 | 101.368 | -0.016 | 5.393 | 100 | 100
0.12026 | 66.2713 | Empty Chamber Correction for 500 to 0 | 141 | 1475420841.1 | 280 | 20161002 09:07:21 | 1475420840.6 | -0.00011273 | -11.6947 | 69.2372 | 5.8369 | 5.5893 | -4.52 | 2.0084 | -0.007689 | 3.9836 | 165.2128 | 26.2105 | 24.9904 | 3.1779 | 63.0511 | 1.9797 | 3.1398 | 1.1982 | -0.034969 | 0.039256 | 3.0318 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012683 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.25 | 843.2109 | 0.843 | 0.196 | 1475420840.6 | 66.2996 | 54.5958 | 23.4825 | 23.5926 | 599.918 | 84.1032 | 0.20028 | 24.7883 | 24.9904 | 999.9 | 10026.9 | 0 | 5.1025 | 999.919 | 0.9 | 0.1 | 0 | 2.3042 | 0 | 132.529 | 39.5 | 42.812 | 41.375 | 41 | 39.625 | 2.3356 | 0.075839 | -1.4332 | 132.6104 | 15 | 2 | 546.77 | 705.986 | 21.9901 | 27.3315 | 30 | 27.3116 | 27.2822 | 5.4768 | 19.9743 | 0.38755 | 21.9994 | 35 | 23.545 | 100.64 | 101.367 | -0.016 | 5.393 | 100 | 100
0.097165 | 62.9316 | Empty Chamber Correction for 500 to 0 | 142 | 1475420843.1 | 282 | 20161002 09:07:23 | 1475420842.6 | -0.00010825 | -11.7386 | 65.9051 | 5.5559 | 5.3072 | -4.5161 | 2.0067 | -0.0073746 | 3.9805 | 165.2133 | 26.2095 | 24.9972 | 3.1791 | 63.0502 | 1.9796 | 3.1397 | 1.1996 | -0.034927 | 0.039208 | 3.0289 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012736 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.22 | 843.1846 | 0.843 | 0.196 | 1475420842.6 | 62.9543 | 51.2097 | 23.4819 | 23.5876 | 600.041 | 84.1017 | 0.20021 | 24.7878 | 24.9972 | 999.9 | 10015 | 0 | 5.1025 | 1000.22 | 0.90003 | 0.09997 | 0 | 2.2244 | 0 | 132.675 | 39.5 | 42.75 | 41.312 | 41.187 | 39.625 | 2.3533 | -0.09799 | -1.3681 | 132.5673 | 15 | 2 | 539.689 | 705.47 | 21.9941 | 27.3315 | 30 | 27.3116 | 27.2822 | 5.234 | 19.9743 | 0.38755 | 22.0043 | 35 | 23.545 | 100.64 | 101.366 | -0.016 | 5.393 | 100 | 100
0.074795 | 59.6104 | Empty Chamber Correction for 500 to 0 | 143 | 1475420845.1 | 284 | 20161002 09:07:25 | 1475420844.6 | -0.0001064 | -11.7816 | 62.5964 | 5.277 | 5.0267 | -4.5049 | 2.0018 | -0.0072393 | 3.9707 | 165.2151 | 26.2099 | 25.0056 | 3.1807 | 63.059 | 1.9796 | 3.1393 | 1.2011 | -0.034805 | 0.039071 | 3.0203 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012791 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.905 | 842.9199 | 0.843 | 0.196 | 1475420844.6 | 59.6276 | 47.8401 | 23.4823 | 23.5862 | 600.024 | 84.1018 | 0.19984 | 24.7857 | 25.0056 | 999.9 | 9980 | 0 | 5.1025 | 999.914 | 0.9 | 0.1 | 0 | 2.5255 | 0 | 132.275 | 39.5 | 42.812 | 41.375 | 41.187 | 39.687 | 2.3717 | -0.18164 | -0.89275 | 132.519 | 15 | 2 | 543.735 | 705.642 | 21.9979 | 27.3315 | 30 | 27.3116 | 27.2822 | 4.9948 | 19.9743 | 0.38755 | 22.0043 | 30 | 23.545 | 100.642 | 101.366 | -0.016 | 5.393 | 100 | 100
0.16634 | 56.288 | Empty Chamber Correction for 500 to 0 | 144 | 1475420847.1 | 286 | 20161002 09:07:27 | 1475420846.6 | -0.00010763 | -11.7099 | 59.2736 | 4.9969 | 4.7486 | -4.5131 | 2.0054 | -0.0073294 | 3.9779 | 165.2644 | 26.2069 | 24.9999 | 3.1797 | 63.065 | 1.9796 | 3.139 | 1.2001 | -0.034894 | 0.039172 | 3.0266 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012702 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.23 | 843.1931 | 0.843 | 0.196 | 1475420846.6 | 56.3283 | 44.6124 | 23.4822 | 23.5873 | 600.002 | 84.1024 | 0.20007 | 24.7842 | 24.9999 | 999.9 | 10005.6 | 0 | 5.1025 | 999.922 | 0.9 | 0.1 | 0 | 1.894 | 0 | 132.477 | 39.5 | 42.812 | 41.375 | 41.125 | 39.625 | 2.3343 | -0.79037 | -0.60178 | 132.4919 | 15 | 2 | 542.272 | 705.667 | 22.0008 | 27.3315 | 30 | 27.3116 | 27.2822 | 4.7698 | 19.9743 | 0.38755 | 22.0043 | 25 | 23.545 | 100.643 | 101.366 | -0.016 | 5.393 | 100 | 100
0.21363 | 52.9544 | Empty Chamber Correction for 500 to 0 | 145 | 1475420849.1 | 288 | 20161002 09:07:29 | 1475420848.6 | -0.0001069 | -11.6828 | 55.9425 | 4.7161 | 4.4686 | -4.5169 | 2.0071 | -0.0072859 | 3.9812 | 165.214 | 26.2052 | 24.9949 | 3.1787 | 63.067 | 1.9797 | 3.139 | 1.199 | -0.034936 | 0.039218 | 3.0295 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012673 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.921 | 842.9334 | 0.843 | 0.196 | 1475420848.6 | 53.0066 | 41.3169 | 23.4831 | 23.5875 | 599.937 | 84.1019 | 0.20007 | 24.7842 | 24.9949 | 999.9 | 10017.5 | 0 | 5.1025 | 999.921 | 0.9 | 0.1 | 0 | 2.1715 | 0 | 132.247 | 39.5 | 42.75 | 41.312 | 41.125 | 39.625 | 2.3032 | -0.64286 | -0.46492 | 132.4566 | 15 | 2 | 540.362 | 705.691 | 22.0036 | 27.3315 | 30.0001 | 27.3116 | 27.2822 | 4.5801 | 19.9743 | 0.38755 | 21.9948 | 25 | 23.5448 | 100.642 | 101.366 | -0.016 | 5.393 | 100 | 100
0.17457 | 48.7197 | Empty Chamber Correction for 500 to 0 | 146 | 1475420851.6 | 290.5 | 20161002 09:07:31 | 1475420851.1 | -0.00010465 | -11.7483 | 51.7158 | 4.3598 | 4.1108 | -4.5148 | 2.0062 | -0.0071372 | 3.9796 | 165.2131 | 26.2053 | 24.9912 | 3.178 | 63.0667 | 1.9797 | 3.1391 | 1.1983 | -0.034914 | 0.039194 | 3.0279 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012751 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.91 | 842.9241 | 0.843 | 0.196 | 1475420851.1 | 48.7623 | 37.0082 | 23.4835 | 23.5857 | 599.964 | 84.1019 | 0.20027 | 24.7846 | 24.9912 | 999.9 | 10011.2 | 0 | 5.1025 | 999.921 | 0.9 | 0.1 | 0 | 2.4745 | 0 | 132.189 | 39.5 | 42.75 | 41.312 | 41 | 39.625 | 2.2873 | -0.15803 | -1.2408 | 132.4128 | 15 | 2 | 535.099 | 705.568 | 22.0042 | 27.3315 | 30 | 27.3116 | 27.2822 | 4.3452 | 19.9743 | 0.38755 | 21.9948 | 20 | 23.5448 | 100.641 | 101.366 | -0.016 | 5.393 | 100 | 100
0.5298 | 46.1615 | Empty Chamber Correction for 500 to 0 | 147 | 1475420853.1 | 292 | 20161002 09:07:33 | 1475420852.6 | -0.0001056 | -11.4061 | 49.1597 | 4.1442 | 3.9026 | -4.5166 | 2.007 | -0.0072044 | 3.9811 | 165.2149 | 26.2064 | 24.9882 | 3.1774 | 63.0592 | 1.9796 | 3.1393 | 1.1978 | -0.034933 | 0.039215 | 3.0293 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012341 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.26 | 843.2193 | 0.843 | 0.196 | 1475420852.6 | 46.2934 | 34.8843 | 23.4828 | 23.5859 | 600.099 | 84.1002 | 0.20043 | 24.7858 | 24.9882 | 999.9 | 10016.9 | 0 | 5.1025 | 999.926 | 0.9 | 0.1 | 0 | 2.4013 | 0 | 132.337 | 39.5 | 42.75 | 41.312 | 41.125 | 39.625 | 2.3194 | -0.13899 | -1.0884 | 132.3999 | 15 | 2 | 517.663 | 705.126 | 22.0016 | 27.3315 | 30.0001 | 27.3116 | 27.2823 | 4.1411 | 19.9743 | 0.38755 | 21.9977 | 15 | 23.5448 | 100.641 | 101.366 | -0.016 | 5.393 | 100 | 100
0.5655 | 42.9517 | Empty Chamber Correction for 500 to 0 | 148 | 1475420855.1 | 294 | 20161002 09:07:35 | 1475420854.6 | -0.00010603 | -11.3899 | 45.9449 | 3.8732 | 3.6327 | -4.5316 | 2.0136 | -0.0072355 | 3.9943 | 165.2145 | 26.2014 | 24.987 | 3.1772 | 63.0635 | 1.9796 | 3.1391 | 1.1976 | -0.035097 | 0.039399 | 3.0407 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012326 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.919 | 842.9317 | 0.843 | 0.196 | 1475420854.6 | 43.0921 | 31.7021 | 23.4833 | 23.5868 | 600.236 | 84.1008 | 0.19936 | 24.7849 | 24.987 | 999.9 | 10063.8 | 0 | 5.1025 | 999.919 | 0.9 | 0.1 | 0 | 2.5226 | 0 | 132.068 | 39.5 | 42.812 | 41.375 | 41.062 | 39.625 | 2.3066 | 0.54896 | -1.5532 | 132.3658 | 15 | 2 | 522.281 | 704.888 | 21.9982 | 27.3315 | 30.0002 | 27.3116 | 27.2828 | 3.9556 | 19.9743 | 0.38755 | 21.9977 | 15 | 23.5446 | 100.64 | 101.366 | -0.016 | 5.393 | 100 | 100
-0.037381 | 39.9028 | Empty Chamber Correction for 500 to 0 | 149 | 1475420857.1 | 296 | 20161002 09:07:37 | 1475420856.6 | -0.00010639 | -12.0163 | 42.9011 | 3.6166 | 3.363 | -4.5328 | 2.0142 | -0.0072523 | 3.9953 | 165.2712 | 26.1993 | 24.994 | 3.1785 | 63.0733 | 1.9797 | 3.1387 | 1.1989 | -0.03511 | 0.039414 | 3.0416 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.013069 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.915 | 842.9283 | 0.843 | 0.196 | 1475420856.6 | 39.8924 | 27.8712 | 23.4833 | 23.5872 | 599.967 | 84.1017 | 0.19944 | 24.7825 | 24.994 | 999.9 | 10067.5 | 0 | 5.1025 | 999.91 | 0.9 | 0.1 | 0 | 2.1332 | 0 | 132.288 | 39.562 | 42.75 | 41.312 | 41.125 | 39.625 | 2.3151 | 0.36668 | -1.6678 | 132.3109 | 15 | 2 | 540.247 | 705.665 | 21.9974 | 27.3315 | 30.0002 | 27.3116 | 27.2839 | 3.7223 | 19.9743 | 0.38755 | 21.9977 | 10 | 23.5444 | 100.642 | 101.366 | -0.016 | 5.393 | 100 | 100
-0.14152 | 36.4895 | Empty Chamber Correction for 500 to 0 | 150 | 1475420859.1 | 298 | 20161002 09:07:39 | 1475420858.6 | -0.0001066 | -12.1423 | 39.5154 | 3.3312 | 3.073 | -4.4997 | 1.9995 | -0.0072722 | 3.9661 | 165.2125 | 26.2086 | 24.9894 | 3.1777 | 63.0751 | 1.9798 | 3.1387 | 1.1979 | -0.034748 | 0.039007 | 3.0163 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.013219 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.911 | 842.925 | 0.843 | 0.196 | 1475420858.6 | 36.4528 | 24.306 | 23.4844 | 23.5885 | 599.97 | 84.1009 | 0.20017 | 24.7828 | 24.9894 | 999.9 | 9963.75 | 0 | 5.1025 | 1000.25 | 0.9 | 0.1 | 0 | 2.0732 | 0 | 132.46 | 39.5 | 42.75 | 41.312 | 41 | 39.562 | 2.2911 | 0.097778 | -0.88126 | 132.2909 | 15 | 2 | 533.895 | 705.427 | 21.9978 | 27.3315 | 30.0001 | 27.3116 | 27.2845 | 3.503 | 19.9743 | 0.38755 | 22.0045 | 5 | 23.5445 | 100.642 | 101.366 | -0.016 | 5.393 | 100 | 100
0.13227 | 32.9516 | Empty Chamber Correction for 500 to 0 | 151 | 1475420861.1 | 300 | 20161002 09:07:41 | 1475420860.6 | -0.00010796 | -11.8882 | 35.9866 | 3.0337 | 2.7806 | -4.4937 | 1.9968 | -0.0073623 | 3.9607 | 165.2141 | 26.2117 | 24.9921 | 3.1782 | 63.0723 | 1.9798 | 3.1389 | 1.1984 | -0.034682 | 0.038934 | 3.0117 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012917 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.91 | 842.9241 | 0.843 | 0.196 | 1475420860.6 | 32.9841 | 21.0956 | 23.4848 | 23.5902 | 600.165 | 84.1006 | 0.19986 | 24.7837 | 24.9921 | 999.9 | 9945 | 0 | 5.1025 | 999.91 | 0.9 | 0.1 | 0 | 2.3138 | 0 | 132.125 | 39.5 | 42.75 | 41.25 | 40.687 | 39.562 | 2.2917 | 0.6404 | -0.84725 | 132.2588 | 15 | 2 | 535.636 | 704.985 | 21.9989 | 27.3315 | 30.0001 | 27.3116 | 27.2845 | 3.3569 | 19.9743 | 0.38755 | 22.0045 | 5 | 23.5444 | 100.642 | 101.366 | -0.016 | 5.393 | 100 | 100
0.2231 | 29.5481 | Empty Chamber Correction for 500 to 0 | 152 | 1475420863.1 | 302 | 20161002 09:07:43 | 1475420862.6 | -0.00010854 | -11.8177 | 32.5811 | 2.7466 | 2.4956 | -4.5043 | 2.0015 | -0.0073928 | 3.97 | 165.2712 | 26.2082 | 24.9995 | 3.1796 | 63.0761 | 1.9798 | 3.1387 | 1.1998 | -0.034798 | 0.039064 | 3.0198 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012833 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.923 | 842.9351 | 0.843 | 0.196 | 1475420862.6 | 29.6035 | 17.7817 | 23.4845 | 23.5905 | 599.955 | 84.1008 | 0.19973 | 24.7825 | 24.9995 | 999.9 | 9978.12 | 0 | 5.1025 | 999.917 | 0.9 | 0.1 | 0 | 2.6006 | 0 | 132.064 | 39.5 | 42.75 | 41.25 | 41.062 | 39.562 | 2.317 | 0.13152 | -0.54051 | 132.2182 | 15 | 2 | 546.536 | 705.673 | 22.0012 | 27.3315 | 30.0001 | 27.3116 | 27.2845 | 3.1324 | 19.9743 | 0.38755 | 22.0086 | 1.0232e-12 | 23.5444 | 100.641 | 101.366 | -0.016 | 5.393 | 100 | 100
0.57041 | 26.1348 | Empty Chamber Correction for 500 to 0 | 153 | 1475420865.1 | 304 | 20161002 09:07:45 | 1475420864.6 | -0.00010915 | -11.4888 | 29.1719 | 2.4592 | 2.2153 | -4.5051 | 2.0019 | -0.0074472 | 3.9707 | 165.2138 | 26.2068 | 24.9887 | 3.1775 | 63.0786 | 1.9797 | 3.1385 | 1.1978 | -0.034807 | 0.039074 | 3.0205 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012443 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.923 | 842.9351 | 0.843 | 0.196 | 1475420864.6 | 26.2777 | 14.785 | 23.4839 | 23.5905 | 599.945 | 84.1012 | 0.20032 | 24.7816 | 24.9887 | 999.9 | 9980.62 | 0 | 5.1025 | 1000.23 | 0.9 | 0.1 | 0 | 2.5405 | 0 | 132.135 | 39.5 | 42.75 | 41.312 | 41.125 | 39.562 | 2.3351 | -0.35364 | -0.30407 | 132.2114 | 15 | 2 | 537.019 | 705.378 | 22.0035 | 27.3315 | 30.0001 | 27.3116 | 27.2845 | 0 | 19.9743 | 0.38755 | 22.0086 | 0 | 23.5443 | 100.642 | 101.366 | -0.016 | 5.393 | 100 | 100
1.1463 | 22.7997 | Empty Chamber Correction for 500 to 0 | 154 | 1475420867.1 | 306 | 20161002 09:07:47 | 1475420866.6 | -0.00010939 | -10.929 | 25.8385 | 2.1783 | 1.9463 | -4.5086 | 2.0034 | -0.007468 | 3.9737 | 165.2159 | 26.2064 | 24.9852 | 3.1769 | 63.0784 | 1.9798 | 3.1386 | 1.1971 | -0.034845 | 0.039116 | 3.0231 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011779 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.931 | 842.9418 | 0.843 | 0.196 | 1475420866.6 | 23.0875 | 12.1582 | 23.4841 | 23.5909 | 600.124 | 84.103 | 0.2 | 24.7821 | 24.9852 | 999.9 | 9991.25 | 0 | 5.1025 | 999.931 | 0.9 | 0.1 | 0 | 2.4046 | 0 | 132.223 | 39.5 | 42.75 | 41.312 | 41 | 39.625 | 2.3349 | -0.41997 | 0.088718 | 132.2193 | 15 | 2 | 533.942 | 704.862 | 22.0056 | 27.3315 | 30.0001 | 27.3116 | 27.2845 | 0 | 19.9743 | 0.38755 | 22.0086 | 0 | 23.5438 | 100.642 | 101.366 | -0.016 | 5.393 | 100 | 100
2.1303 | 19.5965 | Empty Chamber Correction for 500 to 0 | 155 | 1475420869.1 | 308 | 20161002 09:07:49 | 1475420868.6 | -0.00010742 | -9.9572 | 22.6346 | 1.9082 | 1.6972 | -4.514 | 2.0058 | -0.0073281 | 3.9787 | 165.2156 | 26.2036 | 24.9904 | 3.1779 | 63.0828 | 1.9799 | 3.1385 | 1.198 | -0.034904 | 0.039183 | 3.0273 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010623 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.23 | 843.1931 | 0.843 | 0.196 | 1475420868.6 | 20.1314 | 10.1719 | 23.4845 | 23.5894 | 599.991 | 84.1052 | 0.19975 | 24.7816 | 24.9904 | 999.9 | 10008.1 | 0 | 5.1025 | 999.931 | 0.9 | 0.1 | 0 | 2.2553 | 0 | 132.144 | 39.5 | 42.75 | 41.312 | 41.125 | 39.625 | 2.3112 | -0.18831 | 0.49186 | 132.2028 | 15 | 2 | 541.873 | 705.108 | 22.0077 | 27.3315 | 30 | 27.3116 | 27.2845 | 0 | 19.9743 | 0.38755 | 22.017 | 5 | 23.5429 | 100.641 | 101.366 | -0.016 | 5.393 | 100 | 100
3.8031 | 16.6391 | Empty Chamber Correction for 500 to 0 | 156 | 1475420871.1 | 310 | 20161002 09:07:51 | 1475420870.6 | -0.00010915 | -8.2897 | 19.6811 | 1.6593 | 1.4834 | -4.5103 | 2.0042 | -0.0074468 | 3.9753 | 165.2149 | 26.2054 | 24.9906 | 3.1779 | 63.0853 | 1.9799 | 3.1385 | 1.1979 | -0.034864 | 0.039137 | 3.0244 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0086479 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.926 | 842.9376 | 0.843 | 0.196 | 1475420870.6 | 17.5953 | 9.3029 | 23.485 | 23.5916 | 599.94 | 84.1068 | 0.20016 | 24.7817 | 24.9906 | 999.9 | 9996.25 | 0 | 5.1025 | 999.929 | 0.9 | 0.1 | 0 | 2.583 | 0 | 131.944 | 39.5 | 42.75 | 41.25 | 41.125 | 39.562 | 2.3159 | 0.60006 | -0.23569 | 132.2058 | 15 | 2 | 538.457 | 704.96 | 22.01 | 27.3315 | 30 | 27.3116 | 27.2845 | 0 | 19.9743 | 0.38755 | 22.017 | 10 | 23.543 | 100.642 | 101.367 | -0.016 | 5.393 | 100 | 100
5.5363 | 14.3297 | Empty Chamber Correction for 500 to 0 | 157 | 1475420873.1 | 312 | 20161002 09:07:53 | 1475420872.6 | -0.00010998 | -6.5573 | 17.375 | 1.4648 | 1.3256 | -4.5059 | 2.0022 | -0.0075062 | 3.9713 | 165.2168 | 26.2081 | 24.9885 | 3.1775 | 63.083 | 1.98 | 3.1387 | 1.1975 | -0.034815 | 0.039083 | 3.021 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0065927 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.942 | 842.9511 | 0.843 | 0.196 | 1475420872.6 | 15.7234 | 9.1644 | 23.4861 | 23.5935 | 600.004 | 84.1053 | 0.20023 | 24.7828 | 24.9885 | 999.9 | 9982.5 | 0 | 5.1025 | 999.942 | 0.9 | 0.1 | 0 | 2.1619 | 0 | 132.245 | 39.5 | 42.75 | 41.312 | 41.062 | 39.562 | 2.3147 | 0.5755 | -0.4961 | 132.2091 | 15 | 2 | 534.28 | 705.083 | 22.0136 | 27.3315 | 30.0001 | 27.3116 | 27.2845 | 0.019267 | 19.9743 | 0.38755 | 22.0244 | 10 | 23.5431 | 100.642 | 101.367 | -0.016 | 5.393 | 100 | 100
38.8937 | 319.7947 | 500 to 0, with leaf | 1855 | 1475427911.5 | 7350.4 | 20161002 11:05:11 | 1475427911 | 0.006487 | 29.8107 | 351.7193 | 29.6044 | 39.32 | 0.50337 | 2.0027 | 0.44224 | 0.28127 | 165.1964 | 25.9517 | 24.9881 | 3.1774 | 55.459 | 1.9806 | 3.5713 | 1.1968 | -0.034827 | 0.039097 | 3.0219 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.036544 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.12 | 843.1016 | 0.843 | 0.196 | 1475427911 | 467.148 | 499.971 | 23.5307 | 17.1998 | 600.331 | 83.9715 | 0.1989 | 26.9624 | 24.9881 | 999.9 | 10001.9 | 0 | 6.9328 | 1000.12 | 0.89998 | 0.10002 | 0 | 2.9247 | 0 | 7946.04 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5794 | 0.74705 | -244.7911 | 7972.2162 | 15 | 2 | 527.528 | 663.238 | 25.3618 | 27.5948 | 30.0001 | 27.5431 | 27.5157 | 31.4705 | 43.693 | 0 | 25.367 | 500 | 17.2388 | 100.741 | 101.262 | -0.031 | 5.465 | 100 | 100
38.8899 | 319.5787 | 500 to 0, with leaf | 1856 | 1475427913.5 | 7352.4 | 20161002 11:05:13 | 1475427913 | 0.0064748 | 29.8071 | 351.5549 | 29.5908 | 39.3197 | 0.50252 | 2.0005 | 0.44152 | 0.28081 | 165.2454 | 25.9559 | 24.9861 | 3.177 | 55.4557 | 1.9805 | 3.5713 | 1.1965 | -0.034774 | 0.039037 | 3.0182 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.036552 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.8 | 842.8319 | 0.843 | 0.196 | 1475427913 | 467.14 | 500.003 | 23.5295 | 17.201 | 599.43 | 83.9702 | 0.20088 | 26.9627 | 24.9861 | 999.9 | 9986.88 | 0 | 6.9328 | 1000.11 | 0.89998 | 0.10002 | 0 | 2.4517 | 0 | 7939.38 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5529 | 1.0316 | -217.9901 | 7962.9565 | 15 | 2 | 518.015 | 664.059 | 25.3688 | 27.5948 | 30.0001 | 27.5431 | 27.5157 | 31.3896 | 43.693 | 0 | 25.3774 | 495.833 | 17.2388 | 100.741 | 101.263 | -0.031 | 5.465 | 100 | 100
38.6015 | 320.6721 | 500 to 0, with leaf | 1857 | 1475427915.5 | 7354.4 | 20161002 11:05:15 | 1475427915 | 0.0064806 | 29.5167 | 352.6365 | 29.6811 | 39.3189 | 0.50272 | 2.0049 | 0.44179 | 0.28098 | 165.2454 | 25.9575 | 24.9877 | 3.1773 | 55.4503 | 1.9805 | 3.5717 | 1.1968 | -0.034883 | 0.039159 | 3.0258 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.036196 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.11 | 843.0924 | 0.843 | 0.196 | 1475427915 | 467.141 | 499.681 | 23.5301 | 17.2028 | 600.074 | 83.969 | 0.20024 | 26.9644 | 24.9877 | 999.9 | 10018.1 | 0 | 6.9328 | 999.804 | 0.89998 | 0.10002 | 0 | 2.3657 | 0 | 7932.01 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5518 | 0.22067 | -203.2578 | 7956.7004 | 15 | 2 | 500.489 | 662.631 | 25.3734 | 27.5948 | 30 | 27.5431 | 27.5157 | 31.2645 | 43.693 | 0 | 25.3774 | 490.667 | 17.2388 | 100.741 | 101.263 | -0.031 | 5.465 | 100 | 100
37.6864 | 323.7016 | 500 to 0, with leaf | 1858 | 1475427917.5 | 7356.4 | 20161002 11:05:17 | 1475427917 | 0.0064801 | 28.5947 | 355.6888 | 29.9375 | 39.3011 | 0.50278 | 2.0037 | 0.44181 | 0.28099 | 165.2454 | 25.9575 | 24.986 | 3.177 | 55.4444 | 1.9803 | 3.5718 | 1.1967 | -0.034852 | 0.039124 | 3.0236 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.035113 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.797 | 842.8294 | 0.843 | 0.196 | 1475427917 | 466.939 | 498.554 | 23.5286 | 17.2021 | 600.104 | 83.9681 | 0.19945 | 26.9648 | 24.986 | 999.9 | 10009.4 | 0 | 6.9328 | 999.797 | 0.89998 | 0.10002 | 0 | 2.8224 | 0 | 7928.13 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5716 | -0.29404 | -185.134 | 7950.7835 | 15 | 2 | 520.098 | 663.169 | 25.3781 | 27.5948 | 30.0001 | 27.5431 | 27.5157 | 31.1769 | 43.693 | 0 | 25.3864 | 490.667 | 17.2388 | 100.742 | 101.263 | -0.031 | 5.465 | 100 | 100
36.3226 | 327.8274 | 500 to 0, with leaf | 1859 | 1475427919.5 | 7358.4 | 20161002 11:05:19 | 1475427919 | 0.0064772 | 27.2186 | 359.8692 | 30.2901 | 39.248 | 0.50248 | 2.0072 | 0.44167 | 0.28089 | 165.2437 | 25.9601 | 24.9854 | 3.1769 | 55.4445 | 1.9803 | 3.5718 | 1.1966 | -0.034937 | 0.03922 | 3.0296 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.033471 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.1 | 843.0839 | 0.843 | 0.196 | 1475427919 | 466.295 | 496.537 | 23.528 | 17.2025 | 599.938 | 83.9693 | 0.20059 | 26.9648 | 24.9854 | 999.9 | 10033.8 | 0 | 6.9328 | 999.793 | 0.89998 | 0.10002 | 0 | 2.4477 | 0 | 7924.73 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5704 | -0.75477 | -161.3142 | 7943.9231 | 15 | 2 | 498.401 | 662.606 | 25.382 | 27.5948 | 30.0001 | 27.5431 | 27.5157 | 31.01 | 43.693 | 0 | 25.3864 | 485.833 | 17.2388 | 100.742 | 101.262 | -0.031 | 5.465 | 100 | 100
35.0305 | 331.3032 | 500 to 0, with leaf | 1860 | 1475427921.5 | 7360.4 | 20161002 11:05:21 | 1475427921 | 0.0064778 | 25.9113 | 363.388 | 30.5867 | 39.1525 | 0.50272 | 2.0052 | 0.4418 | 0.28098 | 165.2437 | 25.9605 | 24.9837 | 3.1766 | 55.4375 | 1.9803 | 3.5721 | 1.1963 | -0.03489 | 0.039167 | 3.0263 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.03192 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.1 | 843.0839 | 0.843 | 0.196 | 1475427921 | 465.155 | 494.069 | 23.5268 | 17.2037 | 600.217 | 83.9713 | 0.19957 | 26.9663 | 24.9837 | 999.9 | 10020 | 0 | 6.9328 | 1000.1 | 0.90002 | 0.099985 | 0 | 2.3783 | 0 | 7925.51 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5758 | -0.29739 | -140.2147 | 7939.4292 | 15 | 2 | 509.8 | 662.348 | 25.3854 | 27.5948 | 30.0001 | 27.5431 | 27.5157 | 30.8541 | 43.693 | 0 | 25.3864 | 480.667 | 17.2388 | 100.742 | 101.263 | -0.031 | 5.465 | 100 | 100
34.1962 | 332.6935 | 500 to 0, with leaf | 1861 | 1475427923.5 | 7362.4 | 20161002 11:05:23 | 1475427923 | 0.0064714 | 25.062 | 364.8279 | 30.7079 | 39.0174 | 0.50292 | 2.0006 | 0.44184 | 0.28101 | 165.1944 | 25.9625 | 24.9755 | 3.175 | 55.4219 | 1.98 | 3.5725 | 1.1951 | -0.034777 | 0.03904 | 3.0183 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.030913 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.1 | 843.0848 | 0.843 | 0.196 | 1475427923 | 463.55 | 491.622 | 23.5231 | 17.2016 | 599.782 | 83.9704 | 0.20053 | 26.9684 | 24.9755 | 999.9 | 9987.5 | 0 | 6.9328 | 1000.1 | 0.89998 | 0.10002 | 0 | 2.6514 | 0 | 7922.08 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5831 | 0.10168 | -127.5255 | 7935.3577 | 15 | 2 | 507.151 | 663.262 | 25.3894 | 27.5937 | 30.0001 | 27.5431 | 27.5157 | 30.7577 | 43.693 | 0 | 25.3974 | 480.667 | 17.2388 | 100.742 | 101.262 | -0.031 | 5.465 | 100 | 100
33.4136 | 333.5643 | 500 to 0, with leaf | 1862 | 1475427925.5 | 7364.4 | 20161002 11:05:25 | 1475427925 | 0.0064732 | 24.2628 | 365.7663 | 30.7864 | 38.8581 | 0.50278 | 1.9999 | 0.44171 | 0.28093 | 165.2437 | 25.9633 | 24.9779 | 3.1755 | 55.4116 | 1.9798 | 3.5728 | 1.1957 | -0.034759 | 0.03902 | 3.0171 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.029974 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.786 | 842.8201 | 0.843 | 0.196 | 1475427925 | 461.664 | 488.911 | 23.521 | 17.201 | 600.093 | 83.9697 | 0.19996 | 26.9698 | 24.9779 | 999.9 | 9982.5 | 0 | 6.9328 | 1000.1 | 0.89998 | 0.10002 | 0 | 2.4393 | 0 | 7916.58 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5749 | 0.57871 | -114.7197 | 7930.2627 | 15 | 2 | 502.934 | 662.63 | 25.3923 | 27.5925 | 30 | 27.5431 | 27.5157 | 30.5826 | 43.693 | 0 | 25.3974 | 475.833 | 17.2388 | 100.742 | 101.261 | -0.031 | 5.465 | 100 | 100
32.6656 | 334.0955 | 500 to 0, with leaf | 1863 | 1475427927.5 | 7366.4 | 20161002 11:05:27 | 1475427927 | 0.0064684 | 23.497 | 366.3669 | 30.8376 | 38.6795 | 0.50261 | 2.0018 | 0.44163 | 0.28088 | 165.2437 | 25.9675 | 24.9739 | 3.1747 | 55.403 | 1.9796 | 3.5732 | 1.1951 | -0.034805 | 0.039072 | 3.0203 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.029056 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.1 | 843.0839 | 0.843 | 0.196 | 1475427927 | 459.533 | 486.012 | 23.5191 | 17.2005 | 599.783 | 83.9714 | 0.19983 | 26.9714 | 24.9739 | 999.9 | 9995.62 | 0 | 6.9328 | 999.793 | 0.89998 | 0.10002 | 0 | 2.5695 | 0 | 7908.83 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5821 | 1.149 | -112.734 | 7926.4765 | 15 | 2 | 514.468 | 663.398 | 25.3965 | 27.5924 | 30.0001 | 27.5426 | 27.5154 | 30.3988 | 43.693 | 0 | 25.4145 | 470.667 | 17.2388 | 100.742 | 101.261 | -0.031 | 5.465 | 100 | 100
32.1088 | 333.9125 | 500 to 0, with leaf | 1864 | 1475427929.5 | 7368.4 | 20161002 11:05:29 | 1475427929 | 0.0064738 | 22.9219 | 366.1978 | 30.8231 | 38.4788 | 0.50356 | 1.9995 | 0.4423 | 0.28132 | 165.2437 | 25.9646 | 24.9695 | 3.1739 | 55.4025 | 1.9796 | 3.5732 | 1.1943 | -0.034748 | 0.039008 | 3.0164 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.028374 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.1 | 843.0839 | 0.843 | 0.196 | 1475427929 | 457.152 | 483.029 | 23.5192 | 17.1987 | 600.103 | 83.9703 | 0.20037 | 26.9715 | 24.9695 | 999.9 | 9979.38 | 0 | 6.9328 | 1000.1 | 0.90002 | 0.099985 | 0 | 2.3324 | 0 | 7907.52 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.592 | 0.9143 | -114.2082 | 7922.8858 | 15 | 2 | 495.066 | 662.377 | 25.401 | 27.5924 | 30.0001 | 27.5415 | 27.5142 | 30.2667 | 43.693 | 0 | 25.4145 | 470.667 | 17.2388 | 100.742 | 101.26 | -0.031 | 5.465 | 100 | 100
31.474 | 333.4336 | 500 to 0, with leaf | 1865 | 1475427931.5 | 7370.4 | 20161002 11:05:31 | 1475427931 | 0.0064734 | 22.267 | 365.8378 | 30.7915 | 38.2593 | 0.50274 | 1.9984 | 0.44164 | 0.28089 | 165.2454 | 25.9633 | 24.977 | 3.1753 | 55.3994 | 1.9794 | 3.573 | 1.1959 | -0.034721 | 0.038978 | 3.0145 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.027597 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.11 | 843.0924 | 0.843 | 0.196 | 1475427931 | 454.564 | 479.763 | 23.5175 | 17.199 | 600.253 | 83.9678 | 0.19925 | 26.9705 | 24.977 | 999.9 | 9971.88 | 0 | 6.9328 | 1000.11 | 0.90002 | 0.099985 | 0 | 2.1607 | 0 | 7899.92 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5999 | -0.61826 | -127.4431 | 7917.7788 | 15 | 2 | 517.625 | 662.812 | 25.4078 | 27.5924 | 30.0001 | 27.5408 | 27.5134 | 30.0648 | 43.693 | 0 | 25.4145 | 465.833 | 17.2388 | 100.741 | 101.26 | -0.031 | 5.465 | 100 | 100
30.9851 | 332.1542 | 500 to 0, with leaf | 1866 | 1475427933.5 | 7372.4 | 20161002 11:05:33 | 1475427933 | 0.0064684 | 21.7583 | 364.7216 | 30.6981 | 38.0305 | 0.50103 | 2.0014 | 0.44039 | 0.28008 | 165.2454 | 25.9664 | 24.9904 | 3.1779 | 55.4021 | 1.9795 | 3.5729 | 1.1984 | -0.034795 | 0.039061 | 3.0196 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.026994 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.1 | 843.0848 | 0.843 | 0.196 | 1475427933 | 451.837 | 476.531 | 23.5181 | 17.1985 | 599.683 | 83.9679 | 0.20064 | 26.9704 | 24.9904 | 999.9 | 9993.12 | 0 | 6.9328 | 1000.1 | 0.90002 | 0.099985 | 0 | 2.8427 | 0 | 7893.66 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.6172 | -0.56541 | -144.1986 | 7913.6604 | 15 | 2 | 506.541 | 663.349 | 25.4141 | 27.5924 | 30.0001 | 27.5408 | 27.5134 | 29.9284 | 43.693 | 0 | 25.4324 | 460.667 | 17.2388 | 100.741 | 101.259 | -0.031 | 5.465 | 100 | 100
30.873 | 329.7887 | 500 to 0, with leaf | 1867 | 1475427935.5 | 7374.4 | 20161002 11:05:35 | 1475427935 | 0.0064745 | 21.6276 | 362.4094 | 30.5029 | 37.7889 | 0.50106 | 2.0073 | 0.44058 | 0.28018 | 165.1947 | 25.9674 | 24.9941 | 3.1786 | 55.4016 | 1.9796 | 3.5732 | 1.199 | -0.034941 | 0.039224 | 3.0298 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.026839 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.11 | 843.0932 | 0.843 | 0.196 | 1475427935 | 448.975 | 473.501 | 23.5197 | 17.1997 | 600.208 | 83.9672 | 0.19989 | 26.9714 | 24.9941 | 999.9 | 10035 | 0 | 6.9328 | 1000.11 | 0.89998 | 0.10002 | 0 | 2.5274 | 0 | 7887.49 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.617 | -0.64469 | -161.599 | 7909.1696 | 15 | 2 | 501.762 | 662.342 | 25.4212 | 27.5924 | 30.0001 | 27.5408 | 27.5134 | 29.7805 | 43.693 | 0 | 25.4324 | 460.667 | 17.2388 | 100.741 | 101.26 | -0.031 | 5.465 | 100 | 100
30.5424 | 328.138 | 500 to 0, with leaf | 1868 | 1475427937.5 | 7376.4 | 20161002 11:05:37 | 1475427937 | 0.0064665 | 21.2771 | 360.83 | 30.3699 | 37.5453 | 0.50118 | 2.003 | 0.44055 | 0.28018 | 165.2454 | 25.971 | 24.986 | 3.177 | 55.3909 | 1.9795 | 3.5736 | 1.1976 | -0.034834 | 0.039104 | 3.0224 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.026431 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.8 | 842.8319 | 0.843 | 0.196 | 1475427937 | 446.083 | 470.255 | 23.5183 | 17.2012 | 599.744 | 83.9667 | 0.19996 | 26.9736 | 24.986 | 999.9 | 10004.4 | 0 | 6.9328 | 1000.11 | 0.89998 | 0.10002 | 0 | 2.6852 | 0 | 7878.68 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5854 | -0.58772 | -178.227 | 7902.3377 | 15 | 2 | 514.578 | 663.419 | 25.4292 | 27.5924 | 30 | 27.5408 | 27.5134 | 29.5901 | 43.693 | 0 | 25.439 | 455.833 | 17.2388 | 100.742 | 101.261 | -0.031 | 5.465 | 100 | 100
30.2982 | 326.2671 | 500 to 0, with leaf | 1869 | 1475427939.5 | 7378.4 | 20161002 11:05:39 | 1475427939 | 0.0064715 | 21.0132 | 358.9903 | 30.215 | 37.2988 | 0.50192 | 2.0005 | 0.44105 | 0.28051 | 165.2454 | 25.9691 | 24.984 | 3.1766 | 55.3887 | 1.9795 | 3.5738 | 1.1971 | -0.034773 | 0.039036 | 3.0181 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.02611 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.11 | 843.0924 | 0.843 | 0.196 | 1475427939 | 443.154 | 467.03 | 23.5189 | 17.2009 | 600.127 | 83.9665 | 0.20011 | 26.9747 | 24.984 | 999.9 | 9986.88 | 0 | 6.9328 | 999.79 | 0.89998 | 0.10002 | 0 | 2.6955 | 0 | 7869.08 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.6027 | -0.42439 | -191.5279 | 7896.8396 | 15 | 2 | 498.758 | 662.436 | 25.4351 | 27.5924 | 30 | 27.5408 | 27.5134 | 29.3941 | 43.693 | 0 | 25.439 | 450.667 | 17.2388 | 100.742 | 101.261 | -0.031 | 5.465 | 100 | 100
30.128 | 323.9261 | 500 to 0, with leaf | 1870 | 1475427941.5 | 7380.4 | 20161002 11:05:41 | 1475427941 | 0.0064715 | 20.8236 | 356.7374 | 30.0246 | 37.0507 | 0.50147 | 2.0039 | 0.4408 | 0.28034 | 165.2437 | 25.9726 | 24.9865 | 3.1771 | 55.3774 | 1.9793 | 3.5742 | 1.1978 | -0.034857 | 0.03913 | 3.024 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.025885 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.1 | 843.0839 | 0.843 | 0.196 | 1475427941 | 440.217 | 463.883 | 23.5173 | 17.1997 | 600.164 | 83.9654 | 0.19914 | 26.9766 | 24.9865 | 999.9 | 10011.2 | 0 | 6.9328 | 1000.1 | 0.90002 | 0.099985 | 0 | 2.5337 | 0 | 7865.47 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5796 | -0.18092 | -199.7292 | 7891.1723 | 15 | 2 | 522.239 | 662.835 | 25.4384 | 27.5924 | 30 | 27.5408 | 27.5134 | 29.2541 | 43.693 | 0 | 25.439 | 450.667 | 17.2388 | 100.742 | 101.261 | -0.031 | 5.465 | 100 | 100
29.7338 | 322.3279 | 500 to 0, with leaf | 1871 | 1475427943.5 | 7382.4 | 20161002 11:05:43 | 1475427943 | 0.0064664 | 20.4085 | 355.2366 | 29.8995 | 36.8012 | 0.50115 | 2.0008 | 0.44047 | 0.28013 | 165.2454 | 25.9743 | 24.9857 | 3.177 | 55.3698 | 1.9792 | 3.5745 | 1.1978 | -0.034782 | 0.039046 | 3.0187 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.025393 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.09 | 843.0755 | 0.843 | 0.196 | 1475427943 | 437.236 | 460.484 | 23.5148 | 17.1972 | 599.687 | 83.967 | 0.20091 | 26.9778 | 24.9857 | 999.9 | 9989.38 | 0 | 6.9328 | 1000.09 | 0.90002 | 0.099985 | 0 | 2.322 | 0 | 7856.35 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5506 | 0.14612 | -211.9402 | 7883.0212 | 15 | 2 | 504.517 | 662.881 | 25.4413 | 27.5924 | 30 | 27.5408 | 27.5134 | 29.0845 | 43.693 | 0 | 25.4491 | 445.833 | 17.2388 | 100.743 | 101.261 | -0.031 | 5.465 | 100 | 100
29.4764 | 320.3005 | 500 to 0, with leaf | 1872 | 1475427945.5 | 7384.4 | 20161002 11:05:45 | 1475427945 | 0.0064711 | 20.1306 | 353.2678 | 29.7348 | 36.546 | 0.50133 | 2.0029 | 0.44067 | 0.28025 | 165.2454 | 25.9729 | 24.9877 | 3.1773 | 55.3715 | 1.9792 | 3.5744 | 1.1982 | -0.034833 | 0.039104 | 3.0223 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.025063 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.1 | 843.0848 | 0.843 | 0.196 | 1475427945 | 434.189 | 457.123 | 23.5139 | 17.1967 | 600.163 | 83.9706 | 0.20012 | 26.9772 | 24.9877 | 999.9 | 10003.8 | 0 | 6.9328 | 1000.09 | 0.90002 | 0.099985 | 0 | 2.4343 | 0 | 7850.19 | 38.75 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5269 | -0.46443 | -217.2188 | 7876.6819 | 15 | 2 | 496.84 | 661.944 | 25.4445 | 27.5921 | 30 | 27.5408 | 27.5134 | 28.906 | 43.693 | 0 | 25.4491 | 440.667 | 17.2388 | 100.743 | 101.261 | -0.031 | 5.465 | 100 | 100
29.374 | 317.4228 | 500 to 0, with leaf | 1873 | 1475427947.5 | 7386.4 | 20161002 11:05:47 | 1475427947 | 0.0064644 | 20.0082 | 350.5219 | 29.5052 | 36.2862 | 0.50025 | 2.0041 | 0.43986 | 0.27973 | 165.2421 | 25.9769 | 24.9928 | 3.1783 | 55.3662 | 1.9791 | 3.5746 | 1.1992 | -0.034861 | 0.039135 | 3.0243 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.024926 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.787 | 842.8209 | 0.843 | 0.196 | 1475427947 | 431.081 | 453.883 | 23.512 | 17.1976 | 599.814 | 83.9751 | 0.19987 | 26.9783 | 24.9928 | 999.9 | 10011.2 | 0 | 6.9328 | 999.787 | 0.89998 | 0.10002 | 0 | 2.6547 | 0 | 7840.47 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.523 | -1.2184 | -218.1149 | 7870.1392 | 15 | 2 | 513.251 | 662.974 | 25.4494 | 27.5908 | 30 | 27.5408 | 27.5134 | 28.7575 | 43.693 | 0 | 25.4571 | 440.667 | 17.2388 | 100.742 | 101.261 | -0.031 | 5.465 | 100 | 100
29.1925 | 315.0596 | 500 to 0, with leaf | 1874 | 1475427949.5 | 7388.4 | 20161002 11:05:49 | 1475427949 | 0.0064677 | 19.8065 | 348.2281 | 29.3121 | 36.0272 | 0.50032 | 2.003 | 0.43988 | 0.27975 | 165.2421 | 25.9768 | 24.9958 | 3.1789 | 55.3619 | 1.9792 | 3.5749 | 1.1997 | -0.034835 | 0.039105 | 3.0224 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.024679 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.1 | 843.0839 | 0.843 | 0.196 | 1475427949 | 428.003 | 450.575 | 23.5123 | 17.1974 | 600.072 | 83.9746 | 0.20054 | 26.9799 | 24.9958 | 999.9 | 10003.8 | 0 | 6.9328 | 999.786 | 0.89998 | 0.10002 | 0 | 2.3748 | 0 | 7831.8 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5077 | -0.51502 | -221.6523 | 7861.2873 | 15 | 2 | 492.316 | 662.295 | 25.4527 | 27.5901 | 30.0001 | 27.5408 | 27.5134 | 28.5642 | 43.693 | 0 | 25.4571 | 435.833 | 17.2388 | 100.742 | 101.261 | -0.031 | 5.465 | 100 | 100
29.1056 | 312.4426 | 500 to 0, with leaf | 1875 | 1475427951.5 | 7390.4 | 20161002 11:05:51 | 1475427951 | 0.0064693 | 19.7002 | 345.6645 | 29.0951 | 35.7696 | 0.50045 | 2.0074 | 0.4401 | 0.27988 | 165.2421 | 25.9787 | 24.9934 | 3.1784 | 55.357 | 1.979 | 3.575 | 1.1994 | -0.034943 | 0.039226 | 3.03 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.024553 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.09 | 843.0755 | 0.843 | 0.196 | 1475427951 | 424.96 | 447.405 | 23.5119 | 17.1959 | 600.117 | 83.9721 | 0.19953 | 26.9804 | 24.9934 | 999.9 | 10035 | 0 | 6.9328 | 1000.08 | 0.90002 | 0.099985 | 0 | 2.7874 | 0 | 7828.37 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5272 | 0.017928 | -221.4598 | 7854.7092 | 15 | 2 | 503.984 | 662.504 | 25.456 | 27.5901 | 30.0001 | 27.5408 | 27.5132 | 28.3669 | 43.693 | 0 | 25.4571 | 430.667 | 17.2388 | 100.742 | 101.261 | -0.031 | 5.465 | 100 | 100
28.9922 | 309.942 | 500 to 0, with leaf | 1876 | 1475427953.5 | 7392.4 | 20161002 11:05:53 | 1475427953 | 0.0064682 | 19.5672 | 343.2148 | 28.8883 | 35.5135 | 0.50077 | 2.0068 | 0.44034 | 0.28003 | 165.1937 | 25.9792 | 24.9885 | 3.1775 | 55.3534 | 1.979 | 3.5752 | 1.1985 | -0.034928 | 0.03921 | 3.029 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.024395 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.09 | 843.0755 | 0.843 | 0.196 | 1475427953 | 421.927 | 444.223 | 23.5119 | 17.1959 | 600.008 | 83.9694 | 0.20026 | 26.9811 | 24.9885 | 999.9 | 10031.2 | 0 | 6.9328 | 1000.09 | 0.90002 | 0.099985 | 0 | 2.5633 | 0 | 7821.96 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.537 | 0.17715 | -223.4694 | 7848.0315 | 15 | 2 | 495.802 | 662.616 | 25.4592 | 27.5901 | 30.0001 | 27.5397 | 27.5128 | 28.2535 | 43.693 | 0 | 25.4617 | 430.667 | 17.2388 | 100.742 | 101.262 | -0.031 | 5.465 | 100 | 100
28.6628 | 308.0623 | 500 to 0, with leaf | 1877 | 1475427955.5 | 7394.4 | 20161002 11:05:55 | 1475427955 | 0.0064672 | 19.2169 | 341.4288 | 28.7368 | 35.2555 | 0.50047 | 2.0056 | 0.44007 | 0.27986 | 165.2421 | 25.9801 | 24.9911 | 3.178 | 55.3513 | 1.979 | 3.5753 | 1.199 | -0.034899 | 0.039177 | 3.0269 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.02398 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.1 | 843.0839 | 0.843 | 0.196 | 1475427955 | 418.879 | 440.804 | 23.5129 | 17.198 | 600.024 | 83.9668 | 0.19955 | 26.9818 | 24.9911 | 999.9 | 10023.1 | 0 | 6.9328 | 999.789 | 0.89998 | 0.10002 | 0 | 2.8709 | 0 | 7811.23 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5213 | 0.86012 | -217.3265 | 7839.4246 | 15 | 2 | 507.254 | 662.75 | 25.4612 | 27.5901 | 30.0001 | 27.5386 | 27.5122 | 28.0533 | 43.693 | 0 | 25.4617 | 425.833 | 17.2388 | 100.742 | 101.262 | -0.031 | 5.465 | 100 | 100
28.4383 | 305.7013 | 500 to 0, with leaf | 1878 | 1475427957.5 | 7396.4 | 20161002 11:05:57 | 1475427957 | 0.0064664 | 18.9719 | 339.1846 | 28.5481 | 34.9965 | 0.49984 | 2.0016 | 0.43948 | 0.27949 | 165.2404 | 25.9795 | 24.9989 | 3.1795 | 55.3484 | 1.979 | 3.5755 | 1.2005 | -0.034801 | 0.039067 | 3.0201 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.02369 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.08 | 843.0671 | 0.843 | 0.196 | 1475427957 | 415.8 | 437.471 | 23.5128 | 17.1954 | 599.713 | 83.9664 | 0.20035 | 26.9827 | 24.9989 | 999.9 | 9995 | 0 | 6.9328 | 1000.09 | 0.90002 | 0.099985 | 0 | 2.8768 | 0 | 7808.59 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5536 | 1.4258 | -210.7234 | 7833.1792 | 15 | 2 | 499.764 | 663.047 | 25.4629 | 27.5901 | 30 | 27.5385 | 27.5115 | 27.8544 | 43.693 | 0 | 25.4669 | 420.667 | 17.2388 | 100.742 | 101.262 | -0.031 | 5.465 | 100 | 100
28.322 | 303.2212 | 500 to 0, with leaf | 1879 | 1475427959.5 | 7398.4 | 20161002 11:05:59 | 1475427959 | 0.0064716 | 18.8356 | 336.7356 | 28.3417 | 34.7366 | 0.50049 | 2.0021 | 0.43999 | 0.27982 | 165.2404 | 25.9791 | 24.9959 | 3.1789 | 55.3411 | 1.9789 | 3.5758 | 1.2 | -0.034812 | 0.039079 | 3.0208 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.023528 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.08 | 843.0671 | 0.843 | 0.196 | 1475427959 | 412.716 | 434.214 | 23.5117 | 17.1948 | 600.238 | 83.9661 | 0.19986 | 26.984 | 24.9959 | 999.9 | 9998.12 | 0 | 6.9328 | 1000.08 | 0.90002 | 0.099985 | 0 | 2.4477 | 0 | 7803.06 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.559 | 1.4114 | -204.4503 | 7826.8788 | 15 | 2 | 495.613 | 661.918 | 25.4647 | 27.5901 | 30.0002 | 27.5385 | 27.511 | 27.7079 | 43.693 | 0 | 25.4669 | 420.667 | 17.2388 | 100.742 | 101.261 | -0.031 | 5.465 | 100 | 100
27.9897 | 301.2754 | 500 to 0, with leaf | 1880 | 1475427961.5 | 7400.4 | 20161002 11:06:01 | 1475427961 | 0.0064674 | 18.4818 | 334.8833 | 28.186 | 34.4725 | 0.49998 | 2.0093 | 0.43979 | 0.27967 | 165.2404 | 25.985 | 24.9943 | 3.1786 | 55.3353 | 1.9788 | 3.576 | 1.1998 | -0.034991 | 0.03928 | 3.0333 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.023115 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.777 | 842.8125 | 0.843 | 0.196 | 1475427961 | 409.575 | 430.709 | 23.5106 | 17.1943 | 599.906 | 83.9675 | 0.19909 | 26.9851 | 24.9943 | 999.9 | 10049.4 | 0 | 6.9328 | 1000.08 | 0.89998 | 0.10002 | 0 | 2.7815 | 0 | 7795.48 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.6322 | 0.5353 | -196.0803 | 7818.7165 | 15 | 2 | 532.722 | 663.509 | 25.4666 | 27.5901 | 30.0001 | 27.5385 | 27.511 | 27.4947 | 43.693 | 0 | 25.4669 | 415.833 | 17.2388 | 100.741 | 101.262 | -0.031 | 5.465 | 100 | 100
27.7907 | 298.9258 | 500 to 0, with leaf | 1881 | 1475427963.5 | 7402.4 | 20161002 11:06:03 | 1475427963 | 0.0064695 | 18.2621 | 332.6013 | 27.9949 | 34.2094 | 0.50023 | 2.004 | 0.43984 | 0.27972 | 165.2404 | 25.9826 | 24.9969 | 3.1791 | 55.3389 | 1.979 | 3.5762 | 1.2001 | -0.034861 | 0.039134 | 3.0242 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.022848 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.08 | 843.0671 | 0.843 | 0.196 | 1475427963 | 406.435 | 427.329 | 23.5123 | 17.1942 | 599.928 | 83.9686 | 0.20088 | 26.9858 | 24.9969 | 999.9 | 10011.9 | 0 | 6.9328 | 999.771 | 0.89998 | 0.10002 | 0 | 2.5713 | 0 | 7786.96 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.6486 | 0.25395 | -193.6564 | 7812.8542 | 15 | 2 | 499.102 | 662.595 | 25.4684 | 27.5901 | 30 | 27.5385 | 27.511 | 27.3517 | 43.693 | 0 | 25.4701 | 410.667 | 17.2388 | 100.741 | 101.262 | -0.031 | 5.465 | 100 | 100
27.6948 | 296.2934 | 500 to 0, with leaf | 1882 | 1475427965.5 | 7404.4 | 20161002 11:06:05 | 1475427965 | 0.0064708 | 18.1459 | 330.0039 | 27.7759 | 33.9432 | 0.50099 | 1.999 | 0.4403 | 0.28003 | 165.2421 | 25.9798 | 24.9913 | 3.178 | 55.3373 | 1.979 | 3.5762 | 1.1991 | -0.034737 | 0.038995 | 3.0155 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.02271 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.09 | 843.0755 | 0.843 | 0.196 | 1475427965 | 403.277 | 424.027 | 23.5119 | 17.1949 | 600.157 | 83.9684 | 0.19999 | 26.9858 | 24.9913 | 999.9 | 9976.25 | 0 | 6.9328 | 1000.09 | 0.90002 | 0.099985 | 0 | 2.7356 | 0 | 7784.43 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.6606 | 0.27178 | -187.3244 | 7807.1385 | 15 | 2 | 501.143 | 662.243 | 25.4696 | 27.5901 | 30.0001 | 27.5385 | 27.511 | 27.2 | 43.693 | 0 | 25.4701 | 410.667 | 17.2388 | 100.741 | 101.262 | -0.031 | 5.465 | 100 | 100
27.4326 | 293.942 | 500 to 0, with leaf | 1883 | 1475427967.5 | 7406.4 | 20161002 11:06:07 | 1475427967 | 0.0064673 | 17.8625 | 327.7894 | 27.5895 | 33.6773 | 0.49996 | 1.9989 | 0.4395 | 0.27951 | 165.194 | 25.9811 | 24.999 | 3.1795 | 55.3341 | 1.9789 | 3.5763 | 1.2006 | -0.034734 | 0.038992 | 3.0154 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.022381 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.771 | 842.8074 | 0.843 | 0.196 | 1475427967 | 400.118 | 420.573 | 23.5111 | 17.1944 | 599.859 | 83.9681 | 0.20032 | 26.9862 | 24.999 | 999.9 | 9975.62 | 0 | 6.9328 | 1000.09 | 0.90002 | 0.099985 | 0 | 2.5098 | 0 | 7777.08 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.6242 | 0.017644 | -187.7234 | 7799.5512 | 15 | 2 | 495.286 | 662.594 | 25.4707 | 27.5901 | 30.0002 | 27.5385 | 27.511 | 27.004 | 43.693 | 0 | 25.4738 | 405.833 | 17.2388 | 100.741 | 101.261 | -0.031 | 5.465 | 100 | 100
27.3698 | 291.0506 | 500 to 0, with leaf | 1884 | 1475427969.5 | 7408.4 | 20161002 11:06:09 | 1475427969 | 0.0064702 | 17.78 | 324.9906 | 27.3536 | 33.4152 | 0.49952 | 2.0009 | 0.43921 | 0.27932 | 165.1932 | 25.9811 | 25.0056 | 3.1807 | 55.3322 | 1.9788 | 3.5763 | 1.2019 | -0.034784 | 0.039048 | 3.0189 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.022282 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.784 | 842.8184 | 0.843 | 0.196 | 1475427969 | 397.008 | 417.352 | 23.5108 | 17.1942 | 600.138 | 83.9676 | 0.19992 | 26.9864 | 25.0056 | 999.9 | 9990 | 0 | 6.9328 | 999.782 | 0.89998 | 0.10002 | 0 | 2.7548 | 0 | 7768.31 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.6459 | -0.83042 | -190.5012 | 7793.8146 | 15 | 2 | 495.294 | 662.219 | 25.4716 | 27.5892 | 30.0001 | 27.5385 | 27.511 | 26.8026 | 43.693 | 0 | 25.4738 | 400.667 | 17.2388 | 100.741 | 101.261 | -0.031 | 5.465 | 100 | 100
27.3848 | 287.9995 | 500 to 0, with leaf | 1885 | 1475427971.5 | 7410.4 | 20161002 11:06:11 | 1475427971 | 0.0064675 | 17.7761 | 322.0074 | 27.1027 | 33.157 | 0.49946 | 2.0026 | 0.43921 | 0.27931 | 165.2421 | 25.9843 | 25.0025 | 3.1802 | 55.3259 | 1.9787 | 3.5765 | 1.2014 | -0.034826 | 0.039095 | 3.0218 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.022271 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.09 | 843.0755 | 0.843 | 0.196 | 1475427971 | 393.939 | 414.266 | 23.5094 | 17.193 | 599.907 | 83.9678 | 0.20014 | 26.9874 | 25.0025 | 999.9 | 10001.9 | 0 | 6.9328 | 1000.09 | 0.90002 | 0.099985 | 0 | 2.667 | 0 | 7765.75 | 38.75 | 41.812 | 40.437 | 40.625 | 39.187 | 2.6249 | -0.87328 | -187.788 | 7788.31 | 15 | 2 | 499.053 | 662.711 | 25.473 | 27.5878 | 30 | 27.5385 | 27.511 | 26.6577 | 43.693 | 0 | 25.4738 | 400.667 | 17.2388 | 100.741 | 101.262 | -0.031 | 5.465 | 100 | 100
27.0724 | 286.201 | 500 to 0, with leaf | 1886 | 1475427973.5 | 7412.4 | 20161002 11:06:13 | 1475427973 | 0.0064686 | 17.4429 | 320.2602 | 26.9559 | 32.9014 | 0.49974 | 2.0056 | 0.43951 | 0.2795 | 165.1927 | 25.9855 | 24.9993 | 3.1795 | 55.3231 | 1.9787 | 3.5767 | 1.2008 | -0.0349 | 0.039178 | 3.027 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.021876 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.09 | 843.0755 | 0.843 | 0.196 | 1475427973 | 390.898 | 410.869 | 23.5089 | 17.1925 | 600.013 | 83.969 | 0.19984 | 26.9881 | 24.9993 | 999.9 | 10023.1 | 0 | 6.9328 | 1000.09 | 0.90002 | 0.099985 | 0 | 2.6923 | 0 | 7758.77 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5974 | -0.20852 | -186.575 | 7780.9038 | 15 | 2 | 498.374 | 662.22 | 25.474 | 27.5877 | 30 | 27.5385 | 27.511 | 26.4872 | 43.693 | 0 | 25.3408 | 395.833 | 17.2388 | 100.741 | 101.261 | -0.031 | 5.465 | 100 | 100
26.7728 | 284.3137 | 500 to 0, with leaf | 1887 | 1475427975.5 | 7414.4 | 20161002 11:06:15 | 1475427975 | 0.0064704 | 17.1221 | 318.4197 | 26.8025 | 32.6424 | 0.50039 | 1.9989 | 0.43984 | 0.27973 | 165.1912 | 25.9812 | 24.9969 | 3.1791 | 55.3263 | 1.9788 | 3.5765 | 1.2003 | -0.034734 | 0.038992 | 3.0154 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.021502 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.772 | 842.8083 | 0.843 | 0.196 | 1475427975 | 387.798 | 407.431 | 23.5081 | 17.1893 | 599.948 | 83.9737 | 0.19994 | 26.9875 | 24.9969 | 999.9 | 9975 | 0 | 6.9328 | 1000.09 | 0.90002 | 0.099985 | 0 | 2.5737 | 0 | 7753.59 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5985 | -0.15104 | -184.015 | 7775.3 | 15 | 2 | 503.651 | 662.337 | 25.4504 | 27.5877 | 30.0002 | 27.5385 | 27.511 | 26.3016 | 43.693 | 0 | 25.3408 | 390.667 | 17.2388 | 100.74 | 101.26 | -0.031 | 5.465 | 100 | 100
26.7153 | 281.3299 | 500 to 0, with leaf | 1888 | 1475427977.5 | 7416.4 | 20161002 11:06:17 | 1475427977 | 0.0064698 | 17.045 | 315.5568 | 26.5625 | 32.3823 | 0.4996 | 1.997 | 0.43917 | 0.2793 | 165.2421 | 25.982 | 25.0058 | 3.1808 | 55.3216 | 1.9787 | 3.5767 | 1.2021 | -0.034688 | 0.03894 | 3.0121 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.021404 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.09 | 843.0755 | 0.843 | 0.196 | 1475427977 | 384.695 | 404.228 | 23.5067 | 17.1893 | 600.028 | 83.9764 | 0.20006 | 26.9885 | 25.0058 | 999.9 | 9961.25 | 0 | 6.9328 | 1000.09 | 0.90002 | 0.099985 | 0 | 2.6681 | 0 | 7748.41 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.582 | -0.28239 | -180.6687 | 7769.9708 | 15 | 2 | 498.364 | 662.056 | 25.3984 | 27.5877 | 30.0002 | 27.5385 | 27.511 | 26.1496 | 43.693 | 0 | 25.3404 | 390.667 | 17.2388 | 100.74 | 101.26 | -0.031 | 5.465 | 100 | 100
26.5136 | 278.8704 | 500 to 0, with leaf | 1889 | 1475427979.5 | 7418.4 | 20161002 11:06:19 | 1475427979 | 0.0064685 | 16.8226 | 313.2125 | 26.3635 | 32.1175 | 0.49893 | 1.9964 | 0.43863 | 0.27895 | 165.2438 | 25.982 | 25.0102 | 3.1816 | 55.3136 | 1.9784 | 3.5767 | 1.2032 | -0.034673 | 0.038923 | 3.0111 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.02114 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.09 | 843.0755 | 0.843 | 0.196 | 1475427979 | 381.574 | 400.862 | 23.5045 | 17.189 | 600.087 | 83.972 | 0.19921 | 26.9883 | 25.0102 | 999.9 | 9957.5 | 0 | 6.9328 | 1000.08 | 0.90002 | 0.099985 | 0 | 2.3832 | 0 | 7740.76 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5964 | 0.18175 | -179.0072 | 7762.6723 | 15 | 2 | 521.681 | 663.04 | 25.3559 | 27.5877 | 30.0001 | 27.5385 | 27.511 | 25.9303 | 43.693 | 0 | 25.3404 | 385.833 | 17.2389 | 100.739 | 101.26 | -0.031 | 5.465 | 100 | 100
26.3792 | 276.462 | 500 to 0, with leaf | 1890 | 1475427981.5 | 7420.4 | 20161002 11:06:21 | 1475427981 | 0.0064672 | 16.6683 | 310.8321 | 26.1626 | 31.8573 | 0.49954 | 1.9994 | 0.43919 | 0.27931 | 165.2438 | 25.9813 | 25.0002 | 3.1797 | 55.3181 | 1.9783 | 3.5762 | 1.2014 | -0.034746 | 0.039005 | 3.0162 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.020957 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.09 | 843.0764 | 0.843 | 0.196 | 1475427981 | 378.49 | 397.61 | 23.5034 | 17.1869 | 599.875 | 83.9689 | 0.20067 | 26.9858 | 25.0002 | 999.9 | 9978.75 | 0 | 6.9328 | 1000.08 | 0.90002 | 0.099985 | 0 | 2.9307 | 0 | 7734.88 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5836 | 1.0249 | -177.1949 | 7757.0665 | 15 | 2 | 503.691 | 662.806 | 25.3399 | 27.5877 | 30.0001 | 27.5385 | 27.511 | 25.7336 | 43.693 | 0 | 25.3404 | 380.667 | 17.2389 | 100.739 | 101.261 | -0.031 | 5.465 | 100 | 100
26.2051 | 274.0352 | 500 to 0, with leaf | 1891 | 1475427983.5 | 7422.4 | 20161002 11:06:23 | 1475427983 | 0.0064681 | 16.4741 | 308.4963 | 25.9656 | 31.5999 | 0.49896 | 2.0083 | 0.43897 | 0.27914 | 165.2404 | 25.9823 | 25.0037 | 3.1804 | 55.3244 | 1.9782 | 3.5756 | 1.2022 | -0.034965 | 0.039252 | 3.0315 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.020733 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.779 | 842.8142 | 0.843 | 0.196 | 1475427983 | 375.437 | 394.335 | 23.5027 | 17.1881 | 600.143 | 83.9686 | 0.19973 | 26.9831 | 25.0037 | 999.9 | 10041.9 | 0 | 6.9328 | 999.779 | 0.89998 | 0.10002 | 0 | 2.6543 | 0 | 7727.1 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.6165 | 0.75849 | -179.7617 | 7751.6904 | 15 | 2 | 507.042 | 662.056 | 25.33 | 27.5877 | 29.9999 | 27.5385 | 27.511 | 25.6205 | 43.693 | 0 | 25.3353 | 380.667 | 17.2391 | 100.739 | 101.26 | -0.031 | 5.465 | 100 | 100
25.8471 | 272.0983 | 500 to 0, with leaf | 1892 | 1475427985.5 | 7424.4 | 20161002 11:06:25 | 1475427985 | 0.0064651 | 16.0947 | 306.694 | 25.8145 | 31.3388 | 0.49813 | 2.004 | 0.43822 | 0.27867 | 165.2404 | 25.9797 | 25.0119 | 3.1819 | 55.3309 | 1.9782 | 3.5753 | 1.2037 | -0.034859 | 0.039132 | 3.0241 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.020277 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.09 | 843.0755 | 0.843 | 0.196 | 1475427985 | 372.327 | 390.836 | 23.5026 | 17.187 | 599.766 | 83.9698 | 0.20041 | 26.9814 | 25.0119 | 999.9 | 10011.2 | 0 | 6.9328 | 1000.08 | 0.90002 | 0.099985 | 0 | 2.6573 | 0 | 7722.92 | 38.812 | 41.812 | 40.437 | 40.625 | 39.125 | 2.6183 | -0.034458 | -180.9826 | 7744.62 | 15 | 2 | 503.517 | 662.781 | 25.3248 | 27.5877 | 29.9999 | 27.5374 | 27.511 | 25.421 | 43.693 | 0 | 25.3353 | 375.833 | 17.2392 | 100.739 | 101.258 | -0.031 | 5.465 | 100 | 100
25.6883 | 269.7604 | 500 to 0, with leaf | 1893 | 1475427987.5 | 7426.4 | 20161002 11:06:27 | 1475427987 | 0.0064687 | 15.9152 | 304.3617 | 25.6189 | 31.0734 | 0.49933 | 2.0004 | 0.43905 | 0.27922 | 165.1927 | 25.9742 | 25.0043 | 3.1805 | 55.3421 | 1.9784 | 3.5748 | 1.2021 | -0.034771 | 0.039033 | 3.0179 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.02007 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.781 | 842.8159 | 0.843 | 0.196 | 1475427987 | 369.163 | 387.464 | 23.5037 | 17.1878 | 600.071 | 83.9725 | 0.19998 | 26.9792 | 25.0043 | 999.9 | 9985.62 | 0 | 6.9328 | 1000.08 | 0.90002 | 0.099985 | 0 | 2.8745 | 0 | 7717.21 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.6255 | -0.042947 | -180.2 | 7739.1227 | 15 | 2 | 499.616 | 662.15 | 25.3216 | 27.5877 | 30 | 27.5362 | 27.511 | 25.2177 | 43.693 | 0 | 25.3284 | 370.667 | 17.2418 | 100.74 | 101.259 | -0.031 | 5.465 | 100 | 100
25.6184 | 267.018 | 500 to 0, with leaf | 1894 | 1475427989.5 | 7428.4 | 20161002 11:06:29 | 1475427989 | 0.0064652 | 15.8254 | 301.6699 | 25.3924 | 30.8106 | 0.49955 | 2.0039 | 0.43932 | 0.27938 | 165.2421 | 25.9752 | 24.9969 | 3.1791 | 55.3483 | 1.9783 | 3.5743 | 1.2008 | -0.034856 | 0.039129 | 3.0239 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.019957 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.09 | 843.0755 | 0.843 | 0.196 | 1475427989 | 366.041 | 384.232 | 23.5031 | 17.1902 | 600.03 | 83.9729 | 0.19976 | 26.9769 | 24.9969 | 999.9 | 10010 | 0 | 6.9328 | 1000.09 | 0.90002 | 0.099985 | 0 | 2.3669 | 0 | 7711.67 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.6035 | -0.24122 | -180.2205 | 7733.8058 | 15 | 2 | 511.208 | 662.641 | 25.3204 | 27.5877 | 30.0001 | 27.5361 | 27.511 | 25.0724 | 43.693 | 0 | 25.3284 | 370.667 | 17.2394 | 100.74 | 101.259 | -0.031 | 5.465 | 100 | 100
25.3186 | 265.0678 | 500 to 0, with leaf | 1895 | 1475427991.5 | 7430.4 | 20161002 11:06:31 | 1475427991 | 0.006466 | 15.5047 | 299.794 | 25.2343 | 30.5512 | 0.49952 | 2.0048 | 0.43932 | 0.27938 | 165.2421 | 25.9739 | 24.9972 | 3.1791 | 55.3504 | 1.9782 | 3.574 | 1.2009 | -0.034879 | 0.039155 | 3.0255 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.019577 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.09 | 843.0755 | 0.843 | 0.196 | 1475427991 | 362.961 | 380.815 | 23.5022 | 17.1873 | 599.918 | 83.9716 | 0.20058 | 26.9755 | 24.9972 | 999.9 | 10016.9 | 0 | 6.9328 | 1000.09 | 0.90002 | 0.099985 | 0 | 2.2677 | 0 | 7704.31 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.6038 | -0.75183 | -181.3022 | 7726.4888 | 15 | 2 | 493.473 | 662.382 | 25.3192 | 27.5877 | 30.0001 | 27.5361 | 27.511 | 24.8566 | 43.693 | 0 | 25.3284 | 365.833 | 17.243 | 100.74 | 101.258 | -0.031 | 5.465 | 100 | 100
25.0593 | 262.8822 | 500 to 0, with leaf | 1896 | 1475427993.5 | 7432.4 | 20161002 11:06:33 | 1475427993 | 0.0064689 | 15.2245 | 297.7096 | 25.059 | 30.2909 | 0.49906 | 2.0053 | 0.43897 | 0.27915 | 165.2438 | 25.9714 | 25.0042 | 3.1805 | 55.352 | 1.9781 | 3.5737 | 1.2024 | -0.034891 | 0.039168 | 3.0263 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.019244 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.09 | 843.0764 | 0.843 | 0.196 | 1475427993 | 359.867 | 377.416 | 23.5004 | 17.1848 | 600.117 | 83.9726 | 0.19998 | 26.9738 | 25.0042 | 999.9 | 10020 | 0 | 6.9328 | 1000.09 | 0.90002 | 0.099985 | 0 | 2.5438 | 0 | 7698.8 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5727 | -1.4963 | -182.5521 | 7720.9796 | 15 | 2 | 491.109 | 662.076 | 25.3183 | 27.5877 | 30.0001 | 27.5361 | 27.5109 | 24.7087 | 43.693 | 0 | 25.3297 | 360.667 | 17.2447 | 100.741 | 101.258 | -0.031 | 5.465 | 100 | 100
25.0306 | 259.8234 | 500 to 0, with leaf | 1897 | 1475427995.5 | 7434.4 | 20161002 11:06:35 | 1475427995 | 0.0064653 | 15.1763 | 294.7652 | 24.8112 | 30.0292 | 0.49836 | 2.0044 | 0.4384 | 0.27879 | 165.2404 | 25.9707 | 25.0078 | 3.1812 | 55.351 | 1.9779 | 3.5733 | 1.2033 | -0.034869 | 0.039144 | 3.0248 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.019193 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.785 | 842.8192 | 0.843 | 0.196 | 1475427995 | 356.757 | 374.24 | 23.4979 | 17.1845 | 599.995 | 83.9726 | 0.20006 | 26.9723 | 25.0078 | 999.9 | 10013.8 | 0 | 6.9328 | 999.785 | 0.89998 | 0.10002 | 0 | 2.4841 | 0 | 7689.54 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.544 | -1.0856 | -183.293 | 7715.3292 | 15 | 2 | 489.354 | 662.09 | 25.3193 | 27.5877 | 30.0001 | 27.5361 | 27.5101 | 24.5548 | 43.693 | 0 | 25.3297 | 360.667 | 17.244 | 100.741 | 101.258 | -0.031 | 5.465 | 100 | 100
24.7804 | 257.7128 | 500 to 0, with leaf | 1898 | 1475427997.5 | 7436.4 | 20161002 11:06:37 | 1475427997 | 0.0064655 | 14.9055 | 292.7259 | 24.639 | 29.7699 | 0.49841 | 2.0037 | 0.43843 | 0.2788 | 165.2421 | 25.9706 | 25.0076 | 3.1811 | 55.3501 | 1.9779 | 3.5734 | 1.2032 | -0.034851 | 0.039123 | 3.0235 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.018866 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.09 | 843.0755 | 0.843 | 0.196 | 1475427997 | 353.684 | 370.878 | 23.4984 | 17.1842 | 599.939 | 83.971 | 0.19997 | 26.9726 | 25.0076 | 999.9 | 10008.8 | 0 | 6.9328 | 999.788 | 0.89998 | 0.10002 | 0 | 2.6053 | 0 | 7681.34 | 38.75 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5215 | -0.30073 | -190.6209 | 7707.74 | 15 | 2 | 493.049 | 662.22 | 25.3214 | 27.5867 | 30.0001 | 27.5361 | 27.5093 | 24.3554 | 43.693 | 0 | 25.3255 | 355.833 | 17.2438 | 100.741 | 101.258 | -0.031 | 5.465 | 100 | 100
24.6404 | 255.2986 | 500 to 0, with leaf | 1899 | 1475427999.5 | 7438.4 | 20161002 11:06:39 | 1475427999 | 0.0064682 | 14.7455 | 290.3467 | 24.4384 | 29.5106 | 0.499 | 2.003 | 0.43887 | 0.27909 | 165.2437 | 25.9695 | 25.0037 | 3.1804 | 55.3491 | 1.9779 | 3.5734 | 1.2025 | -0.034835 | 0.039106 | 3.0224 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.018676 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.1 | 843.0839 | 0.843 | 0.196 | 1475427999 | 350.609 | 367.622 | 23.4985 | 17.1824 | 600.013 | 83.9698 | 0.19974 | 26.9727 | 25.0037 | 999.9 | 10004.4 | 0 | 6.9328 | 1000.09 | 0.90002 | 0.099985 | 0 | 2.6064 | 0 | 7678.35 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5324 | 0.1074 | -197.0564 | 7702.1208 | 15 | 2 | 496.317 | 662.213 | 25.323 | 27.5855 | 30.0001 | 27.5361 | 27.5088 | 24.1485 | 43.693 | 0 | 25.3255 | 350.667 | 17.2474 | 100.741 | 101.257 | -0.031 | 5.465 | 100 | 100
24.6579 | 252.2966 | 500 to 0, with leaf | 1900 | 1475428001.5 | 7440.4 | 20161002 11:06:41 | 1475428001 | 0.0064665 | 14.7442 | 287.3992 | 24.1896 | 29.2535 | 0.49912 | 2.0055 | 0.43903 | 0.27919 | 165.193 | 25.9708 | 24.9981 | 3.1793 | 55.3414 | 1.9776 | 3.5734 | 1.2017 | -0.034897 | 0.039175 | 3.0268 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.018675 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.08 | 843.0671 | 0.843 | 0.196 | 1475428001 | 347.564 | 364.555 | 23.4959 | 17.1816 | 600.026 | 83.967 | 0.20018 | 26.9727 | 24.9981 | 999.9 | 10022.5 | 0 | 6.9328 | 1000.08 | 0.90002 | 0.099985 | 0 | 2.675 | 0 | 7672.24 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5275 | 0.039938 | -199.3884 | 7696.0688 | 15 | 2 | 496.186 | 662.26 | 25.3232 | 27.5854 | 30.0002 | 27.5361 | 27.5088 | 23.9984 | 43.693 | 0 | 25.3255 | 350.667 | 17.2456 | 100.74 | 101.257 | -0.031 | 5.465 | 100 | 100
24.2935 | 250.6412 | 500 to 0, with leaf | 1901 | 1475428003.5 | 7442.4 | 20161002 11:06:43 | 1475428003 | 0.0064638 | 14.359 | 285.8295 | 24.0571 | 29.0018 | 0.49893 | 2.0057 | 0.43888 | 0.27909 | 165.2404 | 25.9718 | 24.9966 | 3.179 | 55.3383 | 1.9774 | 3.5733 | 1.2016 | -0.034901 | 0.039179 | 3.027 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.018218 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.1 | 843.0839 | 0.843 | 0.196 | 1475428003 | 344.579 | 361.166 | 23.4943 | 17.1821 | 599.976 | 83.9657 | 0.20007 | 26.9722 | 24.9966 | 999.9 | 10023.8 | 0 | 6.9328 | 999.782 | 0.89998 | 0.10002 | 0 | 2.6512 | 0 | 7660.77 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5422 | 1.0018 | -201.7036 | 7688.1873 | 15 | 2 | 493.446 | 662.12 | 25.3227 | 27.5854 | 30.0001 | 27.5361 | 27.5088 | 23.8224 | 43.693 | 0 | 25.3235 | 345.833 | 17.2504 | 100.74 | 101.258 | -0.031 | 5.465 | 100 | 100
23.9192 | 249.0323 | 500 to 0, with leaf | 1902 | 1475428005.5 | 7444.4 | 20161002 11:06:45 | 1475428005 | 0.006466 | 13.9635 | 284.2772 | 23.9262 | 28.7447 | 0.49929 | 2.0034 | 0.4391 | 0.27924 | 165.2421 | 25.9695 | 24.9958 | 3.1789 | 55.3411 | 1.9775 | 3.5732 | 1.2014 | -0.034844 | 0.039116 | 3.0231 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.017749 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.09 | 843.0755 | 0.843 | 0.196 | 1475428005 | 341.528 | 357.699 | 23.4951 | 17.1813 | 600.03 | 83.9653 | 0.19977 | 26.9718 | 24.9958 | 999.9 | 10007.5 | 0 | 6.9328 | 1000.09 | 0.90002 | 0.099985 | 0 | 2.7332 | 0 | 7657.91 | 38.75 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5611 | 1.3463 | -201.7925 | 7682.3327 | 15 | 2 | 497.646 | 661.909 | 25.3225 | 27.5854 | 30 | 27.5361 | 27.5088 | 23.6352 | 43.693 | 0 | 25.3235 | 340.667 | 17.249 | 100.741 | 101.259 | -0.031 | 5.465 | 100 | 100
23.8776 | 246.1095 | 500 to 0, with leaf | 1903 | 1475428007.5 | 7446.4 | 20161002 11:06:47 | 1475428007 | 0.0064639 | 13.9023 | 281.4394 | 23.6876 | 28.484 | 0.49901 | 2.0045 | 0.43891 | 0.27911 | 165.2421 | 25.97 | 24.9964 | 3.179 | 55.3428 | 1.9774 | 3.5731 | 1.2016 | -0.034873 | 0.039147 | 3.0251 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.017676 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.09 | 843.0755 | 0.843 | 0.196 | 1475428007 | 338.426 | 354.516 | 23.4945 | 17.1824 | 599.995 | 83.9658 | 0.20017 | 26.971 | 24.9964 | 999.9 | 10015.6 | 0 | 6.9328 | 1000.09 | 0.90002 | 0.099985 | 0 | 2.3098 | 0 | 7651.39 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5585 | 0.81322 | -201.735 | 7676.1465 | 15 | 2 | 494.399 | 662.073 | 25.3221 | 27.5854 | 30 | 27.5361 | 27.5088 | 23.4734 | 43.693 | 0 | 25.3238 | 340.667 | 17.2533 | 100.741 | 101.259 | -0.031 | 5.465 | 100 | 100
23.678 | 243.7662 | 500 to 0, with leaf | 1904 | 1475428009.5 | 7448.4 | 20161002 11:06:49 | 1475428009 | 0.0064636 | 13.6826 | 279.1955 | 23.4991 | 28.2278 | 0.49863 | 2.0023 | 0.43856 | 0.27889 | 165.2404 | 25.9685 | 25.0004 | 3.1798 | 55.3404 | 1.9773 | 3.5729 | 1.2025 | -0.034819 | 0.039087 | 3.0213 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.017421 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.778 | 842.8133 | 0.843 | 0.196 | 1475428009 | 335.378 | 351.226 | 23.4923 | 17.1815 | 600.089 | 83.9676 | 0.19959 | 26.9704 | 25.0004 | 999.9 | 10000 | 0 | 6.9328 | 1000.09 | 0.89998 | 0.10002 | 0 | 2.8354 | 0 | 7642.51 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.6155 | 0.48966 | -202.5822 | 7667.8804 | 15 | 2 | 503.094 | 662.096 | 25.322 | 27.5854 | 30.0001 | 27.5361 | 27.5088 | 23.2736 | 43.4123 | 0 | 25.3238 | 335.833 | 17.2576 | 100.741 | 101.259 | -0.031 | 5.465 | 100 | 100
23.4935 | 241.1479 | 500 to 0, with leaf | 1905 | 1475428011.5 | 7450.4 | 20161002 11:06:51 | 1475428011 | 0.006444 | 13.4781 | 276.7759 | 23.2965 | 27.9734 | 0.49664 | 2.0021 | 0.43701 | 0.27789 | 165.2404 | 25.9756 | 25.0042 | 3.1805 | 55.3418 | 1.9773 | 3.5729 | 1.2031 | -0.034812 | 0.039079 | 3.0208 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.017173 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.08 | 843.0671 | 0.843 | 0.196 | 1475428011 | 332.34 | 347.966 | 23.4918 | 17.1966 | 599.757 | 83.9711 | 0.20006 | 26.9704 | 25.0042 | 999.9 | 9997.5 | 0 | 6.9328 | 999.778 | 0.89998 | 0.10002 | 0 | 2.796 | 0 | 7633.12 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.6268 | 0.18218 | -207.2875 | 7661.7485 | 15 | 2 | 508.185 | 662.798 | 25.3223 | 27.5854 | 30.0001 | 27.5361 | 27.5088 | 23.0675 | 43.4123 | 0 | 25.3238 | 330.667 | 17.2513 | 100.741 | 101.259 | -0.031 | 5.465 | 100 | 100
23.3368 | 238.5507 | 500 to 0, with leaf | 1906 | 1475428013.5 | 7452.4 | 20161002 11:06:53 | 1475428013 | 0.00643 | 13.3015 | 274.3287 | 23.0905 | 27.7175 | 0.49527 | 2.0071 | 0.43608 | 0.27727 | 165.2404 | 25.9821 | 25.0067 | 3.1809 | 55.3603 | 1.9779 | 3.5727 | 1.2031 | -0.034935 | 0.039218 | 3.0295 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.016963 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.09 | 843.0755 | 0.843 | 0.196 | 1475428013 | 329.3 | 344.715 | 23.4983 | 17.221 | 600.15 | 83.9707 | 0.20036 | 26.9694 | 25.0067 | 999.9 | 10033.1 | 0 | 6.9328 | 1000.09 | 0.90002 | 0.099985 | 0 | 2.307 | 0 | 7630.06 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.6075 | -0.37021 | -207.1631 | 7655.5862 | 15 | 2 | 487.592 | 661.816 | 25.3226 | 27.5854 | 30.0001 | 27.5361 | 27.5088 | 22.9495 | 43.4123 | 0 | 25.3232 | 330.667 | 17.2503 | 100.741 | 101.258 | -0.031 | 5.465 | 100 | 100
23.0925 | 236.4662 | 500 to 0, with leaf | 1907 | 1475428015.5 | 7454.4 | 20161002 11:06:55 | 1475428015 | 0.0064298 | 13.0367 | 272.3108 | 22.9195 | 27.4602 | 0.49549 | 2.0026 | 0.43613 | 0.27732 | 165.1921 | 25.9785 | 25.0091 | 3.1814 | 55.3845 | 1.9786 | 3.5725 | 1.2028 | -0.034826 | 0.039095 | 3.0218 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.016655 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.779 | 842.8142 | 0.843 | 0.196 | 1475428015 | 326.259 | 341.391 | 23.5081 | 17.2305 | 600.098 | 83.9675 | 0.19922 | 26.9682 | 25.0091 | 999.9 | 10001.9 | 0 | 6.9328 | 1000.09 | 0.90002 | 0.099985 | 0 | 2.2468 | 0 | 7621.09 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5829 | -0.96085 | -214.4188 | 7647.1204 | 15 | 2 | 511.163 | 662.331 | 25.3226 | 27.5854 | 30.0001 | 27.5361 | 27.5088 | 22.7393 | 43.4123 | 0 | 25.3232 | 325.833 | 17.2503 | 100.74 | 101.26 | -0.031 | 5.465 | 100 | 100
22.8132 | 234.6345 | 500 to 0, with leaf | 1908 | 1475428017.5 | 7456.4 | 20161002 11:06:57 | 1475428017 | 0.0064354 | 12.7368 | 270.4877 | 22.7661 | 27.2034 | 0.49653 | 2.0014 | 0.43691 | 0.27783 | 165.2404 | 25.9753 | 25.0075 | 3.1811 | 55.4095 | 1.9794 | 3.5723 | 1.2017 | -0.034797 | 0.039062 | 3.0198 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.016294 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.08 | 843.0671 | 0.843 | 0.196 | 1475428017 | 323.208 | 338.03 | 23.5173 | 17.231 | 599.789 | 83.9666 | 0.20031 | 26.9672 | 25.0075 | 999.9 | 9993.75 | 0 | 6.9328 | 1000.08 | 0.90002 | 0.099985 | 0 | 2.8855 | 0 | 7614.04 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5772 | -0.94675 | -215.1087 | 7640.7996 | 15 | 2 | 506.483 | 662.682 | 25.3225 | 27.5854 | 30.0001 | 27.5361 | 27.5088 | 22.5268 | 43.4123 | 0 | 25.3154 | 320.667 | 17.2503 | 100.74 | 101.26 | -0.031 | 5.465 | 100 | 100
22.6921 | 232.3126 | 500 to 0, with leaf | 1909 | 1475428019.5 | 7458.4 | 20161002 11:06:59 | 1475428019 | 0.0064453 | 12.5961 | 268.1451 | 22.5688 | 26.948 | 0.49781 | 2.0059 | 0.43802 | 0.27853 | 165.2404 | 25.9729 | 25.0032 | 3.1803 | 55.423 | 1.9798 | 3.5721 | 1.2005 | -0.034905 | 0.039184 | 3.0274 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.016127 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.08 | 843.0671 | 0.843 | 0.196 | 1475428019 | 320.175 | 334.831 | 23.522 | 17.2299 | 600.151 | 83.9662 | 0.20033 | 26.9664 | 25.0032 | 999.9 | 10025 | 0 | 6.9328 | 1000.08 | 0.90002 | 0.099985 | 0 | 2.636 | 0 | 7607.39 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5485 | -0.74476 | -218.5956 | 7634.3085 | 15 | 2 | 492.717 | 661.722 | 25.3218 | 27.5854 | 30.0001 | 27.5361 | 27.5088 | 22.3698 | 43.4123 | 0 | 25.3154 | 320.667 | 17.2503 | 100.74 | 101.26 | -0.031 | 5.465 | 100 | 100
22.3007 | 230.7454 | 500 to 0, with leaf | 1910 | 1475428021.5 | 7460.4 | 20161002 11:07:01 | 1475428021 | 0.0064434 | 12.1835 | 266.6506 | 22.4432 | 26.6928 | 0.49803 | 1.9998 | 0.43803 | 0.27855 | 165.2388 | 25.9709 | 25.0013 | 3.1799 | 55.4227 | 1.9798 | 3.5721 | 1.2002 | -0.034756 | 0.039016 | 3.0169 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.015642 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.759 | 842.7973 | 0.843 | 0.196 | 1475428021 | 317.141 | 331.37 | 23.5219 | 17.2291 | 599.913 | 83.9669 | 0.20006 | 26.9665 | 25.0013 | 999.9 | 9981.88 | 0 | 6.9328 | 1000.07 | 0.89998 | 0.10002 | 0 | 2.4639 | 0 | 7598.43 | 38.812 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5436 | -0.10887 | -221.4803 | 7625.5323 | 15 | 2 | 496.047 | 662.26 | 25.3192 | 27.5854 | 30 | 27.5361 | 27.5088 | 22.1484 | 43.4123 | 0 | 25.3154 | 315.833 | 17.2503 | 100.74 | 101.26 | -0.031 | 5.465 | 100 | 100
21.9686 | 228.8586 | 500 to 0, with leaf | 1911 | 1475428023.5 | 7462.4 | 20161002 11:07:03 | 1475428023 | 0.0064433 | 11.8302 | 264.8595 | 22.2931 | 26.4337 | 0.49786 | 1.9947 | 0.43776 | 0.27839 | 165.2371 | 25.9671 | 25.0055 | 3.1807 | 55.4299 | 1.9798 | 3.5718 | 1.2009 | -0.03463 | 0.038876 | 3.0081 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.015219 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.06 | 843.0502 | 0.843 | 0.196 | 1475428023 | 314.053 | 327.908 | 23.522 | 17.2297 | 599.944 | 83.9696 | 0.20006 | 26.9649 | 25.0055 | 999.9 | 9945.62 | 0 | 6.9328 | 1000.07 | 0.90002 | 0.099985 | 0 | 2.3595 | 0 | 7591.61 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5319 | -0.61405 | -219.6711 | 7618.8938 | 15 | 2 | 489.66 | 661.886 | 25.316 | 27.5854 | 30.0001 | 27.5361 | 27.5088 | 21.9954 | 43.4123 | 0 | 25.3116 | 310.667 | 17.2503 | 100.74 | 101.26 | -0.031 | 5.465 | 100 | 100
21.8634 | 226.1733 | 500 to 0, with leaf | 1912 | 1475428025.5 | 7464.4 | 20161002 11:07:05 | 1475428025 | 0.0064456 | 11.7048 | 262.2425 | 22.0738 | 26.1709 | 0.49774 | 2.0006 | 0.43782 | 0.27842 | 165.1909 | 25.9681 | 25.0083 | 3.1813 | 55.4373 | 1.9801 | 3.5717 | 1.2012 | -0.034776 | 0.039039 | 3.0183 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.01507 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.07 | 843.0586 | 0.843 | 0.196 | 1475428025 | 310.917 | 324.625 | 23.5236 | 17.23 | 600.035 | 83.9732 | 0.19994 | 26.9645 | 25.0083 | 999.9 | 9986.88 | 0 | 6.9328 | 1000.07 | 0.90002 | 0.099985 | 0 | 2.7772 | 0 | 7583.97 | 38.75 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5358 | 0.63564 | -226.7607 | 7612.0635 | 15 | 2 | 493.112 | 661.909 | 25.3141 | 27.5854 | 30.0001 | 27.5351 | 27.5088 | 21.837 | 43.4123 | 0 | 25.3116 | 310.667 | 17.2503 | 100.74 | 101.259 | -0.031 | 5.465 | 100 | 100
21.6014 | 224.0927 | 500 to 0, with leaf | 1913 | 1475428027.5 | 7466.4 | 20161002 11:07:07 | 1475428027 | 0.0064461 | 11.4218 | 260.2198 | 21.9041 | 25.9089 | 0.49799 | 2.0007 | 0.43802 | 0.27855 | 165.2404 | 25.9706 | 25.0063 | 3.1809 | 55.4314 | 1.9801 | 3.5722 | 1.2008 | -0.034779 | 0.039042 | 3.0185 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.014734 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.07 | 843.0586 | 0.843 | 0.196 | 1475428027 | 307.797 | 321.202 | 23.5235 | 17.2295 | 600.041 | 83.9755 | 0.19995 | 26.9667 | 25.0063 | 999.9 | 9987.5 | 0 | 6.9328 | 1000.08 | 0.90002 | 0.099985 | 0 | 2.4013 | 0 | 7574.28 | 38.75 | 41.812 | 40.437 | 40.5 | 39.125 | 2.5114 | 0.22937 | -229.1176 | 7603.0596 | 15 | 2 | 490.039 | 661.769 | 25.3122 | 27.5854 | 30 | 27.5339 | 27.5088 | 21.6304 | 43.4123 | 0 | 25.3053 | 305.833 | 17.2503 | 100.74 | 101.259 | -0.031 | 5.465 | 100 | 100
21.5103 | 221.4089 | 500 to 0, with leaf | 1914 | 1475428029.5 | 7468.4 | 20161002 11:07:09 | 1475428029 | 0.0064485 | 11.311 | 257.5976 | 21.6836 | 25.6487 | 0.49813 | 2.0003 | 0.43812 | 0.27861 | 165.2421 | 25.9695 | 25.0075 | 3.1811 | 55.4323 | 1.9801 | 3.5722 | 1.201 | -0.034768 | 0.039031 | 3.0178 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.014603 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.08 | 843.0679 | 0.843 | 0.196 | 1475428029 | 304.702 | 317.978 | 23.5236 | 17.2267 | 599.993 | 83.9767 | 0.19973 | 26.9667 | 25.0075 | 999.9 | 9984.38 | 0 | 6.9328 | 1000.08 | 0.90002 | 0.099985 | 0 | 2.7222 | 0 | 7566.4 | 38.687 | 41.75 | 40.375 | 40.375 | 39.062 | 2.5617 | 1.1711 | -232.4465 | 7595.9 | 15 | 2 | 499.6 | 662.073 | 25.3105 | 27.5845 | 30.0001 | 27.5338 | 27.5088 | 21.4214 | 43.4123 | 0 | 25.3053 | 300.667 | 17.2503 | 100.74 | 101.259 | -0.031 | 5.465 | 100 | 100
21.4646 | 218.5871 | 500 to 0, with leaf | 1915 | 1475428031.5 | 7470.4 | 20161002 11:07:11 | 1475428031 | 0.006453 | 11.246 | 254.8497 | 21.4527 | 25.3936 | 0.49781 | 2.0083 | 0.43808 | 0.27857 | 165.1907 | 25.9686 | 25.0136 | 3.1823 | 55.4451 | 1.9803 | 3.5717 | 1.2019 | -0.034964 | 0.039251 | 3.0315 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.014525 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.09 | 843.0764 | 0.843 | 0.196 | 1475428031 | 301.667 | 314.858 | 23.5254 | 17.225 | 600.078 | 83.9776 | 0.20005 | 26.9643 | 25.0136 | 999.9 | 10040.6 | 0 | 6.9328 | 1000.09 | 0.89998 | 0.10002 | 0 | 2.476 | 0 | 7559.14 | 38.75 | 41.812 | 40.437 | 40.5 | 39.125 | 2.5781 | 0.58226 | -240.1385 | 7588.9204 | 15 | 2 | 494.593 | 661.863 | 25.3076 | 27.5831 | 30.0001 | 27.5338 | 27.5088 | 21.2629 | 43.4123 | 0 | 25.3053 | 300.667 | 17.2503 | 100.74 | 101.26 | -0.031 | 5.465 | 100 | 100
21.1247 | 216.771 | 500 to 0, with leaf | 1916 | 1475428033.5 | 7472.4 | 20161002 11:07:13 | 1475428033 | 0.0064536 | 10.8854 | 253.1443 | 21.3077 | 25.1389 | 0.49733 | 2.005 | 0.43762 | 0.27828 | 165.2388 | 25.9661 | 25.0193 | 3.1833 | 55.4436 | 1.9801 | 3.5714 | 1.2032 | -0.034884 | 0.039161 | 3.0259 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.014098 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.07 | 843.0586 | 0.843 | 0.196 | 1475428033 | 298.66 | 311.47 | 23.5246 | 17.2242 | 600.135 | 83.9729 | 0.19929 | 26.9631 | 25.0193 | 999.9 | 10018.1 | 0 | 6.9328 | 1000.07 | 0.90002 | 0.099985 | 0 | 2.3639 | 0 | 7549.65 | 38.75 | 41.75 | 40.5 | 40.437 | 39.125 | 2.584 | 0.11241 | -241.6349 | 7579.4892 | 15 | 2 | 513.345 | 662.379 | 25.3053 | 27.583 | 30 | 27.5338 | 27.5088 | 21.085 | 43.4123 | 0 | 25.2946 | 295.833 | 17.2503 | 100.739 | 101.259 | -0.031 | 5.465 | 100 | 100
20.9077 | 214.5341 | 500 to 0, with leaf | 1917 | 1475428035.5 | 7474.4 | 20161002 11:07:15 | 1475428035 | 0.0064472 | 10.6481 | 250.995 | 21.126 | 24.8812 | 0.49718 | 2.0019 | 0.43743 | 0.27816 | 165.1897 | 25.966 | 25.0142 | 3.1824 | 55.4381 | 1.9799 | 3.5713 | 1.2025 | -0.034808 | 0.039076 | 3.0206 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.013821 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.763 | 842.8007 | 0.843 | 0.196 | 1475428035 | 295.61 | 308.169 | 23.5223 | 17.2242 | 599.761 | 83.9685 | 0.20058 | 26.9625 | 25.0142 | 999.9 | 9996.88 | 0 | 6.9328 | 1000.07 | 0.90002 | 0.099985 | 0 | 2.5198 | 0 | 7541.83 | 38.75 | 41.812 | 40.437 | 40.5 | 39.062 | 2.5738 | -0.20265 | -243.9101 | 7572.1081 | 15 | 2 | 502.718 | 662.844 | 25.3024 | 27.583 | 30 | 27.5338 | 27.5084 | 20.8885 | 43.4123 | 0 | 25.2946 | 290.667 | 17.2503 | 100.739 | 101.258 | -0.031 | 5.465 | 100 | 100
20.7682 | 212.1737 | 500 to 0, with leaf | 1918 | 1475428037.5 | 7476.4 | 20161002 11:07:17 | 1475428037 | 0.0064524 | 10.4887 | 248.6417 | 20.9277 | 24.6233 | 0.4981 | 2.0039 | 0.43819 | 0.27865 | 165.2371 | 25.9647 | 25.0084 | 3.1813 | 55.4414 | 1.9799 | 3.5711 | 1.2014 | -0.034856 | 0.039129 | 3.0239 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.013628 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.06 | 843.0502 | 0.843 | 0.196 | 1475428037 | 292.549 | 304.923 | 23.523 | 17.2236 | 600.115 | 83.9681 | 0.20007 | 26.9618 | 25.0084 | 999.9 | 10010.6 | 0 | 6.9328 | 1000.06 | 0.90002 | 0.099985 | 0 | 2.1447 | 0 | 7534.32 | 38.812 | 41.812 | 40.437 | 40.562 | 39.125 | 2.5632 | -0.29317 | -245.4862 | 7564.8465 | 15 | 2 | 496.988 | 661.856 | 25.2975 | 27.583 | 30 | 27.5338 | 27.5082 | 20.7262 | 43.4123 | 0 | 25.2798 | 290.667 | 17.2503 | 100.739 | 101.258 | -0.031 | 5.465 | 100 | 100
20.3952 | 210.4769 | 500 to 0, with leaf | 1919 | 1475428039.5 | 7478.4 | 20161002 11:07:19 | 1475428039 | 0.0064493 | 10.0948 | 247.0577 | 20.7941 | 24.3697 | 0.49758 | 2.0015 | 0.43773 | 0.27835 | 165.2405 | 25.9643 | 25.0119 | 3.1819 | 55.4428 | 1.9799 | 3.571 | 1.2021 | -0.034797 | 0.039063 | 3.0198 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.01316 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.06 | 843.0502 | 0.843 | 0.196 | 1475428039 | 289.54 | 301.506 | 23.5232 | 17.2236 | 599.807 | 83.9673 | 0.19977 | 26.9613 | 25.0119 | 999.9 | 9993.75 | 0 | 6.9328 | 1000.06 | 0.90002 | 0.099985 | 0 | 2.5553 | 0 | 7524.24 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5803 | -0.31218 | -247.8376 | 7554.8912 | 15 | 2 | 509.782 | 662.597 | 25.2925 | 27.583 | 30 | 27.5338 | 27.5076 | 20.518 | 43.4123 | 0 | 25.2798 | 285.833 | 17.2503 | 100.739 | 101.258 | -0.031 | 5.465 | 100 | 100
20.2504 | 208.0358 | 500 to 0, with leaf | 1920 | 1475428041.5 | 7480.4 | 20161002 11:07:21 | 1475428041 | 0.0064525 | 9.9303 | 244.6915 | 20.5949 | 24.1163 | 0.49748 | 2.0021 | 0.43766 | 0.27831 | 165.2388 | 25.9631 | 25.0161 | 3.1827 | 55.4446 | 1.9799 | 3.571 | 1.2028 | -0.034812 | 0.03908 | 3.0208 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.012965 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.07 | 843.0586 | 0.843 | 0.196 | 1475428041 | 286.53 | 298.308 | 23.5236 | 17.2235 | 600.06 | 83.9667 | 0.20022 | 26.961 | 25.0161 | 999.9 | 9998.12 | 0 | 6.9328 | 1000.07 | 0.90002 | 0.099985 | 0 | 2.8594 | 0 | 7516.08 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.6027 | 0.12012 | -249.506 | 7547.4688 | 15 | 2 | 497.539 | 661.817 | 25.2871 | 27.583 | 30 | 27.5338 | 27.5068 | 20.3057 | 43.4123 | 0 | 25.2798 | 280.667 | 17.2503 | 100.74 | 101.258 | -0.031 | 5.465 | 100 | 100
20.1528 | 205.4764 | 500 to 0, with leaf | 1921 | 1475428043.5 | 7482.4 | 20161002 11:07:23 | 1475428043 | 0.0064533 | 9.8134 | 242.179 | 20.383 | 23.8615 | 0.4978 | 2.0027 | 0.43793 | 0.27848 | 165.2388 | 25.9613 | 25.0127 | 3.1821 | 55.4491 | 1.9799 | 3.5706 | 1.2022 | -0.034827 | 0.039096 | 3.0219 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.012827 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.06 | 843.0502 | 0.843 | 0.196 | 1475428043 | 283.509 | 295.147 | 23.5236 | 17.2248 | 600.257 | 83.9661 | 0.19882 | 26.9592 | 25.0127 | 999.9 | 10002.5 | 0 | 6.9328 | 1000.06 | 0.90002 | 0.099985 | 0 | 2.3823 | 0 | 7508.37 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5978 | -0.86971 | -250.6414 | 7539.7608 | 15 | 2 | 529.288 | 662.28 | 25.2799 | 27.583 | 30 | 27.5338 | 27.5064 | 20.1793 | 43.4123 | 0 | 25.2664 | 280.667 | 17.2503 | 100.74 | 101.258 | -0.031 | 5.465 | 100 | 100
19.8158 | 203.7038 | 500 to 0, with leaf | 1922 | 1475428045.5 | 7484.4 | 20161002 11:07:25 | 1475428045 | 0.006443 | 9.4557 | 240.4973 | 20.2417 | 23.6083 | 0.49771 | 1.9969 | 0.43771 | 0.27835 | 165.2371 | 25.9606 | 25.0057 | 3.1808 | 55.4541 | 1.9798 | 3.5702 | 1.2009 | -0.034685 | 0.038937 | 3.0119 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.012406 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.753 | 842.7923 | 0.843 | 0.196 | 1475428045 | 280.497 | 291.769 | 23.5229 | 17.2264 | 599.518 | 83.9648 | 0.20127 | 26.9574 | 25.0057 | 999.9 | 9961.88 | 0 | 6.9328 | 1000.07 | 0.89998 | 0.10002 | 0 | 2.464 | 0 | 7497.63 | 38.812 | 41.812 | 40.437 | 40.625 | 39.125 | 2.567 | 0.24623 | -256.6612 | 7529.5977 | 15 | 2 | 506.009 | 662.817 | 25.2744 | 27.583 | 30 | 27.5338 | 27.5064 | 19.9655 | 43.4123 | 0 | 25.2664 | 275.833 | 17.2503 | 100.74 | 101.258 | -0.031 | 5.465 | 100 | 100
19.5729 | 201.6673 | 500 to 0, with leaf | 1923 | 1475428047.5 | 7486.4 | 20161002 11:07:27 | 1475428047 | 0.0064537 | 9.1924 | 238.499 | 20.073 | 23.3529 | 0.49817 | 1.9988 | 0.43811 | 0.27861 | 165.2388 | 25.956 | 25.0106 | 3.1817 | 55.4613 | 1.9799 | 3.5699 | 1.2018 | -0.034733 | 0.03899 | 3.0153 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.01209 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.07 | 843.0586 | 0.843 | 0.196 | 1475428047 | 277.469 | 288.449 | 23.5243 | 17.2242 | 600.171 | 83.9639 | 0.20015 | 26.9558 | 25.0106 | 999.9 | 9975.62 | 0 | 6.9328 | 1000.07 | 0.90002 | 0.099985 | 0 | 2.4818 | 0 | 7490.22 | 38.812 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5411 | -0.077268 | -258.4776 | 7522.0319 | 15 | 2 | 492.005 | 661.53 | 25.268 | 27.583 | 30 | 27.5338 | 27.5064 | 19.7498 | 43.4123 | 0 | 25.2576 | 270.667 | 17.2503 | 100.741 | 101.258 | -0.031 | 5.465 | 100 | 100
19.3828 | 199.2482 | 500 to 0, with leaf | 1924 | 1475428049.5 | 7488.4 | 20161002 11:07:29 | 1475428049 | 0.0064519 | 8.982 | 236.1994 | 19.8796 | 23.0947 | 0.49723 | 2.0062 | 0.43758 | 0.27825 | 165.2371 | 25.957 | 25.0158 | 3.1827 | 55.4667 | 1.9798 | 3.5693 | 1.2029 | -0.034914 | 0.039194 | 3.0279 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.011844 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.767 | 842.8041 | 0.843 | 0.196 | 1475428049 | 274.4 | 285.152 | 23.5225 | 17.2226 | 600.024 | 83.965 | 0.19945 | 26.9529 | 25.0158 | 999.9 | 10027.5 | 0 | 6.9328 | 999.767 | 0.89998 | 0.10002 | 0 | 2.6227 | 0 | 7479.47 | 38.812 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5422 | -0.3538 | -258.6147 | 7514.1954 | 15 | 2 | 513.886 | 662.232 | 25.2631 | 27.583 | 30 | 27.5338 | 27.5064 | 19.5852 | 43.4123 | 0 | 25.2576 | 270.667 | 17.2503 | 100.741 | 101.257 | -0.031 | 5.465 | 100 | 100
18.9698 | 197.69 | 500 to 0, with leaf | 1925 | 1475428051.5 | 7490.4 | 20161002 11:07:31 | 1475428051 | 0.0064524 | 8.5475 | 234.731 | 19.7571 | 22.8377 | 0.49708 | 2.0044 | 0.43742 | 0.27815 | 165.2371 | 25.9547 | 25.0196 | 3.1834 | 55.4745 | 1.9799 | 3.569 | 1.2035 | -0.034869 | 0.039144 | 3.0248 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.011325 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.07 | 843.0586 | 0.843 | 0.196 | 1475428051 | 271.331 | 281.631 | 23.5227 | 17.221 | 599.898 | 83.9684 | 0.20063 | 26.9516 | 25.0196 | 999.9 | 10014.4 | 0 | 6.9328 | 1000.07 | 0.90002 | 0.099985 | 0 | 2.7303 | 0 | 7471.91 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5518 | -0.16106 | -258.5942 | 7503.9646 | 15 | 2 | 496.07 | 661.693 | 25.2594 | 27.583 | 30 | 27.5338 | 27.5064 | 19.3631 | 43.4123 | 0 | 25.2576 | 265.833 | 17.2503 | 100.741 | 101.257 | -0.031 | 5.465 | 100 | 100
18.7301 | 195.5413 | 500 to 0, with leaf | 1926 | 1475428053.5 | 7492.4 | 20161002 11:07:33 | 1475428053 | 0.0064549 | 8.2871 | 232.6495 | 19.5816 | 22.578 | 0.4972 | 2.0034 | 0.43748 | 0.27819 | 165.2388 | 25.9537 | 25.0208 | 3.1836 | 55.4724 | 1.9798 | 3.5691 | 1.2038 | -0.034845 | 0.039117 | 3.0231 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.011016 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.06 | 843.0511 | 0.843 | 0.196 | 1475428053 | 268.25 | 278.265 | 23.5226 | 17.2218 | 600.219 | 83.9676 | 0.20009 | 26.9519 | 25.0208 | 999.9 | 10007.5 | 0 | 6.9328 | 1000.05 | 0.90002 | 0.099985 | 0 | 2.6384 | 0 | 7463.93 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5459 | -0.0039077 | -261.972 | 7496.1654 | 15 | 2 | 492.997 | 661.553 | 25.2559 | 27.583 | 29.9999 | 27.5338 | 27.5064 | 19.2033 | 43.4123 | 0 | 25.24 | 260.667 | 17.2503 | 100.74 | 101.257 | -0.031 | 5.465 | 100 | 100
18.7308 | 192.4744 | 500 to 0, with leaf | 1927 | 1475428055.5 | 7494.4 | 20161002 11:07:35 | 1475428055 | 0.00645 | 8.2687 | 229.6733 | 19.3302 | 22.3164 | 0.4969 | 2.0022 | 0.43722 | 0.27802 | 165.2371 | 25.9535 | 25.0182 | 3.1831 | 55.4698 | 1.9796 | 3.5688 | 1.2035 | -0.034815 | 0.039083 | 3.0211 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.010994 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.05 | 843.0418 | 0.843 | 0.196 | 1475428055 | 265.154 | 275.134 | 23.5206 | 17.2216 | 599.934 | 83.964 | 0.20007 | 26.9505 | 25.0182 | 999.9 | 9999.38 | 0 | 6.9328 | 1000.05 | 0.90002 | 0.099985 | 0 | 2.8766 | 0 | 7455.96 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5763 | 0.41615 | -259.8198 | 7488.3273 | 15 | 2 | 496.077 | 662.254 | 25.2497 | 27.583 | 30 | 27.5338 | 27.5064 | 19.0311 | 43.4123 | 0 | 25.24 | 260.667 | 17.2503 | 100.74 | 101.258 | -0.031 | 5.465 | 100 | 100
18.4326 | 190.5425 | 500 to 0, with leaf | 1928 | 1475428057.5 | 7496.4 | 20161002 11:07:37 | 1475428057 | 0.0064501 | 7.9495 | 227.8188 | 19.175 | 22.0602 | 0.4969 | 2.0004 | 0.43717 | 0.278 | 165.2371 | 25.9511 | 25.0196 | 3.1834 | 55.4769 | 1.9797 | 3.5684 | 1.2037 | -0.034771 | 0.039034 | 3.018 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.010619 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.751 | 842.7906 | 0.843 | 0.196 | 1475428057 | 262.098 | 271.739 | 23.5203 | 17.2213 | 599.945 | 83.968 | 0.19987 | 26.9489 | 25.0196 | 999.9 | 9986.25 | 0 | 6.9328 | 1000.06 | 0.89998 | 0.10002 | 0 | 2.7455 | 0 | 7444.57 | 38.75 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5648 | 0.63997 | -263.2756 | 7477.8035 | 15 | 2 | 498.464 | 661.786 | 25.2422 | 27.583 | 29.9999 | 27.5338 | 27.5064 | 18.8211 | 43.4123 | 0 | 25.2206 | 255.833 | 17.2503 | 100.741 | 101.258 | -0.031 | 5.465 | 100 | 100
18.1747 | 188.501 | 500 to 0, with leaf | 1929 | 1475428059.5 | 7498.4 | 20161002 11:07:39 | 1475428059 | 0.006451 | 7.6711 | 225.837 | 19.0092 | 21.8044 | 0.49701 | 2.0036 | 0.43734 | 0.2781 | 165.2371 | 25.9506 | 25.0188 | 3.1832 | 55.4843 | 1.9797 | 3.5681 | 1.2035 | -0.034849 | 0.039121 | 3.0234 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.010285 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.06 | 843.0502 | 0.843 | 0.196 | 1475428059 | 259.045 | 268.388 | 23.52 | 17.2202 | 599.951 | 83.9721 | 0.20018 | 26.9473 | 25.0188 | 999.9 | 10008.1 | 0 | 6.9328 | 1000.06 | 0.90002 | 0.099985 | 0 | 2.6245 | 0 | 7436.73 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5771 | 0.33787 | -263.6964 | 7469.86 | 15 | 2 | 494.169 | 661.645 | 25.2356 | 27.5824 | 30 | 27.5338 | 27.5064 | 18.6038 | 43.4123 | 0 | 25.2206 | 250.667 | 17.2503 | 100.741 | 101.259 | -0.031 | 5.465 | 100 | 100
18.0652 | 186.0536 | 500 to 0, with leaf | 1930 | 1475428061.5 | 7500.4 | 20161002 11:07:41 | 1475428061 | 0.0064552 | 7.5421 | 223.3853 | 18.8025 | 21.5481 | 0.49816 | 2.0021 | 0.43819 | 0.27865 | 165.2371 | 25.947 | 25.0105 | 3.1817 | 55.489 | 1.9797 | 3.5678 | 1.2019 | -0.034814 | 0.039081 | 3.0209 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.010136 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.744 | 842.7847 | 0.843 | 0.196 | 1475428061 | 256.004 | 265.197 | 23.5205 | 17.2183 | 600.108 | 83.9708 | 0.20001 | 26.9459 | 25.0105 | 999.9 | 9998.12 | 0 | 6.9328 | 999.744 | 0.89998 | 0.10002 | 0 | 2.6017 | 0 | 7425.84 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5796 | 0.95104 | -269.1248 | 7461.9123 | 15 | 2 | 492.266 | 661.716 | 25.2276 | 27.5814 | 30.0001 | 27.5333 | 27.5064 | 18.4442 | 43.4123 | 0 | 25.2206 | 250.667 | 17.2503 | 100.741 | 101.259 | -0.031 | 5.465 | 100 | 100
17.7161 | 184.3315 | 500 to 0, with leaf | 1931 | 1475428063.5 | 7502.4 | 20161002 11:07:43 | 1475428063 | 0.006454 | 7.1722 | 221.7513 | 18.6643 | 21.2944 | 0.49798 | 2.0011 | 0.43802 | 0.27855 | 165.2355 | 25.9458 | 25.0111 | 3.1818 | 55.4898 | 1.9796 | 3.5676 | 1.2021 | -0.034789 | 0.039053 | 3.0192 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0096937 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.05 | 843.0418 | 0.843 | 0.196 | 1475428063 | 253 | 261.806 | 23.5202 | 17.2173 | 599.938 | 83.9677 | 0.19995 | 26.9448 | 25.0111 | 999.9 | 9991.25 | 0 | 6.9328 | 1000.06 | 0.90002 | 0.099985 | 0 | 2.7249 | 0 | 7417.45 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.6193 | 0.60238 | -269.3648 | 7451.1654 | 15 | 2 | 495.744 | 662.347 | 25.2195 | 27.5808 | 30.0002 | 27.5322 | 27.5064 | 18.2606 | 43.4123 | 0 | 25.2064 | 245.833 | 17.2503 | 100.742 | 101.259 | -0.031 | 5.465 | 100 | 100
17.4122 | 182.3815 | 500 to 0, with leaf | 1932 | 1475428065.5 | 7504.4 | 20161002 11:07:45 | 1475428065 | 0.0064542 | 6.8475 | 219.9068 | 18.5086 | 21.038 | 0.49738 | 2.0051 | 0.43767 | 0.27831 | 165.1896 | 25.9454 | 25.0156 | 3.1826 | 55.4926 | 1.9795 | 3.5672 | 1.2031 | -0.034888 | 0.039165 | 3.0261 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0093112 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.762 | 842.7998 | 0.843 | 0.196 | 1475428065 | 249.959 | 258.419 | 23.5194 | 17.2176 | 600.055 | 83.9657 | 0.20014 | 26.943 | 25.0156 | 999.9 | 10020 | 0 | 6.9328 | 999.759 | 0.89998 | 0.10002 | 0 | 2.3031 | 0 | 7406.35 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.6206 | -0.062479 | -275.9282 | 7442.9254 | 15 | 2 | 487.88 | 661.856 | 25.2147 | 27.5807 | 30.0001 | 27.5315 | 27.5064 | 18.0602 | 43.4123 | 0 | 25.2064 | 240.667 | 17.2503 | 100.742 | 101.26 | -0.031 | 5.465 | 100 | 100
17.2886 | 179.9072 | 500 to 0, with leaf | 1933 | 1475428067.5 | 7506.4 | 20161002 11:07:47 | 1475428067 | 0.0064554 | 6.7043 | 217.4707 | 18.3034 | 20.7821 | 0.49794 | 2.0021 | 0.43802 | 0.27854 | 165.2371 | 25.9419 | 25.0119 | 3.1819 | 55.5006 | 1.9796 | 3.5668 | 1.2024 | -0.034814 | 0.039081 | 3.0209 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0091386 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.06 | 843.0502 | 0.843 | 0.196 | 1475428067 | 246.921 | 255.219 | 23.5201 | 17.2167 | 600.019 | 83.9651 | 0.19993 | 26.9409 | 25.0119 | 999.9 | 9998.75 | 0 | 6.9328 | 1000.06 | 0.90002 | 0.099985 | 0 | 2.7318 | 0 | 7400.77 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.6136 | -0.22889 | -274.3515 | 7434.7508 | 15 | 2 | 494.742 | 661.786 | 25.2075 | 27.5807 | 30 | 27.5315 | 27.5064 | 17.8906 | 43.4123 | 0 | 25.1926 | 240.667 | 17.2503 | 100.742 | 101.259 | -0.031 | 5.465 | 100 | 100
17.0397 | 177.8592 | 500 to 0, with leaf | 1934 | 1475428069.5 | 7508.4 | 20161002 11:07:49 | 1475428069 | 0.0064568 | 6.4351 | 215.4803 | 18.1359 | 20.5275 | 0.4982 | 2.0006 | 0.43818 | 0.27865 | 165.2355 | 25.9395 | 25.0115 | 3.1819 | 55.5079 | 1.9797 | 3.5665 | 1.2022 | -0.034777 | 0.03904 | 3.0183 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.008822 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.746 | 842.7864 | 0.843 | 0.196 | 1475428069 | 243.896 | 251.906 | 23.5215 | 17.2165 | 599.989 | 83.965 | 0.20015 | 26.9397 | 25.0115 | 999.9 | 9988.12 | 0 | 6.9328 | 1000.06 | 0.89998 | 0.10002 | 0 | 2.5908 | 0 | 7389.52 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5799 | -0.85692 | -277.3802 | 7423.8035 | 15 | 2 | 491.058 | 661.693 | 25.2013 | 27.5807 | 30.0001 | 27.5315 | 27.5064 | 17.6767 | 43.4123 | 0 | 25.1926 | 235.833 | 17.2503 | 100.742 | 101.26 | -0.031 | 5.465 | 100 | 100
16.8457 | 175.6106 | 500 to 0, with leaf | 1935 | 1475428071.5 | 7510.4 | 20161002 11:07:51 | 1475428071 | 0.0064579 | 6.2211 | 213.2988 | 17.9521 | 20.2739 | 0.49812 | 2.0051 | 0.43824 | 0.27868 | 165.2371 | 25.9392 | 25.0116 | 3.1819 | 55.5134 | 1.9797 | 3.5661 | 1.2022 | -0.034887 | 0.039164 | 3.0261 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0085655 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.06 | 843.0502 | 0.843 | 0.196 | 1475428071 | 240.885 | 248.66 | 23.5215 | 17.2169 | 600.135 | 83.9648 | 0.19927 | 26.9378 | 25.0116 | 999.9 | 10020 | 0 | 6.9328 | 1000.05 | 0.90002 | 0.099985 | 0 | 2.761 | 0 | 7381.46 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5749 | -0.30623 | -275.1422 | 7415.505 | 15 | 2 | 507.685 | 661.693 | 25.1951 | 27.5807 | 30.0001 | 27.5315 | 27.5064 | 17.4591 | 43.4123 | 0 | 25.1926 | 230.667 | 17.2503 | 100.741 | 101.261 | -0.031 | 5.465 | 100 | 100
16.6861 | 173.2336 | 500 to 0, with leaf | 1936 | 1475428073.5 | 7512.4 | 20161002 11:07:53 | 1475428073 | 0.0064543 | 6.042 | 211.0133 | 17.7601 | 20.0238 | 0.4979 | 2.0013 | 0.43797 | 0.27851 | 165.2355 | 25.9353 | 25.0121 | 3.182 | 55.5248 | 1.9797 | 3.5654 | 1.2023 | -0.034794 | 0.039059 | 3.0196 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0083531 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.06 | 843.0502 | 0.843 | 0.196 | 1475428073 | 237.909 | 245.489 | 23.521 | 17.2165 | 599.809 | 83.9655 | 0.20033 | 26.9343 | 25.0121 | 999.9 | 9993.12 | 0 | 6.9328 | 1000.06 | 0.90002 | 0.099985 | 0 | 2.4971 | 0 | 7373.13 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5782 | -0.35955 | -275.4704 | 7407.3504 | 15 | 2 | 505.794 | 662.29 | 25.1893 | 27.5807 | 30.0001 | 27.5315 | 27.5055 | 17.3308 | 43.4123 | 0 | 25.1818 | 230.667 | 17.2503 | 100.741 | 101.261 | -0.031 | 5.465 | 100 | 100
16.325 | 171.6345 | 500 to 0, with leaf | 1937 | 1475428075.5 | 7514.4 | 20161002 11:07:55 | 1475428075 | 0.0064594 | 5.6601 | 209.4543 | 17.6288 | 19.771 | 0.49841 | 2.0009 | 0.43836 | 0.27876 | 165.2388 | 25.9292 | 25.0121 | 3.182 | 55.5417 | 1.9798 | 3.5645 | 1.2022 | -0.034783 | 0.039047 | 3.0188 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0079 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.06 | 843.0502 | 0.843 | 0.196 | 1475428075 | 234.907 | 242.084 | 23.5227 | 17.2157 | 600.04 | 83.965 | 0.20017 | 26.9302 | 25.0121 | 999.9 | 9990 | 0 | 6.9328 | 1000.06 | 0.90002 | 0.099985 | 0 | 2.6105 | 0 | 7361.67 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5935 | 0.2748 | -276.6342 | 7396.1265 | 15 | 2 | 497.285 | 661.691 | 25.1849 | 27.5807 | 30.0001 | 27.5315 | 27.5043 | 17.1131 | 43.4123 | 0 | 25.1818 | 225.833 | 17.2503 | 100.741 | 101.261 | -0.031 | 5.465 | 100 | 100
16.059 | 169.6242 | 500 to 0, with leaf | 1938 | 1475428077.5 | 7516.4 | 20161002 11:07:57 | 1475428077 | 0.0064595 | 5.3735 | 207.4996 | 17.4642 | 19.514 | 0.49855 | 2.005 | 0.43857 | 0.27889 | 165.2355 | 25.9287 | 25.009 | 3.1814 | 55.5486 | 1.9798 | 3.564 | 1.2016 | -0.034885 | 0.039162 | 3.026 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0075602 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.05 | 843.0418 | 0.843 | 0.196 | 1475428077 | 231.855 | 238.727 | 23.5225 | 17.2142 | 599.93 | 83.9649 | 0.19998 | 26.9279 | 25.009 | 999.9 | 10019.4 | 0 | 6.9328 | 1000.05 | 0.90002 | 0.099985 | 0 | 2.4131 | 0 | 7353.47 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5519 | -0.21777 | -276.3323 | 7387.9062 | 15 | 2 | 503.812 | 661.945 | 25.1799 | 27.5807 | 30.0001 | 27.5315 | 27.5041 | 16.8892 | 43.4123 | 0 | 25.1699 | 220.667 | 17.2503 | 100.741 | 101.261 | -0.031 | 5.465 | 100 | 100
15.9073 | 167.1985 | 500 to 0, with leaf | 1939 | 1475428079.5 | 7518.4 | 20161002 11:07:59 | 1475428079 | 0.0064611 | 5.2019 | 205.1352 | 17.2655 | 19.258 | 0.49877 | 2.0021 | 0.43867 | 0.27896 | 165.2371 | 25.926 | 25.0093 | 3.1814 | 55.5512 | 1.9798 | 3.5639 | 1.2017 | -0.034812 | 0.039079 | 3.0208 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0073567 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428079 | 228.808 | 235.487 | 23.522 | 17.2141 | 600.113 | 83.9665 | 0.20013 | 26.9271 | 25.0093 | 999.9 | 9998.12 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.6351 | 0 | 7345.19 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.535 | 0.10837 | -276.453 | 7379.4712 | 15 | 2 | 494.673 | 661.501 | 25.1762 | 27.5807 | 30.0001 | 27.5315 | 27.5041 | 16.7191 | 43.4123 | 0 | 25.1699 | 220.667 | 17.2503 | 100.741 | 101.262 | -0.031 | 5.465 | 100 | 100
15.5205 | 165.6733 | 500 to 0, with leaf | 1940 | 1475428081.5 | 7520.4 | 20161002 11:08:01 | 1475428081 | 0.0064642 | 4.7939 | 203.6391 | 17.1394 | 19.0023 | 0.49958 | 1.9977 | 0.43918 | 0.2793 | 165.2371 | 25.9214 | 25.0049 | 3.1806 | 55.5559 | 1.9798 | 3.5635 | 1.2008 | -0.034706 | 0.03896 | 3.0134 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0068726 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.05 | 843.0418 | 0.843 | 0.196 | 1475428081 | 225.772 | 232.023 | 23.5222 | 17.2124 | 600.22 | 83.9669 | 0.19885 | 26.9256 | 25.0049 | 999.9 | 9967.5 | 0 | 6.9328 | 1000.05 | 0.90002 | 0.099985 | 0 | 2.4345 | 0 | 7334.27 | 38.75 | 41.812 | 40.437 | 40.625 | 39.187 | 2.5406 | -0.16967 | -278.0304 | 7368.5573 | 15 | 2 | 523.305 | 661.969 | 25.1723 | 27.5807 | 30.0001 | 27.5315 | 27.5041 | 16.4898 | 43.4123 | 0 | 25.1699 | 215.833 | 17.2503 | 100.741 | 101.261 | -0.031 | 5.465 | 100 | 100
15.2837 | 163.4824 | 500 to 0, with leaf | 1941 | 1475428083.5 | 7522.4 | 20161002 11:08:03 | 1475428083 | 0.0064584 | 4.5365 | 201.5364 | 16.9629 | 18.7436 | 0.49919 | 2.0037 | 0.43903 | 0.27919 | 165.2338 | 25.9239 | 25.0017 | 3.18 | 55.5639 | 1.9798 | 3.5631 | 1.2002 | -0.034852 | 0.039124 | 3.0236 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0065674 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428083 | 222.692 | 228.67 | 23.5217 | 17.2118 | 599.675 | 83.9674 | 0.20077 | 26.9233 | 25.0017 | 999.9 | 10009.4 | 0 | 6.9328 | 999.741 | 0.89998 | 0.10002 | 0 | 2.4346 | 0 | 7322.87 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5534 | -0.0010188 | -278.4947 | 7360.0804 | 15 | 2 | 514.576 | 662.553 | 25.168 | 27.5807 | 30.0001 | 27.5315 | 27.5041 | 16.3266 | 43.4123 | 0 | 25.163 | 210.667 | 17.2503 | 100.741 | 101.26 | -0.031 | 5.465 | 100 | 100
15.1716 | 160.8841 | 500 to 0, with leaf | 1942 | 1475428085.5 | 7524.4 | 20161002 11:08:05 | 1475428085 | 0.0064643 | 4.4046 | 198.9967 | 16.7488 | 18.484 | 0.49942 | 2.0014 | 0.43915 | 0.27927 | 165.2371 | 25.9177 | 25.0059 | 3.1808 | 55.5755 | 1.9798 | 3.5624 | 1.201 | -0.034797 | 0.039062 | 3.0197 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0064108 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.05 | 843.0418 | 0.843 | 0.196 | 1475428085 | 219.613 | 225.436 | 23.523 | 17.2121 | 600.125 | 83.966 | 0.20033 | 26.9203 | 25.0059 | 999.9 | 9993.75 | 0 | 6.9328 | 1000.05 | 0.90002 | 0.099985 | 0 | 2.7025 | 0 | 7317.16 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5741 | -0.25249 | -279.4356 | 7351.7988 | 15 | 2 | 496.82 | 661.29 | 25.1643 | 27.5796 | 30.0001 | 27.5315 | 27.5041 | 16.1536 | 43.4123 | 0 | 25.163 | 210.667 | 17.2503 | 100.741 | 101.26 | -0.031 | 5.465 | 100 | 100
14.8347 | 159.0721 | 500 to 0, with leaf | 1943 | 1475428087.5 | 7526.4 | 20161002 11:08:07 | 1475428087 | 0.0064604 | 4.0466 | 197.2691 | 16.6031 | 18.2259 | 0.49929 | 2.001 | 0.43904 | 0.2792 | 165.2355 | 25.9182 | 25.0027 | 3.1802 | 55.5732 | 1.9797 | 3.5623 | 1.2005 | -0.034785 | 0.03905 | 3.019 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0059862 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.05 | 843.0418 | 0.843 | 0.196 | 1475428087 | 216.55 | 221.996 | 23.5215 | 17.2126 | 599.954 | 83.9653 | 0.19946 | 26.9196 | 25.0027 | 999.9 | 9990.62 | 0 | 6.9328 | 1000.05 | 0.90002 | 0.099985 | 0 | 2.6827 | 0 | 7306.1 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5593 | -0.28118 | -276.2332 | 7340.6792 | 15 | 2 | 513.04 | 662.085 | 25.161 | 27.5785 | 30 | 27.5315 | 27.5041 | 15.9367 | 43.4123 | 0 | 25.1592 | 205.833 | 17.2503 | 100.74 | 101.259 | -0.031 | 5.465 | 100 | 100
14.6019 | 156.939 | 500 to 0, with leaf | 1944 | 1475428089.5 | 7528.4 | 20161002 11:08:09 | 1475428089 | 0.0064608 | 3.7933 | 195.183 | 16.428 | 17.9684 | 0.49964 | 2.002 | 0.43933 | 0.27939 | 165.2372 | 25.9182 | 24.9992 | 3.1795 | 55.5748 | 1.9797 | 3.5622 | 1.1998 | -0.03481 | 0.039077 | 3.0207 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0056857 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.05 | 843.0426 | 0.843 | 0.196 | 1475428089 | 213.485 | 218.658 | 23.5211 | 17.2117 | 599.947 | 83.9666 | 0.20048 | 26.9193 | 24.9992 | 999.9 | 9997.5 | 0 | 6.9328 | 1000.05 | 0.90002 | 0.099985 | 0 | 2.47 | 0 | 7297.57 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5438 | -0.058762 | -279.7617 | 7332.2912 | 15 | 2 | 494.431 | 661.547 | 25.1587 | 27.5783 | 30.0001 | 27.5315 | 27.5041 | 15.7149 | 43.4123 | 0 | 25.1592 | 200.667 | 17.2503 | 100.74 | 101.258 | -0.031 | 5.465 | 100 | 100
14.5743 | 154.1233 | 500 to 0, with leaf | 1945 | 1475428091.5 | 7530.4 | 20161002 11:08:11 | 1475428091 | 0.0064649 | 3.7468 | 192.4066 | 16.1942 | 17.7142 | 0.5001 | 2.0001 | 0.43964 | 0.27959 | 165.2371 | 25.9136 | 24.9984 | 3.1794 | 55.5804 | 1.9796 | 3.5617 | 1.1997 | -0.034764 | 0.039026 | 3.0175 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0056305 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.05 | 843.0418 | 0.843 | 0.196 | 1475428091 | 210.465 | 215.571 | 23.5204 | 17.2093 | 600.162 | 83.9671 | 0.19963 | 26.917 | 24.9984 | 999.9 | 9984.38 | 0 | 6.9328 | 1000.05 | 0.90002 | 0.099985 | 0 | 2.9045 | 0 | 7288.95 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.551 | 1.1795 | -279.4421 | 7323.9477 | 15 | 2 | 500.939 | 661.36 | 25.1571 | 27.5783 | 30.0002 | 27.531 | 27.5041 | 15.5457 | 43.4123 | 0 | 25.1592 | 200.667 | 17.2503 | 100.741 | 101.259 | -0.031 | 5.465 | 100 | 100
14.2728 | 152.1942 | 500 to 0, with leaf | 1946 | 1475428093.5 | 7532.4 | 20161002 11:08:13 | 1475428093 | 0.0064622 | 3.4249 | 190.6007 | 16.0424 | 17.4624 | 0.49928 | 2.0024 | 0.43907 | 0.27922 | 165.2338 | 25.9144 | 25.0049 | 3.1806 | 55.589 | 1.9798 | 3.5615 | 1.2008 | -0.034821 | 0.039089 | 3.0214 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0052503 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.734 | 842.7762 | 0.843 | 0.196 | 1475428093 | 207.473 | 212.24 | 23.5222 | 17.2101 | 599.823 | 83.9672 | 0.20011 | 26.9158 | 25.0049 | 999.9 | 10000.6 | 0 | 6.9328 | 1000.04 | 0.89998 | 0.10002 | 0 | 2.5187 | 0 | 7277.88 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5831 | 0.31405 | -282.6366 | 7312.7246 | 15 | 2 | 505.68 | 661.991 | 25.1561 | 27.5783 | 30.0001 | 27.5299 | 27.5041 | 15.3592 | 43.4123 | 0 | 25.2398 | 195.833 | 17.2503 | 100.741 | 101.259 | -0.031 | 5.465 | 100 | 100
13.8734 | 150.5362 | 500 to 0, with leaf | 1947 | 1475428095.6 | 7534.5 | 20161002 11:08:15 | 1475428095.1 | 0.0064651 | 3.0033 | 189.0028 | 15.9083 | 17.1963 | 0.49955 | 2.0025 | 0.43927 | 0.27935 | 165.2338 | 25.9116 | 25.0058 | 3.1808 | 55.5991 | 1.98 | 3.5611 | 1.2008 | -0.034822 | 0.039091 | 3.0215 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0047487 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428095.1 | 204.305 | 208.629 | 23.5234 | 17.2106 | 600.024 | 83.9698 | 0.20009 | 26.9141 | 25.0058 | 999.9 | 10000.6 | 0 | 6.9328 | 1000.05 | 0.90002 | 0.099985 | 0 | 2.5965 | 0 | 7269.88 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5748 | -0.09759 | -280.667 | 7304.3488 | 15 | 2 | 500.074 | 661.477 | 25.1718 | 27.5783 | 30 | 27.5292 | 27.5041 | 15.1531 | 43.4123 | 0 | 25.2398 | 190.667 | 17.2503 | 100.741 | 101.259 | -0.031 | 5.465 | 100 | 100
13.7175 | 148.1306 | 500 to 0, with leaf | 1948 | 1475428097.6 | 7536.5 | 20161002 11:08:17 | 1475428097.1 | 0.0064636 | 2.8275 | 186.6533 | 15.7111 | 16.9403 | 0.49983 | 2 | 0.43943 | 0.27946 | 165.2338 | 25.9072 | 25.0012 | 3.1799 | 55.6057 | 1.9797 | 3.5603 | 1.2002 | -0.034762 | 0.039024 | 3.0173 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0045401 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428097.1 | 201.257 | 205.385 | 23.5201 | 17.209 | 600.048 | 83.9728 | 0.19967 | 26.9102 | 25.0012 | 999.9 | 9983.12 | 0 | 6.9328 | 1000.05 | 0.90002 | 0.099985 | 0 | 2.4731 | 0 | 7258.74 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5837 | -0.20682 | -277.4318 | 7293.1523 | 15 | 2 | 510.493 | 661.827 | 25.2041 | 27.5783 | 30 | 27.5292 | 27.5041 | 14.9769 | 43.4123 | 0 | 25.2219 | 190.667 | 17.2503 | 100.741 | 101.259 | -0.031 | 5.465 | 100 | 100
13.5196 | 145.8295 | 500 to 0, with leaf | 1949 | 1475428099.6 | 7538.5 | 20161002 11:08:19 | 1475428099.1 | 0.006465 | 2.6095 | 184.4431 | 15.5247 | 16.6838 | 0.49948 | 2.0022 | 0.43921 | 0.27931 | 165.1858 | 25.9058 | 25.0052 | 3.1807 | 55.6087 | 1.9797 | 3.56 | 1.201 | -0.034816 | 0.039084 | 3.0211 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0042828 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.739 | 842.7805 | 0.843 | 0.196 | 1475428099.1 | 198.214 | 202.105 | 23.5199 | 17.2068 | 599.987 | 83.9704 | 0.20032 | 26.9088 | 25.0052 | 999.9 | 9998.75 | 0 | 6.9328 | 1000.05 | 0.90002 | 0.099985 | 0 | 2.7193 | 0 | 7249.82 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5787 | 0.11314 | -283.7275 | 7284.6404 | 15 | 2 | 503.113 | 661.967 | 25.2278 | 27.5783 | 30.0001 | 27.5292 | 27.5041 | 14.7528 | 43.4123 | 0 | 25.2219 | 185.667 | 17.2503 | 100.741 | 101.26 | -0.031 | 5.465 | 100 | 100
13.4196 | 143.3198 | 500 to 0, with leaf | 1950 | 1475428101.6 | 7540.5 | 20161002 11:08:21 | 1475428101.1 | 0.0064701 | 2.4902 | 181.9493 | 15.3137 | 16.4302 | 0.50015 | 2.0051 | 0.43981 | 0.27969 | 165.2338 | 25.9058 | 25.0016 | 3.18 | 55.6108 | 1.9798 | 3.5601 | 1.2002 | -0.034887 | 0.039164 | 3.0261 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.00414 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428101.1 | 195.215 | 198.968 | 23.5227 | 17.2052 | 600.041 | 83.9647 | 0.19978 | 26.9089 | 25.0016 | 999.9 | 10020 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.7223 | 0 | 7241.57 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.6088 | 0.48462 | -281.4383 | 7276.2342 | 15 | 2 | 509.007 | 662.22 | 25.2318 | 27.5783 | 30.0002 | 27.5292 | 27.5037 | 14.5269 | 43.4123 | 0 | 25.2219 | 180.667 | 17.2503 | 100.742 | 101.26 | -0.031 | 5.465 | 100 | 100
13.1926 | 141.2357 | 500 to 0, with leaf | 1951 | 1475428103.6 | 7542.5 | 20161002 11:08:23 | 1475428103.1 | 0.00647 | 2.2435 | 179.9391 | 15.1443 | 16.1818 | 0.50018 | 2.0019 | 0.43975 | 0.27966 | 165.2338 | 25.9038 | 25.0028 | 3.1802 | 55.6156 | 1.9799 | 3.5599 | 1.2003 | -0.034808 | 0.039075 | 3.0206 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0038486 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.74 | 842.7813 | 0.843 | 0.196 | 1475428103.1 | 192.267 | 195.755 | 23.5242 | 17.2055 | 599.918 | 83.963 | 0.2003 | 26.9083 | 25.0028 | 999.9 | 9997.5 | 0 | 6.9328 | 999.735 | 0.89998 | 0.10002 | 0 | 2.7212 | 0 | 7227.85 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.625 | 0.50857 | -284.6383 | 7264.9181 | 15 | 2 | 502.902 | 662.042 | 25.2298 | 27.5783 | 30.0002 | 27.5292 | 27.5025 | 14.3613 | 43.4123 | 0 | 25.2187 | 180.667 | 17.2503 | 100.742 | 101.261 | -0.031 | 5.465 | 100 | 100
12.6916 | 140.1059 | 500 to 0, with leaf | 1952 | 1475428105.6 | 7544.5 | 20161002 11:08:25 | 1475428105.1 | 0.0064718 | 1.7209 | 178.8755 | 15.055 | 15.9311 | 0.50032 | 2.0021 | 0.43986 | 0.27974 | 165.2338 | 25.9037 | 25.0036 | 3.1804 | 55.6175 | 1.98 | 3.56 | 1.2004 | -0.034813 | 0.039081 | 3.0209 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0032275 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428105.1 | 189.286 | 192.232 | 23.5254 | 17.2056 | 599.974 | 83.9641 | 0.2003 | 26.9088 | 25.0036 | 999.9 | 9998.75 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.1516 | 0 | 7223.5 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.6262 | 0.2559 | -279.759 | 7256.6104 | 15 | 2 | 494.414 | 661.517 | 25.2285 | 27.5783 | 30.0001 | 27.5292 | 27.5018 | 14.1686 | 43.4123 | 0 | 25.2187 | 175.667 | 17.2503 | 100.742 | 101.26 | -0.031 | 5.465 | 100 | 100
12.3751 | 138.2518 | 500 to 0, with leaf | 1953 | 1475428107.6 | 7546.5 | 20161002 11:08:27 | 1475428107.1 | 0.0064711 | 1.3836 | 177.0939 | 14.905 | 15.675 | 0.50029 | 2.004 | 0.43989 | 0.27975 | 165.2355 | 25.9048 | 25.003 | 3.1802 | 55.6193 | 1.9801 | 3.56 | 1.2002 | -0.034859 | 0.039132 | 3.0241 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0028273 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.05 | 843.0418 | 0.843 | 0.196 | 1475428107.1 | 186.242 | 188.831 | 23.5261 | 17.2066 | 599.941 | 83.9646 | 0.20001 | 26.9088 | 25.003 | 999.9 | 10011.9 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.6638 | 0 | 7214.82 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5695 | -0.18504 | -275.0072 | 7248.3669 | 15 | 2 | 492.935 | 661.54 | 25.2268 | 27.5783 | 30 | 27.5292 | 27.5018 | 13.9411 | 43.4123 | 0 | 25.2155 | 170.667 | 17.2503 | 100.742 | 101.26 | -0.031 | 5.465 | 100 | 100
12.2116 | 135.8178 | 500 to 0, with leaf | 1954 | 1475428109.6 | 7548.5 | 20161002 11:08:29 | 1475428109.1 | 0.0064714 | 1.1999 | 174.7351 | 14.7065 | 15.416 | 0.50028 | 2.0021 | 0.43983 | 0.27971 | 165.2338 | 25.9034 | 25.0036 | 3.1804 | 55.6186 | 1.98 | 3.5599 | 1.2004 | -0.034813 | 0.039081 | 3.0209 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0026103 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.729 | 842.772 | 0.843 | 0.196 | 1475428109.1 | 183.165 | 185.55 | 23.5251 | 17.2066 | 600.06 | 83.9648 | 0.19993 | 26.9083 | 25.0036 | 999.9 | 9998.75 | 0 | 6.9328 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.655 | 0 | 7203.84 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.579 | -0.51359 | -273.9709 | 7237.4088 | 15 | 2 | 493.406 | 661.26 | 25.2249 | 27.5778 | 30 | 27.5292 | 27.5018 | 13.7708 | 43.4123 | 0 | 25.2155 | 170.667 | 17.2503 | 100.742 | 101.26 | -0.031 | 5.465 | 100 | 100
11.8762 | 134.0194 | 500 to 0, with leaf | 1955 | 1475428111.6 | 7550.5 | 20161002 11:08:31 | 1475428111.1 | 0.006471 | 0.84337 | 173.003 | 14.5609 | 15.1583 | 0.50037 | 2.0032 | 0.43993 | 0.27978 | 165.2338 | 25.904 | 25.0017 | 3.18 | 55.6175 | 1.9799 | 3.5599 | 1.2001 | -0.03484 | 0.039111 | 3.0228 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0021866 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428111.1 | 180.101 | 182.11 | 23.5244 | 17.205 | 599.94 | 83.9655 | 0.20003 | 26.9083 | 25.0017 | 999.9 | 10006.2 | 0 | 6.9328 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.6101 | 0 | 7196.39 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.584 | -0.69173 | -267.9959 | 7229.2385 | 15 | 2 | 498.294 | 661.424 | 25.2223 | 27.5766 | 30.0001 | 27.5292 | 27.5018 | 13.5292 | 43.4123 | 0 | 25.2155 | 165.667 | 17.2503 | 100.742 | 101.26 | -0.031 | 5.465 | 100 | 100
11.6384 | 131.8536 | 500 to 0, with leaf | 1956 | 1475428113.6 | 7552.5 | 20161002 11:08:33 | 1475428113.1 | 0.0064728 | 0.58514 | 170.9181 | 14.3858 | 14.9007 | 0.50015 | 2.0064 | 0.43985 | 0.27971 | 165.2338 | 25.9059 | 25.0049 | 3.1806 | 55.6146 | 1.98 | 3.5602 | 1.2006 | -0.034919 | 0.0392 | 3.0283 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0018803 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428113.1 | 177.036 | 178.767 | 23.5241 | 17.2038 | 600.021 | 83.9677 | 0.19987 | 26.9094 | 25.0049 | 999.9 | 10028.8 | 0 | 6.9328 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.4843 | 0 | 7188.34 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5885 | -0.70779 | -263.0338 | 7221.0869 | 15 | 2 | 498.008 | 661.284 | 25.2197 | 27.576 | 30.0001 | 27.5292 | 27.5018 | 13.3141 | 43.4123 | 0 | 25.2124 | 160.667 | 17.2503 | 100.742 | 101.26 | -0.031 | 5.465 | 100 | 100
11.4991 | 129.3782 | 500 to 0, with leaf | 1957 | 1475428115.6 | 7554.5 | 20161002 11:08:35 | 1475428115.1 | 0.0064743 | 0.42589 | 168.4877 | 14.1814 | 14.6429 | 0.50068 | 2.0007 | 0.44011 | 0.2799 | 165.2321 | 25.9027 | 25.003 | 3.1802 | 55.6159 | 1.98 | 3.5602 | 1.2002 | -0.034778 | 0.039041 | 3.0185 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0016914 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428115.1 | 173.97 | 175.522 | 23.5243 | 17.2032 | 600.087 | 83.9689 | 0.19993 | 26.9094 | 25.003 | 999.9 | 9988.12 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.6988 | 0 | 7177.6 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5725 | -0.49762 | -264.3768 | 7210.5546 | 15 | 2 | 500.712 | 661.448 | 25.2175 | 27.576 | 30.0001 | 27.5292 | 27.5018 | 13.1857 | 43.4123 | 0 | 25.2124 | 160.667 | 17.2503 | 100.742 | 101.261 | -0.031 | 5.465 | 100 | 100
11.2606 | 127.2711 | 500 to 0, with leaf | 1958 | 1475428117.6 | 7556.5 | 20161002 11:08:37 | 1475428117.1 | 0.0064745 | 0.16711 | 166.4394 | 14.0088 | 14.3874 | 0.50087 | 2.0006 | 0.44025 | 0.27999 | 165.1845 | 25.9011 | 25.0012 | 3.1799 | 55.6201 | 1.98 | 3.5599 | 1.1999 | -0.034775 | 0.039038 | 3.0183 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0013844 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.03 | 843.0249 | 0.843 | 0.196 | 1475428117.1 | 170.937 | 172.211 | 23.5249 | 17.2019 | 599.927 | 83.9674 | 0.20027 | 26.9083 | 25.0012 | 999.9 | 9987.5 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.5602 | 0 | 7169.61 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.5695 | 0.40128 | -264.2865 | 7202.4323 | 15 | 2 | 497.939 | 661.634 | 25.2156 | 27.576 | 30 | 27.5292 | 27.5018 | 12.9565 | 43.4123 | 0 | 25.2096 | 155.667 | 17.2503 | 100.742 | 101.261 | -0.031 | 5.465 | 100 | 100
11.0867 | 124.9227 | 500 to 0, with leaf | 1959 | 1475428119.6 | 7558.5 | 20161002 11:08:39 | 1475428119.1 | 0.0064772 | -0.026606 | 164.1762 | 13.818 | 14.1334 | 0.50064 | 2.001 | 0.44008 | 0.27988 | 165.2355 | 25.8986 | 25.0058 | 3.1808 | 55.6255 | 1.98 | 3.5595 | 1.2008 | -0.034786 | 0.03905 | 3.019 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.0011546 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.05 | 843.0418 | 0.843 | 0.196 | 1475428119.1 | 167.924 | 168.985 | 23.5249 | 17.2005 | 600.043 | 83.9657 | 0.19976 | 26.9062 | 25.0058 | 999.9 | 9990.62 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.8784 | 0 | 7161.71 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5504 | 0.67202 | -267.5374 | 7194.6858 | 15 | 2 | 502.549 | 661.447 | 25.2138 | 27.576 | 30 | 27.5281 | 27.5018 | 12.73 | 43.4123 | 0 | 25.2096 | 150.667 | 17.2503 | 100.742 | 101.261 | -0.031 | 5.465 | 100 | 100
10.8732 | 122.744 | 500 to 0, with leaf | 1960 | 1475428121.6 | 7560.5 | 20161002 11:08:41 | 1475428121.1 | 0.0064756 | -0.2601 | 162.0703 | 13.6406 | 13.8808 | 0.50061 | 2.0008 | 0.44005 | 0.27986 | 165.2355 | 25.8979 | 25.0046 | 3.1805 | 55.6297 | 1.98 | 3.5592 | 1.2006 | -0.034781 | 0.039044 | 3.0186 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.00087766 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428121.1 | 164.924 | 165.732 | 23.5252 | 17.201 | 599.914 | 83.9644 | 0.20032 | 26.905 | 25.0046 | 999.9 | 9989.38 | 0 | 6.9328 | 999.725 | 0.89998 | 0.10002 | 0 | 2.522 | 0 | 7148.56 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5634 | 0.47598 | -267.055 | 7183.9485 | 15 | 2 | 496.527 | 661.564 | 25.2115 | 27.576 | 30.0001 | 27.5269 | 27.5018 | 12.5576 | 43.4123 | 0 | 25.2096 | 150.667 | 17.2503 | 100.742 | 101.261 | -0.031 | 5.465 | 100 | 100
10.5265 | 121.0424 | 500 to 0, with leaf | 1961 | 1475428123.6 | 7562.5 | 20161002 11:08:43 | 1475428123.1 | 0.0064769 | -0.62762 | 160.4215 | 13.5018 | 13.6269 | 0.50068 | 2.0088 | 0.44032 | 0.28001 | 165.2355 | 25.8999 | 25.0022 | 3.1801 | 55.6343 | 1.98 | 3.559 | 1.2001 | -0.034979 | 0.039267 | 3.0325 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.00044172 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428123.1 | 161.908 | 162.329 | 23.5255 | 17.2016 | 600.055 | 83.9649 | 0.19986 | 26.9038 | 25.0022 | 999.9 | 10046.2 | 0 | 6.9328 | 1000.05 | 0.90002 | 0.099985 | 0 | 2.4688 | 0 | 7143.89 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.5762 | -0.17008 | -263.7593 | 7175.9915 | 15 | 2 | 496.781 | 661.377 | 25.2096 | 27.576 | 30.0001 | 27.5268 | 27.5018 | 12.3148 | 43.4123 | 0 | 25.205 | 145.667 | 17.2503 | 100.742 | 101.261 | -0.031 | 5.465 | 100 | 100
10.2914 | 118.8966 | 500 to 0, with leaf | 1962 | 1475428125.6 | 7564.5 | 20161002 11:08:45 | 1475428125.1 | 0.0064769 | -0.88305 | 158.3585 | 13.3284 | 13.3718 | 0.50067 | 2.0037 | 0.44017 | 0.27993 | 165.2338 | 25.8969 | 25.0052 | 3.1807 | 55.6405 | 1.9802 | 3.5589 | 1.2005 | -0.034851 | 0.039123 | 3.0236 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.00013873 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428125.1 | 158.874 | 159.02 | 23.5269 | 17.2011 | 599.882 | 83.9662 | 0.19991 | 26.9032 | 25.0052 | 999.9 | 10009.4 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.9596 | 0 | 7136.23 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5803 | -0.094786 | -262.7265 | 7168.2727 | 15 | 2 | 504.235 | 661.728 | 25.2078 | 27.576 | 30.0001 | 27.5268 | 27.5018 | 12.1481 | 43.4123 | 0 | 25.205 | 140.667 | 17.2503 | 100.741 | 101.26 | -0.031 | 5.465 | 100 | 100
10.1215 | 116.5733 | 500 to 0, with leaf | 1963 | 1475428127.6 | 7566.5 | 20161002 11:08:47 | 1475428127.1 | 0.0064804 | -1.0729 | 156.0593 | 13.1353 | 13.1166 | 0.50147 | 1.9989 | 0.44066 | 0.28026 | 165.2355 | 25.8929 | 25.002 | 3.1801 | 55.6442 | 1.9802 | 3.5587 | 1.1998 | -0.034734 | 0.038992 | 3.0154 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -8.6488e-05 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.744 | 842.7847 | 0.843 | 0.196 | 1475428127.1 | 155.837 | 155.774 | 23.527 | 17.2002 | 600.104 | 83.9682 | 0.20021 | 26.9026 | 25.002 | 999.9 | 9975.62 | 0 | 6.9328 | 999.735 | 0.89998 | 0.10002 | 0 | 2.3536 | 0 | 7123.78 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5698 | -0.32175 | -261.2109 | 7157.6869 | 15 | 2 | 491.677 | 660.907 | 25.2055 | 27.576 | 30.0001 | 27.5268 | 27.5014 | 11.9696 | 43.4123 | 0 | 25.2018 | 140.667 | 17.2503 | 100.742 | 101.26 | -0.031 | 5.465 | 100 | 100
9.769 | 114.8772 | 500 to 0, with leaf | 1964 | 1475428129.6 | 7568.5 | 20161002 11:08:49 | 1475428129.1 | 0.0064794 | -1.4462 | 154.4272 | 12.9976 | 12.8623 | 0.50149 | 2.0027 | 0.44079 | 0.28033 | 165.2338 | 25.8928 | 24.9986 | 3.1794 | 55.6477 | 1.9801 | 3.5583 | 1.1993 | -0.034827 | 0.039097 | 3.0219 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.00052932 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428129.1 | 152.819 | 152.363 | 23.5261 | 17.1998 | 600.067 | 83.967 | 0.19958 | 26.9005 | 24.9986 | 999.9 | 10002.5 | 0 | 6.9328 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.6691 | 0 | 7119.2 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5913 | 0.10915 | -255.9282 | 7150.01 | 15 | 2 | 504.747 | 661.29 | 25.2036 | 27.576 | 30 | 27.5268 | 27.5003 | 11.7423 | 43.4123 | 0 | 25.2018 | 135.667 | 17.2503 | 100.743 | 101.26 | -0.031 | 5.465 | 100 | 100
9.6144 | 112.456 | 500 to 0, with leaf | 1965 | 1475428131.6 | 7570.5 | 20161002 11:08:51 | 1475428131.1 | 0.00648 | -1.6206 | 152.0938 | 12.801 | 12.608 | 0.50118 | 2.0045 | 0.44059 | 0.2802 | 165.2321 | 25.8927 | 25.0019 | 3.18 | 55.6503 | 1.9801 | 3.5582 | 1.1999 | -0.034872 | 0.039147 | 3.025 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.00073615 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.03 | 843.0249 | 0.843 | 0.196 | 1475428131.1 | 149.801 | 149.151 | 23.5266 | 17.1979 | 599.891 | 83.9649 | 0.20037 | 26.8998 | 25.0019 | 999.9 | 10015.6 | 0 | 6.9328 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.4986 | 0 | 7112.05 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.598 | 0.16972 | -250.3142 | 7142.4304 | 15 | 2 | 495.8 | 661.7 | 25.202 | 27.576 | 30.0001 | 27.5268 | 27.4995 | 11.5116 | 43.4123 | 0 | 25.2018 | 130.667 | 17.2503 | 100.743 | 101.26 | -0.031 | 5.465 | 100 | 100
9.4833 | 110.0061 | 500 to 0, with leaf | 1966 | 1475428133.6 | 7572.5 | 20161002 11:08:53 | 1475428133.1 | 0.0064851 | -1.7711 | 149.687 | 12.5984 | 12.3553 | 0.50158 | 2.0032 | 0.44086 | 0.28038 | 165.2321 | 25.8897 | 25.003 | 3.1802 | 55.6524 | 1.9801 | 3.558 | 1.2001 | -0.034839 | 0.03911 | 3.0227 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.00091469 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428133.1 | 146.799 | 145.98 | 23.5269 | 17.1952 | 600.079 | 83.9648 | 0.20004 | 26.8993 | 25.003 | 999.9 | 10006.2 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.3081 | 0 | 7102.27 | 38.75 | 41.812 | 40.375 | 40.625 | 39.125 | 2.5887 | -0.28243 | -246.6561 | 7132.5012 | 15 | 2 | 492.458 | 661.185 | 25.2006 | 27.576 | 30.0001 | 27.5268 | 27.4995 | 11.3407 | 43.4123 | 0 | 25.201 | 130.667 | 17.2503 | 100.742 | 101.259 | -0.031 | 5.465 | 100 | 100
9.0769 | 108.5073 | 500 to 0, with leaf | 1967 | 1475428135.6 | 7574.5 | 20161002 11:08:55 | 1475428135.1 | 0.0064807 | -2.1985 | 148.2904 | 12.4807 | 12.1047 | 0.50119 | 2.0027 | 0.44056 | 0.28018 | 165.1854 | 25.8897 | 25.0029 | 3.1802 | 55.6548 | 1.9801 | 3.5578 | 1.2001 | -0.034828 | 0.039098 | 3.0219 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0014221 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.728 | 842.7712 | 0.843 | 0.196 | 1475428135.1 | 143.822 | 142.555 | 23.5266 | 17.1956 | 599.742 | 83.9637 | 0.20045 | 26.8982 | 25.0029 | 999.9 | 10003.1 | 0 | 6.9328 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.6521 | 0 | 7095.38 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.5639 | -0.090732 | -246.3361 | 7124.9242 | 15 | 2 | 489.904 | 661.583 | 25.2 | 27.575 | 30 | 27.5268 | 27.4995 | 11.1445 | 43.4123 | 0 | 25.201 | 125.667 | 17.2503 | 100.742 | 101.259 | -0.031 | 5.465 | 100 | 100
8.9532 | 106.0552 | 500 to 0, with leaf | 1968 | 1475428137.6 | 7576.5 | 20161002 11:08:57 | 1475428137.1 | 0.0064851 | -2.3417 | 145.852 | 12.2754 | 11.8518 | 0.502 | 2.0018 | 0.44115 | 0.28057 | 165.2321 | 25.8862 | 24.9986 | 3.1794 | 55.661 | 1.9801 | 3.5574 | 1.1993 | -0.034807 | 0.039073 | 3.0204 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0015915 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.03 | 843.0249 | 0.843 | 0.196 | 1475428137.1 | 140.819 | 139.391 | 23.5269 | 17.1963 | 600.182 | 83.964 | 0.19947 | 26.8964 | 24.9986 | 999.9 | 9996.88 | 0 | 6.9328 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.5166 | 0 | 7088.39 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.5653 | -0.20683 | -243.1005 | 7117.7335 | 15 | 2 | 496.274 | 661.117 | 25.1996 | 27.5739 | 30 | 27.5268 | 27.4995 | 10.9038 | 43.4123 | 0 | 25.1988 | 120.667 | 17.2503 | 100.743 | 101.259 | -0.031 | 5.465 | 100 | 100
8.6255 | 104.3351 | 500 to 0, with leaf | 1969 | 1475428139.6 | 7578.5 | 20161002 11:08:59 | 1475428139.1 | 0.0064873 | -2.6898 | 144.1724 | 12.134 | 11.602 | 0.50253 | 1.9988 | 0.44149 | 0.28079 | 165.2321 | 25.8829 | 24.9961 | 3.1789 | 55.6655 | 1.9801 | 3.5572 | 1.1988 | -0.034731 | 0.038988 | 3.0151 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0020044 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.03 | 843.0249 | 0.843 | 0.196 | 1475428139.1 | 137.851 | 136.056 | 23.5273 | 17.1945 | 600.173 | 83.9647 | 0.19869 | 26.8953 | 24.9961 | 999.9 | 9975 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.3545 | 0 | 7079.25 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5791 | -0.42039 | -237.3723 | 7108.1346 | 15 | 2 | 533.66 | 662.029 | 25.1993 | 27.5737 | 30 | 27.5268 | 27.4995 | 10.7318 | 43.4123 | 0 | 25.1988 | 120.667 | 17.2503 | 100.743 | 101.259 | -0.031 | 5.465 | 100 | 100
7.9993 | 103.5133 | 500 to 0, with leaf | 1970 | 1475428141.6 | 7580.5 | 20161002 11:09:01 | 1475428141.1 | 0.0064863 | -3.3387 | 143.5001 | 12.0777 | 11.3463 | 0.5016 | 1.9972 | 0.44072 | 0.2803 | 165.2321 | 25.882 | 25.0074 | 3.1811 | 55.674 | 1.9804 | 3.5571 | 1.2007 | -0.034693 | 0.038946 | 3.0125 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0027742 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428141.1 | 134.809 | 132.344 | 23.5295 | 17.1941 | 599.833 | 83.9641 | 0.20131 | 26.8947 | 25.0074 | 999.9 | 9964.38 | 0 | 6.9328 | 999.736 | 0.89998 | 0.10002 | 0 | 2.5452 | 0 | 7070.85 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.5767 | -0.21438 | -232.9265 | 7100.99 | 15 | 2 | 497.045 | 661.445 | 25.1987 | 27.5737 | 30 | 27.5268 | 27.4995 | 10.4869 | 43.4123 | 0 | 25.1988 | 115.667 | 17.2503 | 100.742 | 101.259 | -0.031 | 5.465 | 100 | 100
7.8207 | 101.0773 | 500 to 0, with leaf | 1971 | 1475428143.6 | 7582.5 | 20161002 11:09:03 | 1475428143.1 | 0.006492 | -3.5379 | 141.1507 | 11.8798 | 11.0831 | 0.50125 | 2.0018 | 0.44058 | 0.2802 | 165.2321 | 25.8793 | 25.0145 | 3.1824 | 55.6806 | 1.9803 | 3.5565 | 1.2021 | -0.034805 | 0.039071 | 3.0203 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0030105 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.03 | 843.0249 | 0.843 | 0.196 | 1475428143.1 | 131.685 | 129.003 | 23.529 | 17.1921 | 600.222 | 83.9643 | 0.19955 | 26.892 | 25.0145 | 999.9 | 9996.25 | 0 | 6.9328 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.675 | 0 | 7066.76 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.5967 | -0.13741 | -232.5788 | 7094.1681 | 15 | 2 | 504.454 | 661.047 | 25.1979 | 27.5737 | 30 | 27.5258 | 27.4995 | 10.2678 | 43.4123 | 0 | 25.1976 | 110.667 | 17.2503 | 100.742 | 101.259 | -0.031 | 5.465 | 100 | 100
7.706 | 98.4604 | 500 to 0, with leaf | 1972 | 1475428145.6 | 7584.5 | 20161002 11:09:05 | 1475428145.1 | 0.0064865 | -3.6725 | 138.6116 | 11.6662 | 10.8232 | 0.5009 | 2.0103 | 0.44053 | 0.28015 | 165.2338 | 25.8839 | 25.0101 | 3.1816 | 55.6862 | 1.9803 | 3.5562 | 1.2013 | -0.035016 | 0.039308 | 3.0351 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0031701 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.03 | 843.0249 | 0.843 | 0.196 | 1475428145.1 | 128.596 | 125.756 | 23.5294 | 17.1918 | 599.646 | 83.9648 | 0.19965 | 26.8907 | 25.0101 | 999.9 | 10056.9 | 0 | 6.9328 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.4464 | 0 | 7058.41 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.5471 | -0.13735 | -223.0068 | 7085.1127 | 15 | 2 | 527.932 | 662.706 | 25.1975 | 27.5737 | 30 | 27.5246 | 27.4995 | 10.1351 | 43.4123 | 0 | 25.1976 | 110.667 | 17.2503 | 100.743 | 101.26 | -0.031 | 5.465 | 100 | 100
7.399 | 96.6446 | 500 to 0, with leaf | 1973 | 1475428147.6 | 7586.5 | 20161002 11:09:07 | 1475428147.1 | 0.0064934 | -4.0004 | 136.7421 | 11.509 | 10.566 | 0.50281 | 2.0063 | 0.4419 | 0.28104 | 165.2321 | 25.8805 | 24.9987 | 3.1794 | 55.6904 | 1.9806 | 3.5565 | 1.1988 | -0.034916 | 0.039196 | 3.0281 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.003559 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428147.1 | 125.538 | 122.353 | 23.5322 | 17.1921 | 600.043 | 83.9647 | 0.20098 | 26.8917 | 24.9987 | 999.9 | 10028.1 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.5879 | 0 | 7052.13 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.5594 | 0.29434 | -216.1063 | 7078.455 | 15 | 2 | 493.32 | 661.082 | 25.1971 | 27.5737 | 30 | 27.5245 | 27.4985 | 9.9071 | 43.4123 | 0 | 25.1879 | 105.667 | 17.2503 | 100.743 | 101.26 | -0.031 | 5.465 | 100 | 100
7.2182 | 94.3147 | 500 to 0, with leaf | 1974 | 1475428149.6 | 7588.5 | 20161002 11:09:09 | 1475428149.1 | 0.0064975 | -4.2012 | 134.4414 | 11.3152 | 10.308 | 0.5036 | 1.9995 | 0.44233 | 0.28134 | 165.183 | 25.874 | 24.996 | 3.1789 | 55.6921 | 1.9805 | 3.5561 | 1.1984 | -0.034748 | 0.039008 | 3.0164 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0037985 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.722 | 842.7661 | 0.843 | 0.196 | 1475428149.1 | 122.475 | 119.071 | 23.5312 | 17.1893 | 600.258 | 83.9652 | 0.1993 | 26.8902 | 24.996 | 999.9 | 9980 | 0 | 6.9328 | 999.727 | 0.89998 | 0.10002 | 0 | 2.4352 | 0 | 7043.28 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.5782 | 0.30275 | -214.8844 | 7071.7188 | 15 | 2 | 509.168 | 660.997 | 25.196 | 27.5737 | 30 | 27.5245 | 27.4974 | 9.6789 | 43.4123 | 0 | 25.1879 | 100.667 | 17.2503 | 100.744 | 101.26 | -0.031 | 5.465 | 100 | 100
7.0542 | 91.8714 | 500 to 0, with leaf | 1975 | 1475428151.6 | 7590.5 | 20161002 11:09:11 | 1475428151.1 | 0.006493 | -4.3851 | 132.1624 | 11.1236 | 10.0535 | 0.50222 | 2.001 | 0.44131 | 0.28067 | 165.2305 | 25.8741 | 25.0063 | 3.1809 | 55.7006 | 1.9805 | 3.5556 | 1.2004 | -0.034785 | 0.03905 | 3.019 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0040155 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.02 | 843.0165 | 0.843 | 0.196 | 1475428151.1 | 119.449 | 115.838 | 23.5308 | 17.188 | 599.759 | 83.9654 | 0.20044 | 26.8876 | 25.0063 | 999.9 | 9990.62 | 0 | 6.9328 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.7681 | 0 | 7037.87 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.5784 | 0.085538 | -211.453 | 7063.3862 | 15 | 2 | 505.593 | 661.741 | 25.1926 | 27.5737 | 30 | 27.5245 | 27.4971 | 9.5059 | 43.4123 | 0 | 25.1879 | 100.667 | 17.2503 | 100.745 | 101.26 | -0.031 | 5.465 | 100 | 100
6.7418 | 90.0701 | 500 to 0, with leaf | 1976 | 1475428153.6 | 7592.5 | 20161002 11:09:13 | 1475428153.1 | 0.0064981 | -4.718 | 130.3874 | 10.9739 | 9.8009 | 0.50276 | 2.0047 | 0.44182 | 0.28099 | 165.2305 | 25.8725 | 25.0048 | 3.1806 | 55.7108 | 1.9807 | 3.5553 | 1.1999 | -0.034876 | 0.039151 | 3.0253 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0044103 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.03 | 843.0249 | 0.843 | 0.196 | 1475428153.1 | 116.451 | 112.49 | 23.5338 | 17.1891 | 600.045 | 83.9633 | 0.20034 | 26.8862 | 25.0048 | 999.9 | 10016.9 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.5247 | 0 | 7032.51 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.6024 | 0.35917 | -205.3973 | 7057.1712 | 15 | 2 | 495.693 | 661.156 | 25.1888 | 27.5726 | 30 | 27.5245 | 27.4971 | 9.2581 | 43.4123 | 0 | 25.1875 | 95.6667 | 17.2503 | 100.745 | 101.26 | -0.031 | 5.465 | 100 | 100
6.5332 | 87.8659 | 500 to 0, with leaf | 1977 | 1475428155.6 | 7594.5 | 20161002 11:09:15 | 1475428155.1 | 0.0064998 | -4.9467 | 128.2404 | 10.7929 | 9.5469 | 0.50297 | 2.0037 | 0.44196 | 0.28109 | 165.2305 | 25.8708 | 25.0043 | 3.1805 | 55.7129 | 1.9807 | 3.5552 | 1.1998 | -0.034851 | 0.039124 | 3.0236 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0046817 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.02 | 843.0165 | 0.843 | 0.196 | 1475428155.1 | 113.436 | 109.227 | 23.5345 | 17.1883 | 600.058 | 83.9618 | 0.19955 | 26.8856 | 25.0043 | 999.9 | 10010 | 0 | 6.9328 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.8548 | 0 | 7026.87 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.609 | 0.32352 | -200.0154 | 7051.0904 | 15 | 2 | 505.803 | 661.553 | 25.187 | 27.5715 | 30 | 27.5245 | 27.4971 | 9.0898 | 43.4123 | 0 | 25.1875 | 90.6667 | 17.2503 | 100.745 | 101.259 | -0.031 | 5.465 | 100 | 100
6.3475 | 85.5691 | 500 to 0, with leaf | 1978 | 1475428157.6 | 7596.5 | 20161002 11:09:17 | 1475428157.1 | 0.0064993 | -5.1521 | 126.0415 | 10.608 | 9.2946 | 0.50263 | 2.0018 | 0.44164 | 0.28089 | 165.2321 | 25.8691 | 25.0083 | 3.1813 | 55.7155 | 1.9807 | 3.555 | 1.2006 | -0.034806 | 0.039073 | 3.0204 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0049267 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.729 | 842.772 | 0.843 | 0.196 | 1475428157.1 | 110.435 | 106 | 23.5336 | 17.1863 | 599.914 | 83.9626 | 0.2005 | 26.8845 | 25.0083 | 999.9 | 9996.88 | 0 | 6.9328 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.4871 | 0 | 7019.19 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.588 | -0.07439 | -199.2017 | 7043.1719 | 15 | 2 | 491.194 | 661.225 | 25.1863 | 27.5714 | 29.9999 | 27.5245 | 27.4971 | 8.9059 | 43.4123 | 0 | 25.1822 | 90.6667 | 17.2503 | 100.745 | 101.258 | -0.031 | 5.465 | 100 | 100
5.9597 | 84.0261 | 500 to 0, with leaf | 1979 | 1475428159.6 | 7598.5 | 20161002 11:09:19 | 1475428159.1 | 0.0065005 | -5.561 | 124.5498 | 10.4828 | 9.0433 | 0.50296 | 2.0011 | 0.44188 | 0.28104 | 165.2321 | 25.8661 | 25.0068 | 3.181 | 55.7258 | 1.9808 | 3.5545 | 1.2002 | -0.034788 | 0.039052 | 3.0191 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0054102 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.03 | 843.0249 | 0.843 | 0.196 | 1475428159.1 | 107.446 | 102.584 | 23.5341 | 17.1872 | 600.061 | 83.9661 | 0.19978 | 26.8823 | 25.0068 | 999.9 | 9991.25 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.3651 | 0 | 7014.29 | 38.75 | 41.812 | 40.437 | 40.625 | 39.125 | 2.5954 | 0.39477 | -193.5809 | 7037.2985 | 15 | 2 | 497.58 | 661.015 | 25.1855 | 27.5714 | 30 | 27.5245 | 27.4971 | 8.6774 | 43.4123 | 0 | 25.1822 | 85.6667 | 17.2503 | 100.745 | 101.259 | -0.031 | 5.465 | 100 | 100
5.7465 | 81.8033 | 500 to 0, with leaf | 1980 | 1475428161.6 | 7600.5 | 20161002 11:09:21 | 1475428161.1 | 0.0064988 | -5.7942 | 122.4307 | 10.3046 | 8.7889 | 0.50248 | 2.002 | 0.44153 | 0.28081 | 165.2321 | 25.866 | 25.01 | 3.1816 | 55.7287 | 1.9807 | 3.5542 | 1.2008 | -0.03481 | 0.039077 | 3.0207 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0056869 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.03 | 843.0249 | 0.843 | 0.196 | 1475428161.1 | 104.422 | 99.3056 | 23.5334 | 17.1866 | 599.907 | 83.9668 | 0.20009 | 26.8811 | 25.01 | 999.9 | 9997.5 | 0 | 6.9328 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.4601 | 0 | 7009.62 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.5883 | -0.87674 | -185.92 | 7031.6092 | 15 | 2 | 498.571 | 661.389 | 25.1836 | 27.5714 | 30.0001 | 27.5234 | 27.4971 | 8.4449 | 43.4123 | 0 | 25.1822 | 80.6667 | 17.2503 | 100.745 | 101.259 | -0.031 | 5.465 | 100 | 100
5.594 | 79.4176 | 500 to 0, with leaf | 1981 | 1475428163.6 | 7602.5 | 20161002 11:09:23 | 1475428163.1 | 0.0065036 | -5.9663 | 120.0964 | 10.1079 | 8.5362 | 0.50282 | 1.9984 | 0.4417 | 0.28093 | 165.2338 | 25.8634 | 25.012 | 3.182 | 55.7254 | 1.9807 | 3.5544 | 1.2012 | -0.034722 | 0.038978 | 3.0145 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0058911 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.03 | 843.0249 | 0.843 | 0.196 | 1475428163.1 | 101.423 | 96.1177 | 23.5339 | 17.185 | 600.157 | 83.9645 | 0.20003 | 26.882 | 25.012 | 999.9 | 9972.5 | 0 | 6.9328 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.5228 | 0 | 7002.81 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.5615 | -0.70946 | -176.3644 | 7024.2392 | 15 | 2 | 492.046 | 661.082 | 25.1815 | 27.5714 | 30.0001 | 27.5223 | 27.4968 | 8.267 | 43.4123 | 0 | 25.1737 | 80.6667 | 17.2503 | 100.745 | 101.259 | -0.031 | 5.465 | 100 | 100
5.271 | 77.6652 | 500 to 0, with leaf | 1982 | 1475428165.6 | 7604.5 | 20161002 11:09:25 | 1475428165.1 | 0.0065017 | -6.3097 | 118.3934 | 9.9645 | 8.2863 | 0.50297 | 2.0042 | 0.44197 | 0.28109 | 165.2338 | 25.8664 | 25.0062 | 3.1809 | 55.7257 | 1.9807 | 3.5543 | 1.2002 | -0.034863 | 0.039137 | 3.0244 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0063003 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.727 | 842.7703 | 0.843 | 0.196 | 1475428165.1 | 98.4542 | 92.7843 | 23.5335 | 17.1844 | 599.964 | 83.9645 | 0.19991 | 26.8816 | 25.0062 | 999.9 | 10013.1 | 0 | 6.9328 | 1000.04 | 0.89998 | 0.10002 | 0 | 2.7258 | 0 | 6998.06 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.5666 | -0.56326 | -177.7853 | 7019.0054 | 15 | 2 | 500.065 | 661.348 | 25.1788 | 27.5714 | 30.0001 | 27.5222 | 27.4956 | 8.0706 | 43.4123 | 0 | 25.1737 | 75.6667 | 17.2503 | 100.744 | 101.259 | -0.031 | 5.465 | 100 | 100
4.9495 | 75.8922 | 500 to 0, with leaf | 1983 | 1475428167.6 | 7606.5 | 20161002 11:09:27 | 1475428167.1 | 0.0065031 | -6.6517 | 116.6753 | 9.82 | 8.0356 | 0.50313 | 2.0067 | 0.44216 | 0.28121 | 165.2321 | 25.8664 | 25.0051 | 3.1806 | 55.7292 | 1.9807 | 3.5542 | 1.1999 | -0.034927 | 0.039208 | 3.0288 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0067041 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.03 | 843.0249 | 0.843 | 0.196 | 1475428167.1 | 95.4744 | 89.4434 | 23.5339 | 17.1837 | 599.985 | 83.9653 | 0.19998 | 26.881 | 25.0051 | 999.9 | 10031.2 | 0 | 6.9328 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.7865 | 0 | 6993.94 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.588 | -0.048485 | -171.1474 | 7013.8392 | 15 | 2 | 496.958 | 661.29 | 25.1751 | 27.5714 | 30.0001 | 27.5222 | 27.4948 | 7.8331 | 43.1359 | 0 | 25.1651 | 70.6667 | 17.2503 | 100.743 | 101.259 | -0.031 | 5.465 | 100 | 100
4.7036 | 73.8072 | 500 to 0, with leaf | 1984 | 1475428169.6 | 7608.5 | 20161002 11:09:29 | 1475428169.1 | 0.0064964 | -6.9178 | 114.6978 | 9.6535 | 7.7828 | 0.50283 | 2 | 0.44175 | 0.28096 | 165.2338 | 25.8652 | 25.0038 | 3.1804 | 55.7276 | 1.9806 | 3.5541 | 1.1998 | -0.034761 | 0.039023 | 3.0173 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0070196 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428169.1 | 92.4708 | 86.1546 | 23.5325 | 17.1899 | 600.083 | 83.9651 | 0.19983 | 26.8804 | 25.0038 | 999.9 | 9983.75 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.7939 | 0 | 6988.42 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.5852 | 0.68245 | -160.6646 | 7007.2123 | 15 | 2 | 502.113 | 661.29 | 25.1723 | 27.5703 | 30.0001 | 27.5222 | 27.4948 | 7.6517 | 43.1359 | 0 | 25.1651 | 70.6667 | 17.2503 | 100.743 | 101.259 | -0.031 | 5.465 | 100 | 100
4.3467 | 72.0213 | 500 to 0, with leaf | 1985 | 1475428171.6 | 7610.5 | 20161002 11:09:31 | 1475428171.1 | 0.0064692 | -7.2959 | 113.1463 | 9.523 | 7.5259 | 0.50054 | 2.0035 | 0.44007 | 0.27987 | 165.2338 | 25.8749 | 25.0022 | 3.1801 | 55.7389 | 1.9808 | 3.5537 | 1.1993 | -0.034848 | 0.03912 | 3.0234 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0074682 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428171.1 | 89.4182 | 82.6986 | 23.5344 | 17.2155 | 599.811 | 83.9649 | 0.20048 | 26.8784 | 25.0022 | 999.9 | 10008.8 | 0 | 6.9189 | 999.723 | 0.89998 | 0.10002 | 0 | 2.6065 | 0 | 6982.01 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.5967 | 0.61109 | -154.7118 | 7002.3431 | 15 | 2 | 495.704 | 661.594 | 25.1692 | 27.5691 | 30.0001 | 27.5222 | 27.4948 | 7.4041 | 43.1359 | 0 | 25.1651 | 65.6667 | 17.2503 | 100.743 | 101.26 | -0.031 | 5.465 | 100 | 100
4.061 | 69.9908 | 500 to 0, with leaf | 1986 | 1475428173.6 | 7612.5 | 20161002 11:09:33 | 1475428173.1 | 0.0064562 | -7.6024 | 111.3134 | 9.3686 | 7.2679 | 0.49894 | 2.0007 | 0.43876 | 0.27902 | 165.2321 | 25.8761 | 25.0126 | 3.1821 | 55.7694 | 1.9816 | 3.5532 | 1.2005 | -0.034778 | 0.039042 | 3.0185 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0078317 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.03 | 843.0249 | 0.843 | 0.196 | 1475428173.1 | 86.354 | 79.3111 | 23.5444 | 17.2419 | 600.165 | 83.9645 | 0.19971 | 26.8761 | 25.0126 | 999.9 | 9988.75 | 0 | 6.9189 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.6855 | 0 | 6980.53 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.5753 | -0.024509 | -149.6024 | 6997.935 | 15 | 2 | 495.842 | 660.986 | 25.1654 | 27.569 | 29.9999 | 27.5222 | 27.4948 | 7.18 | 43.1359 | 0 | 25.1616 | 60.6667 | 17.2503 | 100.743 | 101.261 | -0.031 | 5.465 | 100 | 100
3.8625 | 67.7338 | 500 to 0, with leaf | 1987 | 1475428175.6 | 7614.5 | 20161002 11:09:35 | 1475428175.1 | 0.0064594 | -7.8211 | 109.0563 | 9.1786 | 7.0107 | 0.50001 | 1.9976 | 0.4395 | 0.27951 | 165.2338 | 25.8722 | 25.0098 | 3.1815 | 55.7996 | 1.9825 | 3.5529 | 1.199 | -0.034702 | 0.038956 | 3.0131 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0080912 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.03 | 843.0249 | 0.843 | 0.196 | 1475428175.1 | 83.2977 | 76.0138 | 23.5554 | 17.2474 | 599.928 | 83.9645 | 0.19953 | 26.8748 | 25.0098 | 999.9 | 9966.88 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.7756 | 0 | 6975.56 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.571 | 0.19659 | -140.7891 | 6992.1615 | 15 | 2 | 516.087 | 661.734 | 25.1624 | 27.569 | 30 | 27.5222 | 27.4948 | 7.0449 | 43.1359 | 0 | 25.1616 | 60.6667 | 17.2491 | 100.744 | 101.261 | -0.031 | 5.465 | 100 | 100
3.519 | 66.0189 | 500 to 0, with leaf | 1988 | 1475428177.6 | 7616.5 | 20161002 11:09:37 | 1475428177.1 | 0.0064725 | -8.1856 | 107.2823 | 9.0297 | 6.755 | 0.50186 | 1.9964 | 0.4409 | 0.28042 | 165.2338 | 25.8657 | 25.0077 | 3.1811 | 55.8297 | 1.9835 | 3.5527 | 1.1977 | -0.034673 | 0.038923 | 3.0111 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0085235 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428177.1 | 80.256 | 72.589 | 23.5656 | 17.245 | 599.936 | 83.9668 | 0.20074 | 26.8737 | 25.0077 | 999.9 | 9958.12 | 0 | 6.9328 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.6182 | 0 | 6972.39 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.5944 | -0.17653 | -133.1032 | 6987.9919 | 15 | 2 | 494.016 | 661.22 | 25.1603 | 27.569 | 30.0001 | 27.5222 | 27.4948 | 6.8156 | 43.1359 | 0 | 25.1515 | 55.6667 | 17.2388 | 100.743 | 101.261 | -0.031 | 5.465 | 100 | 100
3.2967 | 63.8223 | 500 to 0, with leaf | 1989 | 1475428179.6 | 7618.5 | 20161002 11:09:39 | 1475428179.1 | 0.0064806 | -8.4283 | 105.1501 | 8.8502 | 6.4982 | 0.50175 | 2.0031 | 0.44099 | 0.28046 | 165.1838 | 25.8653 | 25.0166 | 3.1828 | 55.8429 | 1.9839 | 3.5527 | 1.1989 | -0.034839 | 0.039109 | 3.0227 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0088141 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.727 | 842.7703 | 0.843 | 0.196 | 1475428179.1 | 77.2059 | 69.2795 | 23.571 | 17.2444 | 600.115 | 83.9679 | 0.1998 | 26.8736 | 25.0166 | 999.9 | 10005.6 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 3.1794 | 0 | 6968.45 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.5992 | -0.50465 | -128.2591 | 6983.9254 | 15 | 2 | 496.569 | 661.033 | 25.1583 | 27.569 | 30.0001 | 27.5211 | 27.4948 | 6.5806 | 43.1359 | 0 | 25.1515 | 50.6667 | 17.2393 | 100.742 | 101.261 | -0.031 | 5.465 | 100 | 100
3.1649 | 61.3399 | 500 to 0, with leaf | 1990 | 1475428181.6 | 7620.5 | 20161002 11:09:41 | 1475428181.1 | 0.0064814 | -8.5796 | 102.6907 | 8.6433 | 6.2432 | 0.50248 | 1.9999 | 0.44147 | 0.28078 | 165.2338 | 25.8633 | 25.0121 | 3.182 | 55.8528 | 1.9842 | 3.5525 | 1.1978 | -0.034758 | 0.039019 | 3.0171 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0089909 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428181.1 | 74.1759 | 66.0767 | 23.5743 | 17.2454 | 599.974 | 83.9681 | 0.19972 | 26.873 | 25.0121 | 999.9 | 9982.5 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.4806 | 0 | 6964.7 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.5952 | -0.39425 | -123.0174 | 6979.0062 | 15 | 2 | 508.997 | 661.676 | 25.1547 | 27.569 | 30.0001 | 27.52 | 27.4939 | 6.4004 | 43.1359 | 0 | 25.1515 | 50.6667 | 17.232 | 100.742 | 101.261 | -0.031 | 5.465 | 100 | 100
2.9167 | 59.3027 | 500 to 0, with leaf | 1991 | 1475428183.6 | 7622.5 | 20161002 11:09:43 | 1475428183.1 | 0.0064864 | -8.848 | 100.6677 | 8.4729 | 5.9911 | 0.50325 | 1.9999 | 0.44207 | 0.28117 | 165.2338 | 25.86 | 25.0097 | 3.1815 | 55.8651 | 1.9845 | 3.5522 | 1.1971 | -0.03476 | 0.039021 | 3.0172 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0093094 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.03 | 843.0249 | 0.843 | 0.196 | 1475428183.1 | 71.1812 | 62.7952 | 23.5776 | 17.2444 | 600.022 | 83.9668 | 0.20036 | 26.8715 | 25.0097 | 999.9 | 9983.12 | 0 | 6.9328 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.717 | 0 | 6961.63 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.5835 | 0.35017 | -117.2841 | 6975.4038 | 15 | 2 | 500.858 | 661.264 | 25.1507 | 27.5685 | 30 | 27.5198 | 27.4927 | 6.1513 | 43.1359 | 0 | 25.1379 | 45.6667 | 17.2302 | 100.743 | 101.261 | -0.031 | 5.465 | 100 | 100
2.6586 | 57.2952 | 500 to 0, with leaf | 1992 | 1475428185.6 | 7624.5 | 20161002 11:09:45 | 1475428185.1 | 0.0064876 | -9.1262 | 98.7365 | 8.3101 | 5.74 | 0.50294 | 2.0078 | 0.44204 | 0.28113 | 165.2321 | 25.8624 | 25.0115 | 3.1819 | 55.8694 | 1.9845 | 3.5521 | 1.1973 | -0.034953 | 0.039238 | 3.0307 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0096394 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.03 | 843.0249 | 0.843 | 0.196 | 1475428185.1 | 68.2 | 59.5158 | 23.5791 | 17.2442 | 599.97 | 83.9651 | 0.1998 | 26.8708 | 25.0115 | 999.9 | 10038.8 | 0 | 6.9328 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.5662 | 0 | 6958.83 | 38.75 | 41.75 | 40.375 | 40.625 | 39.125 | 2.5606 | 0.30852 | -111.1364 | 6971.96 | 15 | 2 | 506.171 | 661.424 | 25.1464 | 27.5674 | 30 | 27.5198 | 27.4925 | 5.9796 | 43.1359 | 0 | 25.1379 | 40.6667 | 17.2268 | 100.744 | 101.262 | -0.031 | 5.465 | 100 | 100
2.3411 | 55.4685 | 500 to 0, with leaf | 1993 | 1475428187.6 | 7626.5 | 20161002 11:09:47 | 1475428187.1 | 0.0064882 | -9.4644 | 96.9782 | 8.1622 | 5.4865 | 0.50316 | 2.0022 | 0.44206 | 0.28116 | 165.2321 | 25.8589 | 25.0118 | 3.1819 | 55.8714 | 1.9845 | 3.5519 | 1.1974 | -0.034816 | 0.039084 | 3.0211 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.01004 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428187.1 | 65.1872 | 56.1457 | 23.5789 | 17.2436 | 599.996 | 83.9648 | 0.20037 | 26.8701 | 25.0118 | 999.9 | 9999.38 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.6288 | 0 | 6955.79 | 38.75 | 41.75 | 40.375 | 40.625 | 39.125 | 2.5723 | 0.29795 | -107.1241 | 6967.5969 | 15 | 2 | 496.207 | 661.074 | 25.1405 | 27.5667 | 30 | 27.5198 | 27.4925 | 5.7955 | 43.1359 | 0 | 25.1274 | 40.6667 | 17.2223 | 100.744 | 101.262 | -0.031 | 5.465 | 100 | 100
1.9721 | 53.7956 | 500 to 0, with leaf | 1994 | 1475428189.6 | 7628.5 | 20161002 11:09:49 | 1475428189.1 | 0.0064884 | -9.8545 | 95.407 | 8.0298 | 5.2311 | 0.50258 | 2.0092 | 0.4418 | 0.28097 | 165.2321 | 25.861 | 25.0156 | 3.1826 | 55.8748 | 1.9845 | 3.5517 | 1.1981 | -0.034988 | 0.039277 | 3.0331 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010503 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.03 | 843.0249 | 0.843 | 0.196 | 1475428189.1 | 62.1533 | 52.7033 | 23.5793 | 17.2447 | 600.076 | 83.9644 | 0.19948 | 26.8691 | 25.0156 | 999.9 | 10048.8 | 0 | 6.9328 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.2377 | 0 | 6953.93 | 38.687 | 41.75 | 40.375 | 40.437 | 39.062 | 2.5738 | -0.2502 | -99.0349 | 6964.5765 | 15 | 2 | 509.259 | 661.215 | 25.1355 | 27.5667 | 30 | 27.5198 | 27.4925 | 5.5635 | 43.1359 | 0 | 25.1274 | 35.6667 | 17.222 | 100.744 | 101.262 | -0.031 | 5.465 | 100 | 100
1.8059 | 51.4248 | 500 to 0, with leaf | 1995 | 1475428191.6 | 7630.5 | 20161002 11:09:51 | 1475428191.1 | 0.0064893 | -10.0406 | 93.0641 | 7.8326 | 4.9758 | 0.50322 | 2.007 | 0.44223 | 0.28126 | 165.2305 | 25.8581 | 25.0112 | 3.1818 | 55.8848 | 1.9847 | 3.5514 | 1.1971 | -0.034933 | 0.039215 | 3.0293 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010724 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.02 | 843.0165 | 0.843 | 0.196 | 1475428191.1 | 59.1209 | 49.4623 | 23.5813 | 17.2439 | 599.895 | 83.964 | 0.19986 | 26.8675 | 25.0112 | 999.9 | 10033.1 | 0 | 6.9328 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.6023 | 0 | 6951.54 | 38.687 | 41.75 | 40.375 | 40.437 | 39.062 | 2.5758 | -0.16734 | -91.5149 | 6961.7342 | 15 | 2 | 516.288 | 661.869 | 25.1307 | 27.5667 | 30 | 27.5198 | 27.4925 | 5.3293 | 43.1359 | 0 | 25.1274 | 30.6667 | 17.217 | 100.744 | 101.262 | -0.031 | 5.465 | 100 | 100
1.6266 | 49.0972 | 500 to 0, with leaf | 1996 | 1475428193.6 | 7632.5 | 20161002 11:09:53 | 1475428193.1 | 0.0064939 | -10.2398 | 90.7761 | 7.6402 | 4.7208 | 0.50396 | 1.9959 | 0.44251 | 0.28146 | 165.2305 | 25.8497 | 25.0125 | 3.182 | 55.8935 | 1.9848 | 3.5511 | 1.1972 | -0.034661 | 0.03891 | 3.0103 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.01096 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.03 | 843.0249 | 0.843 | 0.196 | 1475428193.1 | 56.0896 | 46.2144 | 23.5822 | 17.2417 | 600.024 | 83.9651 | 0.20078 | 26.8659 | 25.0125 | 999.9 | 9955 | 0 | 6.9328 | 999.717 | 0.89998 | 0.10002 | 0 | 2.0177 | 0 | 6945.94 | 38.687 | 41.75 | 40.375 | 40.437 | 39.062 | 2.5936 | -0.30997 | -84.5463 | 6958.0162 | 15 | 2 | 491.33 | 661.004 | 25.1266 | 27.5667 | 30 | 27.5194 | 27.4925 | 5.1543 | 43.1359 | 0 | 25.1139 | 30.6667 | 17.2101 | 100.744 | 101.261 | -0.031 | 5.465 | 100 | 100
1.2544 | 47.4586 | 500 to 0, with leaf | 1997 | 1475428195.6 | 7634.5 | 20161002 11:09:55 | 1475428195.1 | 0.006496 | -10.633 | 89.2045 | 7.5079 | 4.467 | 0.50397 | 1.9986 | 0.44259 | 0.28151 | 165.2322 | 25.8484 | 25.0134 | 3.1822 | 55.8997 | 1.9848 | 3.5507 | 1.1974 | -0.034727 | 0.038984 | 3.0148 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011427 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.02 | 843.0165 | 0.843 | 0.196 | 1475428195.1 | 53.0744 | 42.7875 | 23.5826 | 17.2406 | 600.08 | 83.9653 | 0.19948 | 26.8641 | 25.0134 | 999.9 | 9973.75 | 0 | 6.9328 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.3288 | 0 | 6947.02 | 38.687 | 41.75 | 40.375 | 40.437 | 39.062 | 2.5647 | -0.35123 | -79.9279 | 6955.6435 | 15 | 2 | 503.148 | 661.118 | 25.1218 | 27.5667 | 29.9999 | 27.5182 | 27.4921 | 4.9552 | 43.1359 | 0 | 25.1139 | 25.6667 | 17.2105 | 100.744 | 101.261 | -0.031 | 5.465 | 100 | 100
0.97971 | 45.4925 | 500 to 0, with leaf | 1998 | 1475428197.6 | 7636.5 | 20161002 11:09:57 | 1475428197.1 | 0.006493 | -10.9281 | 87.2784 | 7.3458 | 4.2143 | 0.50427 | 2.0032 | 0.44294 | 0.28173 | 165.2338 | 25.8504 | 25.0055 | 3.1807 | 55.9033 | 1.9848 | 3.5504 | 1.1959 | -0.03484 | 0.03911 | 3.0227 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.011777 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428197.1 | 50.0713 | 39.4655 | 23.5823 | 17.2407 | 599.84 | 83.9651 | 0.20029 | 26.8629 | 25.0055 | 999.9 | 10006.2 | 0 | 6.9328 | 1000.02 | 0.90002 | 0.099985 | 0 | 2.4538 | 0 | 6945.2 | 38.687 | 41.687 | 40.375 | 40.375 | 39.062 | 2.5809 | 0.036499 | -73.7203 | 6953.3277 | 15 | 2 | 498.619 | 661.454 | 25.1153 | 27.5662 | 29.9999 | 27.5175 | 27.491 | 4.7171 | 43.1359 | 0 | 25.103 | 20.6667 | 17.205 | 100.745 | 101.261 | -0.031 | 5.465 | 100 | 100
0.77611 | 43.2599 | 500 to 0, with leaf | 1999 | 1475428199.6 | 7638.5 | 20161002 11:09:59 | 1475428199.1 | 0.006497 | -11.1516 | 85.1031 | 7.1628 | 3.9604 | 0.50442 | 2.0036 | 0.44308 | 0.28181 | 165.2305 | 25.8475 | 25.0071 | 3.181 | 55.9064 | 1.9847 | 3.5501 | 1.1963 | -0.034851 | 0.039123 | 3.0235 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012042 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.01 | 843.0081 | 0.843 | 0.196 | 1475428199.1 | 47.055 | 36.212 | 23.5813 | 17.2392 | 600.162 | 83.9658 | 0.19992 | 26.8613 | 25.0071 | 999.9 | 10009.4 | 0 | 6.9328 | 1000.02 | 0.90002 | 0.099985 | 0 | 2.6937 | 0 | 6943.36 | 38.687 | 41.75 | 40.375 | 40.437 | 39.062 | 2.5623 | -0.026137 | -64.9569 | 6950.5273 | 15 | 2 | 497.186 | 660.788 | 25.1103 | 27.565 | 29.9999 | 27.5175 | 27.4901 | 4.5344 | 43.1359 | 0 | 25.103 | 20.6667 | 17.206 | 100.745 | 101.261 | -0.031 | 5.465 | 100 | 100
0.36151 | 41.7634 | 500 to 0, with leaf | 2000 | 1475428201.6 | 7640.5 | 20161002 11:10:01 | 1475428201.1 | 0.006494 | -11.5876 | 83.7287 | 7.0469 | 3.7064 | 0.50387 | 2.0008 | 0.44257 | 0.28149 | 165.2321 | 25.8445 | 25.0105 | 3.1817 | 55.912 | 1.9846 | 3.5495 | 1.1971 | -0.034781 | 0.039044 | 3.0186 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012559 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.02 | 843.0165 | 0.843 | 0.196 | 1475428201.1 | 44.0374 | 32.7329 | 23.5803 | 17.2378 | 599.846 | 83.9642 | 0.19961 | 26.8585 | 25.0105 | 999.9 | 9989.38 | 0 | 6.9328 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.904 | 0 | 6942.07 | 38.625 | 41.75 | 40.375 | 40.437 | 39.062 | 2.6021 | 0.41036 | -61.8123 | 6948.7158 | 15 | 2 | 515.949 | 661.841 | 25.1054 | 27.5643 | 29.9999 | 27.5175 | 27.4901 | 4.306 | 43.1359 | 0 | 25.103 | 15.6667 | 17.1985 | 100.745 | 101.262 | -0.031 | 5.465 | 100 | 100
0.11935 | 39.6209 | 500 to 0, with leaf | 2001 | 1475428203.6 | 7642.5 | 20161002 11:10:03 | 1475428203.1 | 0.0064984 | -11.8502 | 81.6494 | 6.872 | 3.4486 | 0.50376 | 2.0083 | 0.44269 | 0.28155 | 165.2321 | 25.8452 | 25.0142 | 3.1824 | 55.9214 | 1.9848 | 3.5493 | 1.1976 | -0.034965 | 0.039252 | 3.0315 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.012871 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.03 | 843.0249 | 0.843 | 0.196 | 1475428203.1 | 40.9749 | 29.3925 | 23.5826 | 17.2383 | 600.081 | 83.9639 | 0.2004 | 26.8574 | 25.0142 | 999.9 | 10042.5 | 0 | 6.9328 | 999.724 | 0.89998 | 0.10002 | 0 | 2.5763 | 0 | 6938.25 | 38.687 | 41.75 | 40.375 | 40.437 | 39.062 | 2.6056 | 0.57623 | -58.5703 | 6946.8885 | 15 | 2 | 499.996 | 661.117 | 25.1012 | 27.5643 | 30 | 27.5175 | 27.4901 | 4.0779 | 43.1359 | 0 | 25.095 | 10.6667 | 17.198 | 100.745 | 101.261 | -0.031 | 5.465 | 100 | 100
-0.0016598 | 37.0636 | 500 to 0, with leaf | 2002 | 1475428205.6 | 7644.5 | 20161002 11:10:05 | 1475428205.1 | 0.0065021 | -11.9909 | 79.1338 | 6.6605 | 3.1919 | 0.50409 | 2.0097 | 0.44298 | 0.28173 | 165.2305 | 25.8432 | 25.0145 | 3.1824 | 55.9291 | 1.9849 | 3.549 | 1.1975 | -0.035 | 0.03929 | 3.0339 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.013037 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.03 | 843.0249 | 0.843 | 0.196 | 1475428205.1 | 37.923 | 26.185 | 23.5833 | 17.238 | 600.324 | 83.968 | 0.19902 | 26.8561 | 25.0145 | 999.9 | 10051.9 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.6634 | 0 | 6939.1 | 38.687 | 41.75 | 40.5 | 40.5 | 39.062 | 2.6092 | 0.29736 | -50.7026 | 6944.6604 | 15 | 2 | 525.21 | 660.836 | 25.097 | 27.5643 | 30 | 27.5175 | 27.4901 | 3.9411 | 43.1359 | 0 | 25.095 | 10.6667 | 17.1931 | 100.745 | 101.26 | -0.031 | 5.465 | 100 | 100
-0.31127 | 35.2073 | 500 to 0, with leaf | 2003 | 1475428207.6 | 7646.5 | 20161002 11:10:07 | 1475428207.1 | 0.0064928 | -12.3213 | 77.3385 | 6.5093 | 2.9383 | 0.50448 | 2.0021 | 0.44308 | 0.28182 | 165.2321 | 25.8409 | 25.0033 | 3.1803 | 55.9325 | 1.9848 | 3.5485 | 1.1955 | -0.034812 | 0.039079 | 3.0208 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.013429 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428207.1 | 34.9105 | 22.8042 | 23.5817 | 17.2359 | 599.42 | 83.9661 | 0.2004 | 26.8538 | 25.0033 | 999.9 | 9998.12 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.7471 | 0 | 6938.23 | 38.687 | 41.75 | 40.312 | 40.375 | 39.062 | 2.6272 | 0.46993 | -44.4824 | 6943.1835 | 15 | 2 | 534.226 | 662.612 | 25.0933 | 27.5643 | 30.0001 | 27.5175 | 27.4901 | 3.7071 | 43.1359 | 0 | 25.0822 | 5.6667 | 17.1887 | 100.745 | 101.261 | -0.031 | 5.465 | 100 | 100
-0.53068 | 33.0589 | 500 to 0, with leaf | 2004 | 1475428209.6 | 7648.5 | 20161002 11:10:09 | 1475428209.1 | 0.0065001 | -12.5605 | 75.1809 | 6.3276 | 2.6871 | 0.50557 | 2.0009 | 0.44389 | 0.28234 | 165.2338 | 25.8365 | 24.9985 | 3.1794 | 55.9347 | 1.9847 | 3.5483 | 1.1947 | -0.034782 | 0.039046 | 3.0187 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.013713 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428209.1 | 31.9274 | 19.5771 | 23.5817 | 17.2362 | 600.128 | 83.9633 | 0.20101 | 26.8527 | 24.9985 | 999.9 | 9990 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.8643 | 0 | 6937.17 | 38.687 | 41.75 | 40.312 | 40.437 | 39.062 | 2.6207 | 0.37646 | -42.2622 | 6941.9827 | 15 | 2 | 494.235 | 661.047 | 25.0902 | 27.5643 | 30.0001 | 27.5175 | 27.4901 | 3.4661 | 43.1359 | 0 | 25.0822 | 0.66667 | 17.1919 | 100.745 | 101.262 | -0.031 | 5.465 | 100 | 100
-0.709 | 30.7726 | 500 to 0, with leaf | 2005 | 1475428211.6 | 7650.5 | 20161002 11:10:11 | 1475428211.1 | 0.0064972 | -12.7584 | 73.0604 | 6.1489 | 2.436 | 0.50424 | 2.0017 | 0.44288 | 0.28169 | 165.1848 | 25.8355 | 25.0077 | 3.1811 | 55.9304 | 1.9843 | 3.5479 | 1.1968 | -0.034802 | 0.039068 | 3.0201 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.013948 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0342 | 0.843 | 0.196 | 1475428211.1 | 28.9443 | 16.3799 | 23.5775 | 17.2365 | 600.284 | 83.9633 | 0.1992 | 26.8506 | 25.0077 | 999.9 | 9995.62 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.3338 | 0 | 6936.19 | 38.687 | 41.75 | 40.375 | 40.437 | 39.062 | 2.6176 | -0.013573 | -39.7942 | 6940.2269 | 15 | 2 | 513.329 | 661.07 | 25.0858 | 27.5633 | 30.0001 | 27.5165 | 27.4901 | 3.2848 | 43.1359 | 0 | 25.0822 | 0.66667 | 17.1884 | 100.746 | 101.262 | -0.031 | 5.465 | 100 | 100
-1.0768 | 29.1486 | 500 to 0, with leaf | 2006 | 1475428213.6 | 7652.5 | 20161002 11:10:13 | 1475428213.1 | 0.0064908 | -13.1472 | 71.5763 | 6.0241 | 2.1845 | 0.50338 | 2.0008 | 0.44219 | 0.28125 | 165.2338 | 25.8355 | 25.0107 | 3.1817 | 55.9346 | 1.9842 | 3.5474 | 1.1975 | -0.03478 | 0.039043 | 3.0186 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.014409 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.03 | 843.0249 | 0.843 | 0.196 | 1475428213.1 | 25.9557 | 12.9702 | 23.5759 | 17.2348 | 599.685 | 83.9625 | 0.20066 | 26.8483 | 25.0107 | 999.9 | 9989.38 | 0 | 6.9328 | 1000.03 | 0.90002 | 0.099985 | 0 | 2.579 | 0 | 6935.96 | 38.687 | 41.75 | 40.375 | 40.375 | 39.062 | 2.6199 | -0.33941 | -35.3815 | 6939.1704 | 15 | 2 | 507.864 | 661.69 | 25.0813 | 27.5621 | 30.0001 | 27.5153 | 27.4892 | 0 | 43.1359 | 0 | 25.0792 | 0 | 17.1841 | 100.746 | 101.261 | -0.031 | 5.465 | 100 | 100
-1.9393 | 29.2095 | 500 to 0, with leaf | 2007 | 1475428215.6 | 7654.5 | 20161002 11:10:15 | 1475428215.1 | 0.0064986 | -14.0337 | 71.6672 | 6.0317 | 1.9294 | 0.50408 | 2.0015 | 0.44276 | 0.28161 | 165.2321 | 25.8301 | 25.0106 | 3.1817 | 55.946 | 1.9843 | 3.5468 | 1.1974 | -0.034797 | 0.039063 | 3.0198 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.015461 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.03 | 843.0249 | 0.843 | 0.196 | 1475428215.1 | 22.9246 | 9.0422 | 23.577 | 17.2327 | 600.102 | 83.9621 | 0.20035 | 26.8455 | 25.0106 | 999.9 | 9994.38 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.692 | 0 | 6936.41 | 38.75 | 41.75 | 40.312 | 40.437 | 39.062 | 2.6463 | -0.38071 | -29.0885 | 6938.2654 | 15 | 2 | 490.895 | 660.578 | 25.0787 | 27.562 | 30.0001 | 27.5152 | 27.4881 | 0.0011673 | 43.1359 | 0 | 25.0792 | 5 | 17.1884 | 100.746 | 101.261 | -0.031 | 5.465 | 100 | 100
-1.8845 | 25.8457 | 500 to 0, with leaf | 2008 | 1475428217.6 | 7656.5 | 20161002 11:10:17 | 1475428217.1 | 0.0064962 | -13.9987 | 68.3927 | 5.756 | 1.6559 | 0.50378 | 2.003 | 0.44256 | 0.28148 | 165.2338 | 25.83 | 25.01 | 3.1816 | 55.9471 | 1.9841 | 3.5464 | 1.1974 | -0.034836 | 0.039106 | 3.0225 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.015424 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.738 | 842.7796 | 0.843 | 0.196 | 1475428217.1 | 19.676 | 5.806 | 23.5754 | 17.2328 | 600.038 | 83.9615 | 0.19939 | 26.8437 | 25.01 | 999.9 | 10005.6 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.5761 | 0 | 6936.48 | 38.687 | 41.75 | 40.312 | 40.437 | 39.062 | 2.6034 | -0.32938 | -18.4359 | 6937.375 | 15 | 2 | 509.802 | 661.134 | 25.0767 | 27.562 | 30.0001 | 27.5152 | 27.4879 | 0.011028 | 43.1359 | 0 | 25.0688 | 5 | 17.1855 | 100.745 | 101.263 | -0.031 | 5.465 | 100 | 100
-0.11751 | 16.6922 | 500 to 0, with leaf | 2009 | 1475428219.6 | 7658.5 | 20161002 11:10:19 | 1475428219.1 | 0.0064927 | -12.2396 | 59.2618 | 4.9875 | 1.4021 | 0.50378 | 2.0051 | 0.44262 | 0.28151 | 165.2321 | 25.832 | 25.0052 | 3.1807 | 55.9452 | 1.984 | 3.5464 | 1.1966 | -0.034887 | 0.039164 | 3.0261 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.013332 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428219.1 | 16.6594 | 4.5266 | 23.5746 | 17.2343 | 599.933 | 83.9593 | 0.2006 | 26.8435 | 25.0052 | 999.9 | 10020.6 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.8731 | 0 | 6931.5 | 38.687 | 41.75 | 40.312 | 40.437 | 39.062 | 2.6031 | -0.22862 | -19.9788 | 6936.6523 | 15 | 2 | 492.81 | 660.807 | 25.0748 | 27.562 | 30.0001 | 27.5152 | 27.4879 | 0.12054 | 43.1359 | 0 | 25.0688 | 10 | 17.1835 | 100.746 | 101.263 | -0.031 | 5.465 | 100 | 100
1.9299 | 7.415 | 500 to 0, with leaf | 2010 | 1475428221.6 | 7660.5 | 20161002 11:10:21 | 1475428221.1 | 0.0064948 | -10.1928 | 49.9705 | 4.2055 | 1.2221 | 0.50395 | 2.0069 | 0.4428 | 0.28162 | 165.2338 | 25.8306 | 25.0041 | 3.1805 | 55.9467 | 1.9839 | 3.5461 | 1.1965 | -0.034931 | 0.039213 | 3.0291 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.010904 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428221.1 | 14.5208 | 4.4228 | 23.5736 | 17.2323 | 600.033 | 83.959 | 0.1998 | 26.8421 | 25.0041 | 999.9 | 10033.1 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.4996 | 0 | 6923.64 | 38.687 | 41.75 | 40.375 | 40.437 | 39.062 | 2.5922 | -0.4454 | -35.7966 | 6935.1588 | 15 | 2 | 495.857 | 660.783 | 25.0711 | 27.562 | 30 | 27.5152 | 27.4879 | 0.29244 | 43.1359 | 0 | 25.0688 | 15 | 17.1833 | 100.745 | 101.262 | -0.031 | 5.465 | 100 | 100
3.2007 | 1.6878 | 500 to 0, with leaf | 2011 | 1475428223.6 | 7662.5 | 20161002 11:10:23 | 1475428223.1 | 0.006496 | -8.9223 | 44.2306 | 3.7225 | 1.1127 | 0.50431 | 2.0008 | 0.44291 | 0.28171 | 165.2338 | 25.8251 | 25.0045 | 3.1805 | 55.9583 | 1.9841 | 3.5456 | 1.1965 | -0.034781 | 0.039045 | 3.0186 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0094002 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 999.732 | 842.7746 | 0.843 | 0.196 | 1475428223.1 | 13.2214 | 4.3837 | 23.5749 | 17.2311 | 599.91 | 83.9598 | 0.20038 | 26.8398 | 25.0045 | 999.9 | 9990 | 0 | 6.9328 | 999.736 | 0.89998 | 0.10002 | 0 | 2.9011 | 0 | 6909.94 | 38.687 | 41.75 | 40.375 | 40.437 | 39.062 | 2.5927 | -0.081822 | -82.1046 | 6931.6092 | 15 | 2 | 481.896 | 660.854 | 25.0671 | 27.562 | 30 | 27.5152 | 27.4879 | 0.46168 | 43.1359 | 0 | 25.0635 | 15 | 17.1762 | 100.745 | 101.262 | -0.031 | 5.465 | 100 | 100
4.0013 | -1.892 | 500 to 0, with leaf | 2012 | 1475428225.6 | 7664.5 | 20161002 11:10:25 | 1475428225.1 | 0.0064969 | -8.1216 | 40.6369 | 3.42 | 1.046 | 0.50449 | 2.001 | 0.44306 | 0.2818 | 165.2321 | 25.8222 | 25.0033 | 3.1803 | 55.9668 | 1.9841 | 3.5451 | 1.1962 | -0.034785 | 0.039049 | 3.0189 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0084477 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.03 | 843.0249 | 0.843 | 0.196 | 1475428225.1 | 12.4284 | 4.3887 | 23.5751 | 17.2323 | 600.09 | 83.9595 | 0.1999 | 26.8372 | 25.0033 | 999.9 | 9991.25 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.5277 | 0 | 6905.08 | 38.75 | 41.75 | 40.437 | 40.625 | 39.125 | 2.5734 | -0.44915 | -121.7798 | 6928.1777 | 15 | 2 | 484.016 | 660.504 | 25.0646 | 27.561 | 30 | 27.5148 | 27.4879 | 0.68859 | 43.1359 | 0 | 25.0635 | 20 | 17.1779 | 100.746 | 101.263 | -0.031 | 5.465 | 100 | 100
4.5145 | -4.1919 | 500 to 0, with leaf | 2013 | 1475428227.6 | 7666.5 | 20161002 11:10:27 | 1475428227.1 | 0.0064929 | -7.6084 | 38.3558 | 3.228 | 1.0035 | 0.50419 | 2.0026 | 0.44287 | 0.28168 | 165.2338 | 25.8227 | 25.0012 | 3.1799 | 55.968 | 1.9839 | 3.5447 | 1.196 | -0.034824 | 0.039093 | 3.0217 | 0 | 0 | 0 | 0 | 0 | -1 | - | 0 | 0 | 0 | 0.5 | 0 | 0 | 0 | 0 | -0.0078388 | 0 | - | 0 | 0 | 0 | 0 | 0 | 0 | 1000.04 | 843.0334 | 0.843 | 0.196 | 1475428227.1 | 11.9245 | 4.3927 | 23.5736 | 17.233 | 599.931 | 83.9588 | 0.19941 | 26.8355 | 25.0012 | 999.9 | 10002.5 | 0 | 6.9328 | 1000.04 | 0.90002 | 0.099985 | 0 | 2.5523 | 0 | 6898.15 | 38.75 | 41.75 | 40.375 | 40.562 | 39.125 | 2.5688 | -0.33722 | -166.8311 | 6924.1504 | 15 | 2 | 509.772 | 661.321 | 25.0622 | 27.5598 | 30 | 27.5136 | 27.4879 | 1.0193 | 43.1359 | 0 | 25.0601 | 25 | 17.1754 | 100.745 | 101.263 | -0.031 | 5.465 | 100 | 100

## Data
Atmospheric Pressure: 85.4 kPa
ID: indicated machine (LI1 or LI2) and plant measured (A, B or C), plants were only measured with one LI 6800
Curve #: applicable to replicated standard Aci curve
Curve Type: RACiR (with range of CO2) or standard
Ci: intecellular CO2 concentration, ppm
Photo: net CO2 assimilation rate, umol m-2 s-1
Tleaf: leaf temperature, degrees Celsius
PARi: Irradiance used in umol photons m-2 s-1
E: transpiration rate, mol m-2 s-1
gs: stomatal conductance to CO2, mol m-2 s-1
Cs: CO2 concentration at sample IRGA, ppm
Cr: CO2 concentration at reference IRGA, ppm
Species | ID | Curve# | Curve Type | Ci | Photo | Tleaf | PARi | E | gs | Cs | Cr
Populus deltoides | LI1A |  | RACiR 0 to 500 | 107.1789 | 6.3275 | 24.9551 | 1000.29 | 0.009008 | 0.40196 | 125.528 | 139.376
Populus deltoides | LI1A |  | RACiR 0 to 500 | 108.8986 | 6.4802 | 24.9576 | 999.933 | 0.0090008 | 0.40134 | 127.698 | 141.738
Populus deltoides | LI1A |  | RACiR 0 to 500 | 110.3718 | 6.7371 | 24.9599 | 999.929 | 0.009001 | 0.40112 | 129.863 | 144.196
Populus deltoides | LI1A |  | RACiR 0 to 500 | 111.8219 | 6.7617 | 24.9623 | 999.925 | 0.0090008 | 0.40094 | 131.417 | 145.799
Populus deltoides | LI1A |  | RACiR 0 to 500 | 113.2373 | 7.0401 | 24.9587 | 999.931 | 0.0089957 | 0.4009 | 133.567 | 148.27
Populus deltoides | LI1A |  | RACiR 0 to 500 | 115.09 | 7.1486 | 24.9582 | 999.927 | 0.0089986 | 0.40102 | 135.73 | 150.571
Populus deltoides | LI1A |  | RACiR 0 to 500 | 116.0608 | 7.3641 | 24.959 | 1000.27 | 0.0089939 | 0.40074 | 137.28 | 152.368
Populus deltoides | LI1A |  | RACiR 0 to 500 | 117.649 | 7.5719 | 24.9626 | 999.931 | 0.0089947 | 0.40052 | 139.441 | 154.772
Populus deltoides | LI1A |  | RACiR 0 to 500 | 119.2709 | 7.7623 | 24.9721 | 999.933 | 0.0089915 | 0.39975 | 141.623 | 157.186
Populus deltoides | LI1A |  | RACiR 0 to 500 | 120.362 | 7.9228 | 24.9743 | 1000.27 | 0.00899 | 0.39951 | 143.158 | 158.909
Populus deltoides | LI1A |  | RACiR 0 to 500 | 121.9937 | 8.1318 | 24.9679 | 999.924 | 0.008994 | 0.40009 | 145.323 | 161.314
Populus deltoides | LI1A |  | RACiR 0 to 500 | 123.6496 | 8.3293 | 24.9578 | 999.934 | 0.0089873 | 0.40042 | 147.494 | 163.725
Populus deltoides | LI1A |  | RACiR 0 to 500 | 125.1626 | 8.3371 | 24.9569 | 999.937 | 0.0089878 | 0.40052 | 149.055 | 165.319
Populus deltoides | LI1A |  | RACiR 0 to 500 | 126.7881 | 8.5223 | 24.9622 | 999.932 | 0.0089857 | 0.4001 | 151.21 | 167.695
Populus deltoides | LI1A |  | RACiR 0 to 500 | 128.6097 | 8.6274 | 24.9691 | 999.925 | 0.0089867 | 0.3997 | 153.364 | 169.99
Populus deltoides | LI1A |  | RACiR 0 to 500 | 129.4347 | 8.8922 | 24.971 | 1000.28 | 0.0089864 | 0.39956 | 154.887 | 171.803
Populus deltoides | LI1A |  | RACiR 0 to 500 | 130.887 | 9.1474 | 24.9764 | 999.927 | 0.0089874 | 0.39921 | 157.042 | 174.252
Populus deltoides | LI1A |  | RACiR 0 to 500 | 132.439 | 9.3709 | 24.9778 | 999.928 | 0.0089843 | 0.39895 | 159.212 | 176.686
Populus deltoides | LI1A |  | RACiR 0 to 500 | 133.5077 | 9.5583 | 24.976 | 999.935 | 0.0089875 | 0.39921 | 160.763 | 178.452
Populus deltoides | LI1A |  | RACiR 0 to 500 | 135.3556 | 9.7041 | 24.9694 | 999.931 | 0.0089976 | 0.40007 | 162.966 | 180.838
Populus deltoides | LI1A |  | RACiR 0 to 500 | 136.8031 | 10.0018 | 24.962 | 999.935 | 0.0090104 | 0.40105 | 165.134 | 183.346
Populus deltoides | LI1A |  | RACiR 0 to 500 | 138.4259 | 9.9802 | 24.9578 | 999.934 | 0.0090234 | 0.40182 | 166.689 | 184.902
Populus deltoides | LI1A |  | RACiR 0 to 500 | 139.93 | 10.212 | 24.9586 | 999.944 | 0.0090226 | 0.40151 | 168.833 | 187.322
Populus deltoides | LI1A |  | RACiR 0 to 500 | 141.6022 | 10.3905 | 24.9561 | 999.93 | 0.0090213 | 0.40142 | 170.999 | 189.699
Populus deltoides | LI1A |  | RACiR 0 to 500 | 142.438 | 10.6522 | 24.9577 | 999.937 | 0.0090151 | 0.40093 | 172.548 | 191.537
Populus deltoides | LI1A |  | RACiR 0 to 500 | 144.1567 | 10.7894 | 24.9592 | 1000.28 | 0.008994 | 0.39976 | 174.734 | 193.899
Populus deltoides | LI1A |  | RACiR 0 to 500 | 145.6618 | 11.0024 | 24.9648 | 999.925 | 0.0089755 | 0.39854 | 176.901 | 196.306
Populus deltoides | LI1A |  | RACiR 0 to 500 | 146.7558 | 11.1504 | 24.964 | 999.923 | 0.0089576 | 0.39786 | 178.443 | 198.024
Populus deltoides | LI1A |  | RACiR 0 to 500 | 148.2862 | 11.3991 | 24.9624 | 1000.27 | 0.0089547 | 0.39803 | 180.625 | 200.489
Populus deltoides | LI1A |  | RACiR 0 to 500 | 149.7956 | 11.6654 | 24.9599 | 999.917 | 0.0089596 | 0.39864 | 182.796 | 202.967
Populus deltoides | LI1A |  | RACiR 0 to 500 | 151.2333 | 11.715 | 24.9587 | 999.908 | 0.0089622 | 0.39894 | 184.368 | 204.62
Populus deltoides | LI1A |  | RACiR 0 to 500 | 152.8489 | 11.9237 | 24.9607 | 999.922 | 0.0089695 | 0.39928 | 186.524 | 207.018
Populus deltoides | LI1A |  | RACiR 0 to 500 | 154.778 | 11.9975 | 24.9588 | 999.914 | 0.0089677 | 0.39937 | 188.675 | 209.283
Populus deltoides | LI1A |  | RACiR 0 to 500 | 155.6209 | 12.2602 | 24.9585 | 1000.26 | 0.0089693 | 0.39952 | 190.19 | 211.087
Populus deltoides | LI1A |  | RACiR 0 to 500 | 156.9651 | 12.5691 | 24.9629 | 1000.26 | 0.0089728 | 0.39943 | 192.356 | 213.599
Populus deltoides | LI1A |  | RACiR 0 to 500 | 158.4309 | 12.8458 | 24.9636 | 999.909 | 0.0089743 | 0.39952 | 194.548 | 216.104
Populus deltoides | LI1A |  | RACiR 0 to 500 | 159.5299 | 13.0074 | 24.9663 | 999.907 | 0.0089721 | 0.39923 | 196.107 | 217.855
Populus deltoides | LI1A |  | RACiR 0 to 500 | 161.3692 | 13.1486 | 24.9583 | 999.905 | 0.008974 | 0.39986 | 198.288 | 220.212
Populus deltoides | LI1A |  | RACiR 0 to 500 | 162.5878 | 13.5263 | 24.9553 | 1000.26 | 0.0089713 | 0.39993 | 200.482 | 222.83
Populus deltoides | LI1A |  | RACiR 0 to 500 | 163.9179 | 13.6272 | 24.954 | 999.902 | 0.0089761 | 0.40024 | 202.069 | 224.536
Populus deltoides | LI1A |  | RACiR 0 to 500 | 165.5442 | 13.843 | 24.955 | 999.906 | 0.0089736 | 0.40007 | 204.293 | 227.019
Populus deltoides | LI1A |  | RACiR 0 to 500 | 167.6808 | 13.838 | 24.9605 | 999.899 | 0.0089693 | 0.39951 | 206.519 | 229.277
Populus deltoides | LI1A |  | RACiR 0 to 500 | 168.7254 | 14.03 | 24.9598 | 1000.26 | 0.0089689 | 0.39953 | 208.071 | 231.044
Populus deltoides | LI1A |  | RACiR 0 to 500 | 170.0344 | 14.3459 | 24.9622 | 999.904 | 0.0089666 | 0.39925 | 210.237 | 233.568
Populus deltoides | LI1A |  | RACiR 0 to 500 | 171.7466 | 14.5492 | 24.9602 | 1000.25 | 0.0089798 | 0.39999 | 212.433 | 236
Populus deltoides | LI1A |  | RACiR 0 to 500 | 172.868 | 14.7108 | 24.9653 | 999.904 | 0.0089862 | 0.39993 | 213.998 | 237.766
Populus deltoides | LI1A |  | RACiR 0 to 500 | 174.5948 | 14.9058 | 24.9684 | 999.914 | 0.0090059 | 0.40056 | 216.2 | 240.197
Populus deltoides | LI1A |  | RACiR 0 to 500 | 175.8047 | 15.2833 | 24.9689 | 999.906 | 0.0090074 | 0.40045 | 218.404 | 242.827
Populus deltoides | LI1A |  | RACiR 0 to 500 | 177.2794 | 15.3388 | 24.9632 | 999.913 | 0.0090091 | 0.40077 | 220.018 | 244.517
Populus deltoides | LI1A |  | RACiR 0 to 500 | 178.7023 | 15.6338 | 24.9569 | 999.917 | 0.0090005 | 0.40061 | 222.231 | 247.067
Populus deltoides | LI1A |  | RACiR 0 to 500 | 180.4287 | 15.8077 | 24.9585 | 999.908 | 0.0089996 | 0.40043 | 224.455 | 249.5
Populus deltoides | LI1A |  | RACiR 0 to 500 | 181.1768 | 16.1255 | 24.9598 | 1000.25 | 0.0089992 | 0.40034 | 226.033 | 251.423
Populus deltoides | LI1A |  | RACiR 0 to 500 | 182.7537 | 16.3758 | 24.9646 | 999.915 | 0.0090014 | 0.40016 | 228.3 | 253.985
Populus deltoides | LI1A |  | RACiR 0 to 500 | 184.3269 | 16.6358 | 24.9657 | 999.911 | 0.009003 | 0.40012 | 230.572 | 256.551
Populus deltoides | LI1A |  | RACiR 0 to 500 | 185.5142 | 16.7776 | 24.9698 | 999.9 | 0.0090004 | 0.3997 | 232.192 | 258.344
Populus deltoides | LI1A |  | RACiR 0 to 500 | 187.3168 | 16.9462 | 24.9675 | 999.905 | 0.0090003 | 0.39979 | 234.452 | 260.805
Populus deltoides | LI1A |  | RACiR 0 to 500 | 188.9225 | 17.1815 | 24.9695 | 1000.24 | 0.0089967 | 0.39947 | 236.727 | 263.364
Populus deltoides | LI1A |  | RACiR 0 to 500 | 190.5934 | 17.1457 | 24.9714 | 999.887 | 0.0089991 | 0.39944 | 238.35 | 264.974
Populus deltoides | LI1A |  | RACiR 0 to 500 | 192.4944 | 17.2331 | 24.9793 | 999.9 | 0.0089981 | 0.39889 | 240.582 | 267.334
Populus deltoides | LI1A |  | RACiR 0 to 500 | 194.2533 | 17.38 | 24.9829 | 999.887 | 0.0090008 | 0.39884 | 242.761 | 269.696
Populus deltoides | LI1A |  | RACiR 0 to 500 | 195.0693 | 17.6603 | 24.9819 | 999.901 | 0.0090002 | 0.39886 | 244.303 | 271.543
Populus deltoides | LI1A |  | RACiR 0 to 500 | 196.3687 | 18.0246 | 24.9728 | 999.899 | 0.0089949 | 0.39919 | 246.511 | 274.163
Populus deltoides | LI1A |  | RACiR 0 to 500 | 198.0228 | 18.2567 | 24.9675 | 999.897 | 0.0089983 | 0.3997 | 248.728 | 276.638
Populus deltoides | LI1A |  | RACiR 0 to 500 | 199.1199 | 18.4356 | 24.967 | 999.895 | 0.0089944 | 0.39958 | 250.316 | 278.439
Populus deltoides | LI1A |  | RACiR 0 to 500 | 200.9381 | 18.5643 | 24.9693 | 1000.23 | 0.0089919 | 0.39931 | 252.535 | 280.817
Populus deltoides | LI1A |  | RACiR 0 to 500 | 202.3285 | 18.8375 | 24.9714 | 999.9 | 0.0089794 | 0.39864 | 254.731 | 283.33
Populus deltoides | LI1A |  | RACiR 0 to 500 | 203.8415 | 18.8528 | 24.9719 | 1000.25 | 0.0089799 | 0.3987 | 256.309 | 284.947
Populus deltoides | LI1A |  | RACiR 0 to 500 | 205.3321 | 19.0979 | 24.9758 | 1000.24 | 0.0089782 | 0.39853 | 258.478 | 287.398
Populus deltoides | LI1A |  | RACiR 0 to 500 | 206.9985 | 19.294 | 24.9788 | 999.878 | 0.0089849 | 0.39881 | 260.646 | 289.802
Populus deltoides | LI1A |  | RACiR 0 to 500 | 207.7277 | 19.6224 | 24.9781 | 999.933 | 0.0089863 | 0.39901 | 262.197 | 291.709
Populus deltoides | LI1A |  | RACiR 0 to 500 | 209.3602 | 19.8539 | 24.9764 | 1000.24 | 0.0089903 | 0.39943 | 264.398 | 294.186
Populus deltoides | LI1A |  | RACiR 0 to 500 | 211.3747 | 19.9496 | 24.9748 | 999.905 | 0.0090017 | 0.40015 | 266.607 | 296.521
Populus deltoides | LI1A |  | RACiR 0 to 500 | 212.639 | 20.0621 | 24.9781 | 1000.25 | 0.0090096 | 0.40031 | 268.166 | 298.231
Populus deltoides | LI1A |  | RACiR 0 to 500 | 214.4065 | 20.237 | 24.9849 | 999.91 | 0.0090272 | 0.40064 | 270.38 | 300.653
Populus deltoides | LI1A |  | RACiR 0 to 500 | 215.6953 | 20.576 | 24.9898 | 1000.26 | 0.0090326 | 0.40041 | 272.59 | 303.255
Populus deltoides | LI1A |  | RACiR 0 to 500 | 217.0612 | 20.6281 | 24.9945 | 999.903 | 0.0090315 | 0.39992 | 274.189 | 304.93
Populus deltoides | LI1A |  | RACiR 0 to 500 | 218.5452 | 20.8792 | 24.9961 | 1000.24 | 0.0090304 | 0.39961 | 276.386 | 307.411
Populus deltoides | LI1A |  | RACiR 0 to 500 | 220.4511 | 20.9702 | 24.9937 | 999.897 | 0.009027 | 0.3995 | 278.58 | 309.739
Populus deltoides | LI1A |  | RACiR 0 to 500 | 221.3057 | 21.2659 | 24.9895 | 1000.25 | 0.0090294 | 0.39986 | 280.151 | 311.623
Populus deltoides | LI1A |  | RACiR 0 to 500 | 223.0606 | 21.3843 | 24.9888 | 999.904 | 0.0090121 | 0.3991 | 282.349 | 313.985
Populus deltoides | LI1A |  | RACiR 0 to 500 | 224.4962 | 21.6226 | 24.9923 | 999.901 | 0.0089919 | 0.39795 | 284.582 | 316.478
Populus deltoides | LI1A |  | RACiR 0 to 500 | 225.5406 | 21.7814 | 24.9955 | 1000.25 | 0.008976 | 0.39708 | 286.178 | 318.268
Populus deltoides | LI1A |  | RACiR 0 to 500 | 227.1497 | 22.0264 | 24.9921 | 999.907 | 0.008971 | 0.39725 | 288.419 | 320.781
Populus deltoides | LI1A |  | RACiR 0 to 500 | 228.342 | 22.4337 | 24.9923 | 999.906 | 0.0089731 | 0.39749 | 290.639 | 323.457
Populus deltoides | LI1A |  | RACiR 0 to 500 | 229.8453 | 22.4833 | 24.9931 | 999.91 | 0.0089786 | 0.39784 | 292.251 | 325.143
Populus deltoides | LI1A |  | RACiR 0 to 500 | 231.6746 | 22.6447 | 24.9953 | 999.912 | 0.0089834 | 0.39804 | 294.502 | 327.591
Populus deltoides | LI1A |  | RACiR 0 to 500 | 233.9482 | 22.6288 | 24.9938 | 999.894 | 0.0089858 | 0.39838 | 296.735 | 329.849
Populus deltoides | LI1A |  | RACiR 0 to 500 | 235.1778 | 22.7833 | 24.9904 | 999.911 | 0.0089883 | 0.39878 | 298.323 | 331.62
Populus deltoides | LI1A |  | RACiR 0 to 500 | 236.7377 | 23.0101 | 24.9917 | 999.889 | 0.0089875 | 0.39868 | 300.509 | 334.073
Populus deltoides | LI1A |  | RACiR 0 to 500 | 238.2444 | 23.2669 | 24.997 | 1000.25 | 0.0089922 | 0.39862 | 302.714 | 336.575
Populus deltoides | LI1A |  | RACiR 0 to 500 | 239.4051 | 23.4106 | 25.0017 | 999.897 | 0.0089935 | 0.39839 | 304.306 | 338.339
Populus deltoides | LI1A |  | RACiR 0 to 500 | 240.9876 | 23.6655 | 25.0036 | 1000.24 | 0.0089959 | 0.39844 | 306.565 | 340.892
Populus deltoides | LI1A |  | RACiR 0 to 500 | 242.5931 | 23.9402 | 24.9974 | 999.902 | 0.0089985 | 0.399 | 308.811 | 343.456
Populus deltoides | LI1A |  | RACiR 0 to 500 | 244.0671 | 24.0017 | 24.9972 | 999.941 | 0.0090012 | 0.39922 | 310.44 | 345.17
Populus deltoides | LI1A |  | RACiR 0 to 500 | 245.5494 | 24.278 | 25.0009 | 999.906 | 0.0090023 | 0.39905 | 312.685 | 347.739
Populus deltoides | LI1A |  | RACiR 0 to 500 | 247.4495 | 24.3873 | 25.0028 | 999.904 | 0.0090014 | 0.39891 | 314.929 | 350.119
Populus deltoides | LI1A |  | RACiR 0 to 500 | 248.266 | 24.6927 | 25.0002 | 999.904 | 0.0090006 | 0.39905 | 316.514 | 352.037
Populus deltoides | LI1A |  | RACiR 0 to 500 | 249.7029 | 25.0133 | 24.9993 | 999.902 | 0.0089993 | 0.39909 | 318.788 | 354.679
Populus deltoides | LI1A |  | RACiR 0 to 500 | 251.4073 | 25.2176 | 25.0081 | 999.899 | 0.0090029 | 0.39873 | 321.115 | 357.249
Populus deltoides | LI1A |  | RACiR 0 to 500 | 252.8203 | 25.2836 | 25.0129 | 999.899 | 0.0090033 | 0.39843 | 322.782 | 359.015
Populus deltoides | LI1A |  | RACiR 0 to 500 | 255.0386 | 25.3331 | 25.0182 | 1000.25 | 0.0090211 | 0.3989 | 325.106 | 361.426
Populus deltoides | LI1A |  | RACiR 0 to 500 | 256.8262 | 25.5013 | 25.0224 | 999.899 | 0.0090281 | 0.39883 | 327.379 | 363.915
Populus deltoides | LI1A |  | RACiR 0 to 500 | 258.5042 | 25.4741 | 25.0233 | 1000.26 | 0.0090346 | 0.39893 | 329.014 | 365.547
Populus deltoides | LI1A |  | RACiR 0 to 500 | 260.0447 | 25.7442 | 25.0162 | 999.903 | 0.0090288 | 0.3989 | 331.274 | 368.125
Populus deltoides | LI1A |  | RACiR 0 to 500 | 261.9211 | 25.8926 | 25.0083 | 1000.25 | 0.0090242 | 0.39903 | 333.543 | 370.567
Populus deltoides | LI1A |  | RACiR 0 to 500 | 262.8172 | 26.1655 | 25.0046 | 999.899 | 0.0090215 | 0.39908 | 335.141 | 372.47
Populus deltoides | LI1A |  | RACiR 0 to 500 | 264.5335 | 26.3586 | 25.0037 | 1000.26 | 0.0090197 | 0.39892 | 337.413 | 374.973
Populus deltoides | LI1A |  | RACiR 0 to 500 | 266.3703 | 26.5309 | 24.9991 | 999.895 | 0.0090211 | 0.39924 | 339.67 | 377.438
Populus deltoides | LI1A |  | RACiR 0 to 500 | 267.6849 | 26.6237 | 24.999 | 1000.25 | 0.0090147 | 0.39894 | 341.302 | 379.207
Populus deltoides | LI1A |  | RACiR 0 to 500 | 269.9012 | 26.6536 | 24.9951 | 999.9 | 0.009018 | 0.39939 | 343.563 | 381.517
Populus deltoides | LI1A |  | RACiR 0 to 500 | 271.5078 | 26.871 | 24.9954 | 1000.24 | 0.0090119 | 0.39899 | 345.827 | 384.052
Populus deltoides | LI1A |  | RACiR 0 to 500 | 273.1215 | 26.8632 | 24.9932 | 999.895 | 0.0090134 | 0.39916 | 347.427 | 385.664
Populus deltoides | LI1A |  | RACiR 0 to 500 | 275.046 | 26.9445 | 24.9955 | 1000.24 | 0.0090094 | 0.39873 | 349.679 | 388.029
Populus deltoides | LI1A |  | RACiR 0 to 500 | 276.8296 | 27.0847 | 24.9935 | 999.896 | 0.0090024 | 0.39848 | 351.902 | 390.434
Populus deltoides | LI1A |  | RACiR 0 to 500 | 277.7863 | 27.3344 | 24.9921 | 1000.24 | 0.0090037 | 0.3986 | 353.492 | 392.292
Populus deltoides | LI1A |  | RACiR 0 to 500 | 279.3011 | 27.6303 | 24.987 | 999.892 | 0.0090013 | 0.39881 | 355.749 | 394.89
Populus deltoides | LI1A |  | RACiR 0 to 500 | 281.4243 | 27.6987 | 24.9841 | 999.929 | 0.0090043 | 0.39916 | 358.03 | 397.278
Populus deltoides | LI1A |  | RACiR 0 to 500 | 282.9396 | 27.7558 | 24.9815 | 999.891 | 0.0090071 | 0.3995 | 359.66 | 398.996
Populus deltoides | LI1A |  | RACiR 0 to 500 | 285.0921 | 27.7948 | 24.9833 | 999.895 | 0.0090119 | 0.39968 | 361.93 | 401.34
Populus deltoides | LI1A |  | RACiR 0 to 500 | 286.6451 | 28.0393 | 24.9845 | 999.895 | 0.0090106 | 0.39954 | 364.163 | 403.867
Populus deltoides | LI1A |  | RACiR 0 to 500 | 288.1631 | 28.0699 | 24.9839 | 999.888 | 0.0090105 | 0.39955 | 365.79 | 405.542
Populus deltoides | LI1A |  | RACiR 0 to 500 | 290.0036 | 28.2533 | 24.975 | 999.888 | 0.0090064 | 0.39991 | 368.063 | 408.042
Populus deltoides | LI1A |  | RACiR 0 to 500 | 292.0053 | 28.3389 | 24.9738 | 1000.24 | 0.0090067 | 0.39996 | 370.317 | 410.412
Populus deltoides | LI1A |  | RACiR 0 to 500 | 292.8454 | 28.6268 | 24.9745 | 999.889 | 0.0090071 | 0.39989 | 371.918 | 412.327
Populus deltoides | LI1A |  | RACiR 0 to 500 | 294.4798 | 28.8352 | 24.9751 | 1000.24 | 0.0090021 | 0.39952 | 374.188 | 414.851
Populus deltoides | LI1A |  | RACiR 0 to 500 | 296.3894 | 28.9708 | 24.9728 | 999.896 | 0.0089998 | 0.39949 | 376.489 | 417.32
Populus deltoides | LI1A |  | RACiR 0 to 500 | 297.8509 | 29.0386 | 24.9688 | 999.889 | 0.0089988 | 0.39971 | 378.11 | 419.044
Populus deltoides | LI1A |  | RACiR 0 to 500 | 299.8597 | 29.0971 | 24.9732 | 999.885 | 0.0089958 | 0.39927 | 380.399 | 421.421
Populus deltoides | LI1A |  | RACiR 0 to 500 | 301.3305 | 29.4071 | 24.9692 | 1000.24 | 0.0089901 | 0.39925 | 382.687 | 424.073
Populus deltoides | LI1A |  | RACiR 0 to 500 | 302.7897 | 29.4857 | 24.9692 | 999.891 | 0.0089943 | 0.39945 | 384.341 | 425.819
Populus deltoides | LI1A |  | RACiR 0 to 500 | 304.7803 | 29.6464 | 24.9662 | 999.901 | 0.0090077 | 0.40027 | 386.626 | 428.311
Populus deltoides | LI1A |  | RACiR 0 to 500 | 307.0949 | 29.6214 | 24.9715 | 999.89 | 0.009021 | 0.40051 | 388.892 | 430.602
Populus deltoides | LI1A |  | RACiR 0 to 500 | 308.4736 | 29.7245 | 24.9722 | 999.899 | 0.0090324 | 0.40092 | 390.488 | 432.324
Populus deltoides | LI1A |  | RACiR 0 to 500 | 310.1105 | 29.8974 | 24.9766 | 999.896 | 0.009028 | 0.40026 | 392.732 | 434.797
Populus deltoides | LI1A |  | RACiR 0 to 500 | 312.0413 | 30.0378 | 24.9731 | 1000.23 | 0.0090315 | 0.40049 | 395.016 | 437.233
Populus deltoides | LI1A |  | RACiR 0 to 500 | 313.3335 | 30.1404 | 24.9716 | 999.883 | 0.0090263 | 0.40024 | 396.645 | 439.005
Populus deltoides | LI1A |  | RACiR 0 to 500 | 315.2594 | 30.2778 | 24.9633 | 1000.24 | 0.00902 | 0.4004 | 398.922 | 441.46
Populus deltoides | LI1A |  | RACiR 0 to 500 | 316.5902 | 30.6555 | 24.9591 | 999.901 | 0.0090178 | 0.40046 | 401.223 | 444.179
Populus deltoides | LI1A |  | RACiR 0 to 500 | 318.1955 | 30.6521 | 24.959 | 1000.23 | 0.0090149 | 0.40029 | 402.891 | 445.871
Populus deltoides | LI1A |  | RACiR 0 to 500 | 319.9104 | 30.8397 | 24.9665 | 1000.24 | 0.0090183 | 0.39987 | 405.211 | 448.415
Populus deltoides | LI1A |  | RACiR 0 to 500 | 322.0716 | 30.8393 | 24.9674 | 999.895 | 0.0090075 | 0.39928 | 407.538 | 450.777
Populus deltoides | LI1A |  | RACiR 0 to 500 | 322.9519 | 31.1059 | 24.9679 | 999.887 | 0.0090005 | 0.39895 | 409.18 | 452.711
Populus deltoides | LI1A |  | RACiR 0 to 500 | 324.6251 | 31.3431 | 24.9622 | 1000.25 | 0.0089885 | 0.39889 | 411.494 | 455.294
Populus deltoides | LI1A |  | RACiR 0 to 500 | 326.508 | 31.5159 | 24.9598 | 999.896 | 0.008984 | 0.39898 | 413.834 | 457.843
Populus deltoides | LI1A |  | RACiR 0 to 500 | 327.8482 | 31.6173 | 24.9614 | 1000.24 | 0.0089797 | 0.39875 | 415.509 | 459.66
Populus deltoides | LI1A |  | RACiR 0 to 500 | 329.9634 | 31.718 | 24.9603 | 999.9 | 0.0089843 | 0.3991 | 417.854 | 462.135
Populus deltoides | LI1A |  | RACiR 0 to 500 | 331.6955 | 31.883 | 24.9687 | 999.909 | 0.0089791 | 0.39831 | 420.216 | 464.714
Populus deltoides | LI1A |  | RACiR 0 to 500 | 333.5351 | 31.8122 | 24.9704 | 999.905 | 0.0089838 | 0.39842 | 421.897 | 466.335
Populus deltoides | LI1A |  | RACiR 0 to 500 | 335.8796 | 31.7603 | 24.9712 | 999.904 | 0.0089821 | 0.39831 | 424.188 | 468.624
Populus deltoides | LI1A |  | RACiR 0 to 500 | 338.0188 | 31.8173 | 24.9667 | 999.903 | 0.0089841 | 0.39874 | 426.425 | 470.951
Populus deltoides | LI1A |  | RACiR 0 to 500 | 339.0049 | 32.0761 | 24.9632 | 1000.25 | 0.0089863 | 0.39909 | 428.014 | 472.824
Populus deltoides | LI1A |  | RACiR 0 to 500 | 340.5473 | 32.3149 | 24.9677 | 1000.26 | 0.008983 | 0.39864 | 430.296 | 475.392
Populus deltoides | LI1A |  | RACiR 0 to 500 | 342.3337 | 32.5178 | 24.9656 | 999.912 | 0.0089849 | 0.39885 | 432.59 | 477.918
Populus deltoides | LI1A |  | RACiR 0 to 500 | 343.7507 | 32.6069 | 24.9629 | 999.906 | 0.008983 | 0.39893 | 434.245 | 479.709
Populus deltoides | LI1A |  | RACiR 0 to 500 | 345.8688 | 32.7094 | 24.9588 | 999.903 | 0.0089878 | 0.39945 | 436.558 | 482.144
Populus deltoides | LI1A |  | RACiR 0 to 500 | 347.6758 | 32.8627 | 24.9576 | 1000.25 | 0.0089802 | 0.39915 | 438.855 | 484.652
Populus deltoides | LI1A |  | RACiR 0 to 500 | 349.5126 | 32.7598 | 24.9619 | 999.912 | 0.0089842 | 0.39909 | 440.491 | 486.192
Populus deltoides | LI1A |  | RACiR 0 to 500 | 351.2607 | 32.9093 | 24.9645 | 999.919 | 0.0089807 | 0.39876 | 442.731 | 488.636
Populus deltoides | LI1A |  | RACiR 0 to 500 | 353.4231 | 32.9329 | 24.9658 | 999.91 | 0.0089838 | 0.39883 | 444.989 | 490.941
Populus deltoides | LI1A |  | RACiR 0 to 500 | 355.2471 | 32.8247 | 24.966 | 1000.26 | 0.0089826 | 0.39878 | 446.591 | 492.463
Populus deltoides | LI1A | 1 | Standard | 274.6465 | 29.4767 | 25.0272 | 1000.06 | 0.0076968 | 0.34327 | 367.719 | 400.026
Populus deltoides | LI1A | 1 | Standard | 239.3628 | 26.319 | 24.9913 | 999.706 | 0.0078082 | 0.34819 | 321.236 | 350.063
Populus deltoides | LI1A | 1 | Standard | 205.241 | 22.7693 | 24.9877 | 1000.04 | 0.0079018 | 0.35307 | 275.105 | 300.048
Populus deltoides | LI1A | 1 | Standard | 172.8188 | 18.6136 | 25.009 | 999.99 | 0.0079873 | 0.35611 | 229.601 | 250.048
Populus deltoides | LI1A | 1 | Standard | 140.4173 | 14.4368 | 25.0247 | 999.973 | 0.0080648 | 0.3602 | 184.13 | 200.052
Populus deltoides | LI1A | 1 | Standard | 109.8534 | 9.7314 | 25.0009 | 999.963 | 0.0082154 | 0.36711 | 139.148 | 150.023
Populus deltoides | LI1A | 1 | Standard | 78.7957 | 5.0097 | 24.9997 | 999.928 | 0.0083462 | 0.37197 | 94.2044 | 100
Populus deltoides | LI1A | 1 | Standard | 48.5483 | -0.020543 | 25.0209 | 999.926 | 0.0083919 | 0.37658 | 49.5872 | 49.9828
Populus deltoides | LI1A | 1 | Standard | 285.7512 | 28.8046 | 25.0474 | 999.974 | 0.008598 | 0.38379 | 368.129 | 400.099
Populus deltoides | LI1A | 1 | Standard | 362.3903 | 34.9464 | 24.9857 | 1000.06 | 0.0087531 | 0.39076 | 461.046 | 500.027
Populus deltoides | LI1A | 1 | Standard | 445.6355 | 39.2871 | 25.0264 | 1000.09 | 0.0088903 | 0.39682 | 555.858 | 600.088
Populus deltoides | LI1A | 1 | Standard | 534.2398 | 41.9488 | 24.9935 | 1000.11 | 0.0089929 | 0.40072 | 652.238 | 700.053
Populus deltoides | LI1A | 1 | Standard | 626.8435 | 43.2577 | 25.0326 | 1000.12 | 0.0089721 | 0.40152 | 749.962 | 799.949
Populus deltoides | LI1A | 1 | Standard | 720.0892 | 44.3194 | 25.002 | 1000.11 | 0.0089933 | 0.40154 | 848.023 | 899.967
Populus deltoides | LI1A | 1 | Standard | 814.9075 | 44.9827 | 25.0195 | 999.764 | 0.0089906 | 0.40194 | 946.521 | 1000.01
Populus deltoides | LI1A | 1 | Standard | 908.7961 | 45.8471 | 25.0308 | 999.769 | 0.0089719 | 0.40125 | 1044.9 | 1100.12
Populus deltoides | LI1A | 1 | Standard | 1004.568 | 46.2939 | 25.0073 | 1000.09 | 0.0089953 | 0.40295 | 1143.43 | 1200.01
Populus deltoides | LI1A | 1 | Standard | 1100.4885 | 46.5008 | 25.0063 | 1000.08 | 0.0089421 | 0.40138 | 1242.44 | 1300.05
Populus deltoides | LI1A | 1 | Standard | 1197.3173 | 46.7197 | 25.0208 | 1000.05 | 0.0089313 | 0.40288 | 1341.42 | 1400.12
Populus deltoides | LI1A | 1 | Standard | 1293.5665 | 46.7654 | 25.0056 | 1000.06 | 0.0090226 | 0.40265 | 1440.34 | 1500.1
Populus deltoides | LI1A | 1 | Standard | 1581.4893 | 47.3649 | 25.0361 | 1000.04 | 0.0089404 | 0.39961 | 1737.14 | 1800.03
Populus deltoides | LI1A | 1 | Standard | 1773.244 | 47.8209 | 24.9976 | 1000.03 | 0.008902 | 0.39775 | 1934.97 | 2000.02
Populus deltoides | LI1A |  | RACiR 800 to 300 | 594.8789 | 42.9211 | 25.0082 | 999.902 | 0.0087895 | 0.38656 | 720.871 | 764.049
Populus deltoides | LI1A |  | RACiR 800 to 300 | 592.89 | 42.9935 | 25.011 | 1000.25 | 0.0087771 | 0.38595 | 719.207 | 762.429
Populus deltoides | LI1A |  | RACiR 800 to 300 | 591.7916 | 42.8747 | 25.0095 | 999.9 | 0.008773 | 0.38607 | 717.725 | 760.8
Populus deltoides | LI1A |  | RACiR 800 to 300 | 590.5199 | 42.7725 | 25.0058 | 999.896 | 0.0087645 | 0.3861 | 716.133 | 759.083
Populus deltoides | LI1A |  | RACiR 800 to 300 | 589.2347 | 42.7807 | 24.9997 | 999.886 | 0.0087637 | 0.38664 | 714.66 | 757.602
Populus deltoides | LI1A |  | RACiR 800 to 300 | 588.2079 | 42.6625 | 24.9978 | 999.887 | 0.0087749 | 0.38746 | 713.051 | 755.843
Populus deltoides | LI1A |  | RACiR 800 to 300 | 587.2687 | 42.5007 | 24.9986 | 999.881 | 0.0087776 | 0.38765 | 711.61 | 754.224
Populus deltoides | LI1A |  | RACiR 800 to 300 | 585.4698 | 42.5808 | 24.9981 | 1000.23 | 0.0087794 | 0.38785 | 709.918 | 752.594
Populus deltoides | LI1A |  | RACiR 800 to 300 | 583.9678 | 42.5886 | 25.0009 | 999.893 | 0.0087817 | 0.3878 | 708.423 | 751.08
Populus deltoides | LI1A |  | RACiR 800 to 300 | 582.709 | 42.4403 | 25.0028 | 1000.23 | 0.0087794 | 0.38762 | 706.801 | 749.286
Populus deltoides | LI1A |  | RACiR 800 to 300 | 581.0934 | 42.5098 | 25.0035 | 999.88 | 0.0087802 | 0.38768 | 705.311 | 747.85
Populus deltoides | LI1A |  | RACiR 800 to 300 | 579.7691 | 42.3687 | 25.0073 | 999.886 | 0.0087802 | 0.38748 | 703.654 | 746.023
Populus deltoides | LI1A |  | RACiR 800 to 300 | 578.6062 | 42.2502 | 25.0113 | 1000.23 | 0.0087831 | 0.38743 | 702.175 | 744.401
Populus deltoides | LI1A |  | RACiR 800 to 300 | 576.7133 | 42.3661 | 25.0116 | 999.882 | 0.0087884 | 0.38774 | 700.453 | 742.774
Populus deltoides | LI1A |  | RACiR 800 to 300 | 575.7223 | 42.1876 | 25.0113 | 1000.23 | 0.0087913 | 0.38793 | 698.916 | 741.034
Populus deltoides | LI1A |  | RACiR 800 to 300 | 574.3589 | 42.085 | 25.0136 | 999.885 | 0.0087915 | 0.38781 | 697.291 | 739.281
Populus deltoides | LI1A |  | RACiR 800 to 300 | 572.926 | 42.0486 | 25.0129 | 999.889 | 0.0087899 | 0.3878 | 695.733 | 737.666
Populus deltoides | LI1A |  | RACiR 800 to 300 | 571.5136 | 41.9168 | 25.0137 | 999.884 | 0.0087832 | 0.38745 | 694.045 | 735.818
Populus deltoides | LI1A |  | RACiR 800 to 300 | 570.0482 | 41.8596 | 25.0157 | 999.877 | 0.0087779 | 0.38709 | 692.504 | 734.195
Populus deltoides | LI1A |  | RACiR 800 to 300 | 568.0582 | 41.981 | 25.0152 | 1000.23 | 0.0087777 | 0.38716 | 690.76 | 732.539
Populus deltoides | LI1A |  | RACiR 800 to 300 | 566.7175 | 41.9221 | 25.0168 | 999.877 | 0.0087793 | 0.3872 | 689.226 | 730.93
Populus deltoides | LI1A |  | RACiR 800 to 300 | 565.0418 | 41.935 | 25.0186 | 1000.22 | 0.0087812 | 0.38722 | 687.542 | 729.236
Populus deltoides | LI1A |  | RACiR 800 to 300 | 563.0595 | 42.1155 | 25.0214 | 999.87 | 0.0087839 | 0.38722 | 685.992 | 727.843
Populus deltoides | LI1A |  | RACiR 800 to 300 | 561.6714 | 41.9998 | 25.0256 | 999.876 | 0.0087836 | 0.387 | 684.339 | 726.054
Populus deltoides | LI1A |  | RACiR 800 to 300 | 560.6011 | 41.825 | 25.027 | 1000.22 | 0.0087815 | 0.38681 | 682.843 | 724.365
Populus deltoides | LI1A |  | RACiR 800 to 300 | 558.6107 | 41.9447 | 25.0302 | 999.876 | 0.0087866 | 0.38689 | 681.103 | 722.711
Populus deltoides | LI1A |  | RACiR 800 to 300 | 557.601 | 41.7468 | 25.0321 | 1000.22 | 0.008786 | 0.38674 | 679.599 | 720.984
Populus deltoides | LI1A |  | RACiR 800 to 300 | 556.1603 | 41.6431 | 25.0305 | 1000.23 | 0.0087812 | 0.38663 | 677.882 | 719.144
Populus deltoides | LI1A |  | RACiR 800 to 300 | 553.997 | 41.8487 | 25.0318 | 1000.22 | 0.008775 | 0.38628 | 676.308 | 717.736
Populus deltoides | LI1A |  | RACiR 800 to 300 | 552.424 | 41.7647 | 25.0325 | 999.875 | 0.0087636 | 0.38575 | 674.632 | 715.943
Populus deltoides | LI1A |  | RACiR 800 to 300 | 550.8539 | 41.7592 | 25.0309 | 1000.23 | 0.0087548 | 0.38556 | 673.058 | 714.345
Populus deltoides | LI1A |  | RACiR 800 to 300 | 548.5748 | 41.9644 | 25.0308 | 999.881 | 0.0087477 | 0.38534 | 671.325 | 712.782
Populus deltoides | LI1A |  | RACiR 800 to 300 | 547.2469 | 41.9509 | 25.0265 | 1000.22 | 0.0087465 | 0.38565 | 669.829 | 711.25
Populus deltoides | LI1A |  | RACiR 800 to 300 | 545.4713 | 42.0128 | 25.0247 | 999.876 | 0.0087465 | 0.38582 | 668.119 | 709.588
Populus deltoides | LI1A |  | RACiR 800 to 300 | 543.8977 | 42.0394 | 25.0288 | 999.878 | 0.0087497 | 0.38574 | 666.609 | 708.074
Populus deltoides | LI1A |  | RACiR 800 to 300 | 542.621 | 41.8917 | 25.0342 | 1000.23 | 0.0087549 | 0.38573 | 664.93 | 706.22
Populus deltoides | LI1A |  | RACiR 800 to 300 | 541.5582 | 41.7161 | 25.0358 | 999.875 | 0.0087537 | 0.38562 | 663.414 | 704.515
Populus deltoides | LI1A |  | RACiR 800 to 300 | 539.9572 | 41.7352 | 25.0336 | 1000.22 | 0.0087569 | 0.38596 | 661.722 | 702.816
Populus deltoides | LI1A |  | RACiR 800 to 300 | 538.8809 | 41.598 | 25.0303 | 999.876 | 0.0087576 | 0.38622 | 660.18 | 701.106
Populus deltoides | LI1A |  | RACiR 800 to 300 | 537.314 | 41.5709 | 25.0284 | 1000.21 | 0.0087538 | 0.38618 | 658.513 | 699.397
Populus deltoides | LI1A |  | RACiR 800 to 300 | 535.5547 | 41.6376 | 25.03 | 1000.22 | 0.0087521 | 0.38594 | 656.962 | 697.884
Populus deltoides | LI1A |  | RACiR 800 to 300 | 533.9718 | 41.5794 | 25.0333 | 1000.22 | 0.0087518 | 0.38572 | 655.261 | 696.091
Populus deltoides | LI1A |  | RACiR 800 to 300 | 532.0725 | 41.7096 | 25.0336 | 999.875 | 0.008748 | 0.38551 | 653.719 | 694.67
Populus deltoides | LI1A |  | RACiR 800 to 300 | 530.1304 | 41.8295 | 25.0311 | 999.87 | 0.0087513 | 0.38582 | 651.954 | 692.995
Populus deltoides | LI1A |  | RACiR 800 to 300 | 528.7936 | 41.7827 | 25.0306 | 999.888 | 0.0087536 | 0.386 | 650.41 | 691.379
Populus deltoides | LI1A |  | RACiR 800 to 300 | 527.6247 | 41.6212 | 25.0281 | 999.88 | 0.0087548 | 0.38626 | 648.709 | 689.498
Populus deltoides | LI1A |  | RACiR 800 to 300 | 526.6725 | 41.4582 | 25.0264 | 1000.23 | 0.0087589 | 0.38659 | 647.21 | 687.808
Populus deltoides | LI1A |  | RACiR 800 to 300 | 525.7903 | 41.1793 | 25.0266 | 999.883 | 0.008758 | 0.38653 | 645.596 | 685.886
Populus deltoides | LI1A |  | RACiR 800 to 300 | 524.4789 | 41.0974 | 25.0277 | 999.886 | 0.0087546 | 0.38624 | 644.128 | 684.32
Populus deltoides | LI1A |  | RACiR 800 to 300 | 522.2713 | 41.2535 | 25.0321 | 999.888 | 0.0087532 | 0.3858 | 642.412 | 682.73
Populus deltoides | LI1A |  | RACiR 800 to 300 | 520.5718 | 41.2871 | 25.0344 | 999.876 | 0.0087496 | 0.38542 | 640.877 | 681.204
Populus deltoides | LI1A |  | RACiR 800 to 300 | 518.483 | 41.3989 | 25.0354 | 999.897 | 0.0087472 | 0.38517 | 639.109 | 679.525
Populus deltoides | LI1A |  | RACiR 800 to 300 | 516.8431 | 41.4682 | 25.0335 | 999.89 | 0.0087496 | 0.38539 | 637.547 | 678.004
Populus deltoides | LI1A |  | RACiR 800 to 300 | 515.6426 | 41.3188 | 25.0296 | 1000.24 | 0.0087475 | 0.38553 | 635.879 | 676.166
Populus deltoides | LI1A |  | RACiR 800 to 300 | 514.5963 | 41.1353 | 25.0292 | 999.901 | 0.0087443 | 0.38539 | 634.369 | 674.451
Populus deltoides | LI1A |  | RACiR 800 to 300 | 512.9051 | 41.1507 | 25.0281 | 999.891 | 0.0087446 | 0.38546 | 632.656 | 672.719
Populus deltoides | LI1A |  | RACiR 800 to 300 | 511.6776 | 41.0551 | 25.0263 | 999.896 | 0.0087424 | 0.38546 | 631.147 | 671.092
Populus deltoides | LI1A |  | RACiR 800 to 300 | 510.12 | 40.9963 | 25.0259 | 999.888 | 0.008739 | 0.3853 | 629.444 | 669.307
Populus deltoides | LI1A |  | RACiR 800 to 300 | 508.0585 | 41.2028 | 25.0253 | 1000.24 | 0.0087404 | 0.38533 | 627.869 | 667.91
Populus deltoides | LI1A |  | RACiR 800 to 300 | 506.6437 | 41.1048 | 25.0245 | 999.883 | 0.0087392 | 0.38532 | 626.168 | 666.09
Populus deltoides | LI1A |  | RACiR 800 to 300 | 505.2846 | 41.0394 | 25.0246 | 999.877 | 0.0087365 | 0.38512 | 624.665 | 664.503
Populus deltoides | LI1A |  | RACiR 800 to 300 | 503.1005 | 41.2259 | 25.0247 | 999.921 | 0.0087372 | 0.38505 | 622.943 | 662.932
Populus deltoides | LI1A |  | RACiR 800 to 300 | 501.7561 | 41.1594 | 25.0238 | 999.886 | 0.0087348 | 0.38497 | 621.414 | 661.313
Populus deltoides | LI1A |  | RACiR 800 to 300 | 500.3317 | 41.0731 | 25.0241 | 999.898 | 0.0087316 | 0.3848 | 619.779 | 659.576
Populus deltoides | LI1A |  | RACiR 800 to 300 | 498.6235 | 41.1383 | 25.0272 | 999.884 | 0.0087311 | 0.3845 | 618.295 | 658.123
Populus deltoides | LI1A |  | RACiR 800 to 300 | 497.3018 | 41.0009 | 25.0272 | 999.887 | 0.008729 | 0.3844 | 616.612 | 656.275
Populus deltoides | LI1A |  | RACiR 800 to 300 | 496.1416 | 40.8802 | 25.0255 | 1000.24 | 0.0087262 | 0.38432 | 615.127 | 654.659
Populus deltoides | LI1A |  | RACiR 800 to 300 | 494.3463 | 40.9354 | 25.025 | 999.89 | 0.0087295 | 0.38443 | 613.406 | 652.959
Populus deltoides | LI1A |  | RACiR 800 to 300 | 493.0667 | 40.8261 | 25.0258 | 1000.25 | 0.0087266 | 0.38423 | 611.869 | 651.29
Populus deltoides | LI1A |  | RACiR 800 to 300 | 491.4332 | 40.778 | 25.0272 | 999.899 | 0.0087192 | 0.38378 | 610.202 | 649.548
Populus deltoides | LI1A |  | RACiR 800 to 300 | 489.7807 | 40.81 | 25.0291 | 999.886 | 0.0087136 | 0.38343 | 608.695 | 648.048
Populus deltoides | LI1A |  | RACiR 800 to 300 | 488.5339 | 40.6504 | 25.0292 | 1000.24 | 0.0087104 | 0.38334 | 607.024 | 646.183
Populus deltoides | LI1A |  | RACiR 800 to 300 | 487.0776 | 40.6487 | 25.0279 | 999.891 | 0.0087082 | 0.38338 | 605.514 | 644.649
Populus deltoides | LI1A |  | RACiR 800 to 300 | 485.1282 | 40.7588 | 25.0288 | 1000.24 | 0.0087087 | 0.3834 | 603.804 | 643.032
Populus deltoides | LI1A |  | RACiR 800 to 300 | 483.8247 | 40.7091 | 25.0274 | 999.891 | 0.008708 | 0.38349 | 602.312 | 641.465
Populus deltoides | LI1A |  | RACiR 800 to 300 | 482.2651 | 40.6824 | 25.0249 | 999.887 | 0.0087073 | 0.38359 | 600.613 | 639.712
Populus deltoides | LI1A |  | RACiR 800 to 300 | 480.9393 | 40.6514 | 25.0215 | 999.89 | 0.0087062 | 0.38375 | 599.122 | 638.168
Populus deltoides | LI1A |  | RACiR 800 to 300 | 479.9659 | 40.4224 | 25.0153 | 999.889 | 0.0087043 | 0.38404 | 597.43 | 636.22
Populus deltoides | LI1A |  | RACiR 800 to 300 | 478.6101 | 40.3633 | 25.0134 | 999.892 | 0.0087013 | 0.384 | 595.896 | 634.599
Populus deltoides | LI1A |  | RACiR 800 to 300 | 476.7787 | 40.3853 | 25.0118 | 1000.25 | 0.0086904 | 0.38357 | 594.2 | 632.893
Populus deltoides | LI1A |  | RACiR 800 to 300 | 475.1544 | 40.3971 | 25.01 | 999.902 | 0.0086816 | 0.38326 | 592.651 | 631.331
Populus deltoides | LI1A |  | RACiR 800 to 300 | 473.605 | 40.3535 | 25.0074 | 1000.24 | 0.0086749 | 0.38319 | 590.965 | 629.573
Populus deltoides | LI1A |  | RACiR 800 to 300 | 471.8357 | 40.4483 | 25.0071 | 999.896 | 0.0086715 | 0.3831 | 589.427 | 628.109
Populus deltoides | LI1A |  | RACiR 800 to 300 | 470.8282 | 40.2472 | 25.0071 | 999.893 | 0.0086756 | 0.38344 | 587.768 | 626.215
Populus deltoides | LI1A |  | RACiR 800 to 300 | 469.5977 | 40.1434 | 25.0079 | 999.896 | 0.0086737 | 0.38334 | 586.263 | 624.592
Populus deltoides | LI1A |  | RACiR 800 to 300 | 467.6093 | 40.2578 | 25.0116 | 999.916 | 0.008675 | 0.38319 | 584.58 | 622.994
Populus deltoides | LI1A |  | RACiR 800 to 300 | 466.7262 | 40.0546 | 25.0098 | 999.898 | 0.0086771 | 0.38341 | 583.076 | 621.258
Populus deltoides | LI1A |  | RACiR 800 to 300 | 464.9321 | 40.1065 | 25.0063 | 999.91 | 0.0086727 | 0.38342 | 581.367 | 619.595
Populus deltoides | LI1A |  | RACiR 800 to 300 | 463.3446 | 40.1873 | 25.0042 | 999.904 | 0.0086773 | 0.38377 | 579.855 | 618.134
Populus deltoides | LI1A |  | RACiR 800 to 300 | 462.4018 | 39.9775 | 24.9998 | 999.932 | 0.0086815 | 0.38426 | 578.194 | 616.233
Populus deltoides | LI1A |  | RACiR 800 to 300 | 460.8825 | 39.9587 | 25.0011 | 1000.25 | 0.0086765 | 0.38393 | 576.686 | 614.693
Populus deltoides | LI1A |  | RACiR 800 to 300 | 458.7799 | 40.1068 | 25.0056 | 999.91 | 0.0086798 | 0.3838 | 574.969 | 613.096
Populus deltoides | LI1A |  | RACiR 800 to 300 | 457.6603 | 39.9684 | 25.0065 | 999.916 | 0.0086818 | 0.38378 | 573.466 | 611.424
Populus deltoides | LI1A |  | RACiR 800 to 300 | 455.8277 | 39.9853 | 25.0116 | 1000.25 | 0.0086779 | 0.38328 | 571.784 | 609.747
Populus deltoides | LI1A |  | RACiR 800 to 300 | 454.0994 | 40.0427 | 25.0135 | 999.915 | 0.0086746 | 0.38297 | 570.26 | 608.259
Populus deltoides | LI1A |  | RACiR 800 to 300 | 452.9861 | 39.8654 | 25.0117 | 999.908 | 0.0086739 | 0.38301 | 568.64 | 606.42
Populus deltoides | LI1A |  | RACiR 800 to 300 | 451.7933 | 39.7238 | 25.0132 | 999.946 | 0.0086664 | 0.38258 | 567.166 | 604.79
Populus deltoides | LI1A |  | RACiR 800 to 300 | 450.0571 | 39.7531 | 25.0126 | 999.921 | 0.0086638 | 0.38245 | 565.502 | 603.126
Populus deltoides | LI1A |  | RACiR 800 to 300 | 448.9904 | 39.628 | 25.0087 | 999.916 | 0.0086625 | 0.38266 | 564.017 | 601.489
Populus deltoides | LI1A |  | RACiR 800 to 300 | 447.4888 | 39.6017 | 25.0065 | 1000.25 | 0.0086648 | 0.38295 | 562.326 | 599.749
Populus deltoides | LI1A |  | RACiR 800 to 300 | 445.6452 | 39.7422 | 25.0052 | 1000.25 | 0.0086654 | 0.38308 | 560.773 | 598.309
Populus deltoides | LI1A |  | RACiR 800 to 300 | 444.3063 | 39.6367 | 25.0062 | 999.901 | 0.0086666 | 0.38311 | 559.117 | 596.53
Populus deltoides | LI1A |  | RACiR 800 to 300 | 442.9383 | 39.5815 | 25.0094 | 999.907 | 0.008667 | 0.38296 | 557.616 | 594.956
Populus deltoides | LI1A |  | RACiR 800 to 300 | 441.0578 | 39.656 | 25.0108 | 1000.25 | 0.0086688 | 0.38292 | 555.904 | 593.286
Populus deltoides | LI1A |  | RACiR 800 to 300 | 440.2044 | 39.4405 | 25.0074 | 999.917 | 0.008668 | 0.38313 | 554.399 | 591.549
Populus deltoides | LI1A |  | RACiR 800 to 300 | 438.8561 | 39.3242 | 25.0019 | 999.915 | 0.0086607 | 0.38311 | 552.708 | 589.717
Populus deltoides | LI1A |  | RACiR 800 to 300 | 437.3339 | 39.3363 | 25.0017 | 1000.27 | 0.0086599 | 0.38307 | 551.194 | 588.18
Populus deltoides | LI1A |  | RACiR 800 to 300 | 436.2956 | 39.0747 | 25.0044 | 999.922 | 0.008656 | 0.38273 | 549.538 | 586.236
Populus deltoides | LI1A |  | RACiR 800 to 300 | 434.6194 | 39.0936 | 25.0101 | 999.922 | 0.0086511 | 0.38217 | 548.036 | 584.737
Populus deltoides | LI1A |  | RACiR 800 to 300 | 432.8179 | 39.1489 | 25.0145 | 1000.27 | 0.0086551 | 0.38208 | 546.372 | 583.099
Populus deltoides | LI1A |  | RACiR 800 to 300 | 431.5735 | 39.0946 | 25.0135 | 999.928 | 0.0086572 | 0.38225 | 544.907 | 581.554
Populus deltoides | LI1A |  | RACiR 800 to 300 | 430.0318 | 39.0771 | 25.0082 | 999.928 | 0.00865 | 0.3822 | 543.29 | 579.907
Populus deltoides | LI1A |  | RACiR 800 to 300 | 428.5181 | 39.1052 | 25.006 | 999.922 | 0.0086506 | 0.38229 | 541.788 | 578.407
Populus deltoides | LI1A |  | RACiR 800 to 300 | 427.1703 | 38.9842 | 25.0071 | 1000.27 | 0.0086501 | 0.38218 | 540.123 | 576.586
Populus deltoides | LI1A |  | RACiR 800 to 300 | 425.8928 | 38.8713 | 25.0074 | 999.928 | 0.0086461 | 0.38192 | 538.589 | 574.927
Populus deltoides | LI1A |  | RACiR 800 to 300 | 424.1907 | 38.8851 | 25.0102 | 999.922 | 0.0086512 | 0.38199 | 536.869 | 573.187
Populus deltoides | LI1A |  | RACiR 800 to 300 | 423.2583 | 38.6901 | 25.0087 | 1000.27 | 0.0086506 | 0.38206 | 535.377 | 571.473
Populus deltoides | LI1A |  | RACiR 800 to 300 | 421.8321 | 38.6153 | 25.005 | 999.919 | 0.0086475 | 0.3821 | 533.705 | 569.711
Populus deltoides | LI1A |  | RACiR 800 to 300 | 420.1265 | 38.6818 | 25.0047 | 999.925 | 0.0086451 | 0.38196 | 532.175 | 568.22
Populus deltoides | LI1A |  | RACiR 800 to 300 | 418.5477 | 38.5976 | 25.0069 | 999.913 | 0.0086351 | 0.38135 | 530.505 | 566.433
Populus deltoides | LI1A |  | RACiR 800 to 300 | 416.7527 | 38.6251 | 25.0116 | 1000.27 | 0.00862 | 0.38034 | 529.023 | 564.96
Populus deltoides | LI1A |  | RACiR 800 to 300 | 414.7002 | 38.7282 | 25.0125 | 999.926 | 0.0086101 | 0.37986 | 527.329 | 563.333
Populus deltoides | LI1A |  | RACiR 800 to 300 | 413.3717 | 38.6712 | 25.0098 | 999.925 | 0.0086037 | 0.37979 | 525.833 | 561.745
Populus deltoides | LI1A |  | RACiR 800 to 300 | 411.894 | 38.654 | 25.0037 | 1000.27 | 0.0085985 | 0.38003 | 524.197 | 560.077
Populus deltoides | LI1A |  | RACiR 800 to 300 | 410.4871 | 38.7037 | 24.998 | 1000.26 | 0.0086021 | 0.38063 | 522.716 | 558.619
Populus deltoides | LI1A |  | RACiR 800 to 300 | 409.4181 | 38.5266 | 24.9955 | 999.93 | 0.0086023 | 0.38085 | 521.086 | 556.788
Populus deltoides | LI1A |  | RACiR 800 to 300 | 408.4334 | 38.3457 | 24.9978 | 999.939 | 0.0086018 | 0.38072 | 519.638 | 555.14
Populus deltoides | LI1A |  | RACiR 800 to 300 | 406.7723 | 38.3218 | 25.0022 | 999.965 | 0.0085999 | 0.38033 | 517.987 | 553.436
Populus deltoides | LI1A |  | RACiR 800 to 300 | 405.5488 | 38.2013 | 25.0039 | 999.93 | 0.0085985 | 0.38012 | 516.476 | 551.784
Populus deltoides | LI1A |  | RACiR 800 to 300 | 403.9191 | 38.1845 | 25.0058 | 999.926 | 0.008598 | 0.38 | 514.798 | 550.065
Populus deltoides | LI1A |  | RACiR 800 to 300 | 402.5974 | 38.1585 | 25.0043 | 999.928 | 0.0086018 | 0.38032 | 513.288 | 548.504
Populus deltoides | LI1A |  | RACiR 800 to 300 | 401.2648 | 38.0649 | 25.0013 | 999.929 | 0.0086047 | 0.38067 | 511.577 | 546.677
Populus deltoides | LI1A |  | RACiR 800 to 300 | 400.0375 | 37.993 | 24.9998 | 999.934 | 0.0086062 | 0.38085 | 510.079 | 545.086
Populus deltoides | LI1A |  | RACiR 800 to 300 | 398.4115 | 37.9997 | 25.0019 | 999.932 | 0.00861 | 0.38091 | 508.421 | 543.413
Populus deltoides | LI1A |  | RACiR 800 to 300 | 397.1627 | 37.8735 | 25.0026 | 999.948 | 0.0086045 | 0.38061 | 506.89 | 541.729
Populus deltoides | LI1A |  | RACiR 800 to 300 | 395.7361 | 37.7793 | 25 | 1000.3 | 0.0085943 | 0.38027 | 505.266 | 539.987
Populus deltoides | LI1A |  | RACiR 800 to 300 | 394.2256 | 37.766 | 24.9993 | 1000.29 | 0.008587 | 0.37999 | 503.759 | 538.435
Populus deltoides | LI1A |  | RACiR 800 to 300 | 393.1624 | 37.5554 | 24.9941 | 999.942 | 0.0085818 | 0.38013 | 502.063 | 536.494
Populus deltoides | LI1A |  | RACiR 800 to 300 | 391.5886 | 37.5836 | 24.9913 | 999.942 | 0.008576 | 0.38002 | 500.555 | 535
Populus deltoides | LI1A |  | RACiR 800 to 300 | 389.7048 | 37.6676 | 24.9946 | 999.948 | 0.0085795 | 0.38004 | 498.848 | 533.348
Populus deltoides | LI1A |  | RACiR 800 to 300 | 388.5548 | 37.5502 | 24.9946 | 1000.29 | 0.0085801 | 0.38009 | 497.347 | 531.709
Populus deltoides | LI1A |  | RACiR 800 to 300 | 386.9671 | 37.5234 | 24.9941 | 999.94 | 0.0085774 | 0.37997 | 495.682 | 530
Populus deltoides | LI1A |  | RACiR 800 to 300 | 385.54 | 37.5206 | 24.9951 | 999.945 | 0.0085803 | 0.38008 | 494.187 | 528.476
Populus deltoides | LI1A |  | RACiR 800 to 300 | 384.3447 | 37.3685 | 24.9949 | 999.947 | 0.0085811 | 0.38017 | 492.535 | 526.644
Populus deltoides | LI1A |  | RACiR 800 to 300 | 383.1356 | 37.2662 | 24.9976 | 999.944 | 0.008581 | 0.38003 | 491.065 | 525.052
Populus deltoides | LI1A |  | RACiR 800 to 300 | 381.3167 | 37.3062 | 25.0036 | 999.979 | 0.0085815 | 0.37971 | 489.406 | 523.407
Populus deltoides | LI1A |  | RACiR 800 to 300 | 380.1969 | 37.1662 | 25.0047 | 1000.28 | 0.0085818 | 0.37966 | 487.902 | 521.742
Populus deltoides | LI1A |  | RACiR 800 to 300 | 378.8695 | 37.0609 | 25.0056 | 999.935 | 0.0085822 | 0.37961 | 486.278 | 519.986
Populus deltoides | LI1A |  | RACiR 800 to 300 | 377.0984 | 37.1697 | 25.0053 | 999.947 | 0.0085813 | 0.37957 | 484.768 | 518.565
Populus deltoides | LI1A |  | RACiR 800 to 300 | 375.7718 | 37.0765 | 25.0017 | 999.942 | 0.0085808 | 0.37975 | 483.11 | 516.791
Populus deltoides | LI1A |  | RACiR 800 to 300 | 374.6566 | 36.954 | 25.0023 | 999.939 | 0.0085821 | 0.37976 | 481.641 | 515.175
Populus deltoides | LI1A |  | RACiR 800 to 300 | 372.8094 | 37.0167 | 25.0052 | 1000.3 | 0.0085827 | 0.37962 | 479.959 | 513.532
Populus deltoides | LI1A |  | RACiR 800 to 300 | 371.5997 | 36.9079 | 25.0052 | 999.96 | 0.0085791 | 0.37941 | 478.488 | 511.932
Populus deltoides | LI1A |  | RACiR 800 to 300 | 370.4002 | 36.7739 | 25.005 | 999.955 | 0.0085826 | 0.37962 | 476.849 | 510.132
Populus deltoides | LI1A |  | RACiR 800 to 300 | 368.763 | 36.8345 | 25.0051 | 999.965 | 0.0085813 | 0.37953 | 475.358 | 508.677
Populus deltoides | LI1A |  | RACiR 800 to 300 | 367.8259 | 36.6098 | 25.0028 | 999.958 | 0.0085796 | 0.37962 | 473.775 | 506.852
Populus deltoides | LI1A |  | RACiR 800 to 300 | 366.6531 | 36.5169 | 25.0027 | 999.956 | 0.0085832 | 0.37979 | 472.282 | 505.239
Populus deltoides | LI1A |  | RACiR 800 to 300 | 364.8755 | 36.5728 | 25.0029 | 1000.3 | 0.0085825 | 0.37974 | 470.628 | 503.617
Populus deltoides | LI1A |  | RACiR 800 to 300 | 363.68 | 36.4476 | 25.0044 | 999.956 | 0.00858 | 0.37953 | 469.126 | 501.968
Populus deltoides | LI1A |  | RACiR 800 to 300 | 361.9367 | 36.4154 | 25.0068 | 999.952 | 0.0085684 | 0.37881 | 467.448 | 500.227
Populus deltoides | LI1A |  | RACiR 800 to 300 | 360.2293 | 36.4715 | 25.007 | 999.956 | 0.0085583 | 0.37837 | 465.965 | 498.772
Populus deltoides | LI1A |  | RACiR 800 to 300 | 358.7349 | 36.3729 | 25.0115 | 999.966 | 0.0085492 | 0.37777 | 464.33 | 497.011
Populus deltoides | LI1A |  | RACiR 800 to 300 | 357.6649 | 36.2113 | 25.0154 | 999.954 | 0.0085424 | 0.37736 | 462.913 | 495.41
Populus deltoides | LI1A |  | RACiR 800 to 300 | 356.2093 | 36.1854 | 25.016 | 999.961 | 0.0085455 | 0.37761 | 461.288 | 493.731
Populus deltoides | LI1A |  | RACiR 800 to 300 | 355.2675 | 36.0183 | 25.0148 | 1000.31 | 0.0085456 | 0.37778 | 459.829 | 492.087
Populus deltoides | LI1A |  | RACiR 800 to 300 | 354.3281 | 35.8133 | 25.0103 | 999.968 | 0.0085458 | 0.37813 | 458.221 | 490.251
Populus deltoides | LI1A |  | RACiR 800 to 300 | 352.8782 | 35.8514 | 25.0046 | 999.969 | 0.0085483 | 0.37862 | 456.707 | 488.75
Populus deltoides | LI1A |  | RACiR 800 to 300 | 351.8269 | 35.6692 | 25.0008 | 999.97 | 0.0085489 | 0.37896 | 455.051 | 486.892
Populus deltoides | LI1A |  | RACiR 800 to 300 | 350.5843 | 35.5661 | 25.0054 | 999.963 | 0.0085519 | 0.37886 | 453.537 | 485.253
Populus deltoides | LI1A |  | RACiR 800 to 300 | 348.809 | 35.5723 | 25.0144 | 999.964 | 0.0085536 | 0.37843 | 451.857 | 483.552
Populus deltoides | LI1A |  | RACiR 800 to 300 | 347.7158 | 35.4139 | 25.019 | 999.968 | 0.0085535 | 0.37818 | 450.384 | 481.898
Populus deltoides | LI1A |  | RACiR 800 to 300 | 346.346 | 35.3159 | 25.0201 | 999.968 | 0.0085516 | 0.37803 | 448.76 | 480.155
Populus deltoides | LI1A |  | RACiR 800 to 300 | 344.8465 | 35.3627 | 25.0175 | 999.962 | 0.0085516 | 0.3782 | 447.306 | 478.723
Populus deltoides | LI1A |  | RACiR 800 to 300 | 343.8289 | 35.1761 | 25.0132 | 999.968 | 0.0085511 | 0.37844 | 445.7 | 476.905
Populus deltoides | LI1A |  | RACiR 800 to 300 | 342.6899 | 35.0581 | 25.0132 | 999.966 | 0.0085505 | 0.3784 | 444.23 | 475.302
Populus deltoides | LI1A |  | RACiR 800 to 300 | 340.8513 | 35.1173 | 25.016 | 999.962 | 0.0085502 | 0.37814 | 442.576 | 473.676
Populus deltoides | LI1A |  | RACiR 800 to 300 | 339.7164 | 34.9948 | 25.016 | 999.961 | 0.0085517 | 0.37816 | 441.084 | 472.04
Populus deltoides | LI1A |  | RACiR 800 to 300 | 338.1239 | 34.9637 | 25.0153 | 999.96 | 0.0085501 | 0.37809 | 439.389 | 470.295
Populus deltoides | LI1A |  | RACiR 800 to 300 | 336.4607 | 35.052 | 25.0124 | 999.962 | 0.0085508 | 0.37831 | 437.866 | 468.834
Populus deltoides | LI1A |  | RACiR 800 to 300 | 335.4235 | 34.8601 | 25.012 | 999.962 | 0.0085515 | 0.37843 | 436.261 | 467.015
Populus deltoides | LI1A |  | RACiR 800 to 300 | 334.3656 | 34.7022 | 25.0143 | 999.992 | 0.0085508 | 0.37825 | 434.804 | 465.378
Populus deltoides | LI1A |  | RACiR 800 to 300 | 332.6855 | 34.7133 | 25.0181 | 999.96 | 0.0085505 | 0.37797 | 433.191 | 463.743
Populus deltoides | LI1A |  | RACiR 800 to 300 | 331.4543 | 34.6084 | 25.0219 | 1000.31 | 0.0085493 | 0.37764 | 431.737 | 462.166
Populus deltoides | LI1A |  | RACiR 800 to 300 | 330.0837 | 34.5013 | 25.0251 | 999.956 | 0.0085493 | 0.37741 | 430.109 | 460.411
Populus deltoides | LI1A |  | RACiR 800 to 300 | 328.601 | 34.531 | 25.0264 | 999.958 | 0.0085533 | 0.37754 | 428.643 | 458.95
Populus deltoides | LI1A |  | RACiR 800 to 300 | 327.4499 | 34.3678 | 25.0247 | 999.946 | 0.0085513 | 0.37755 | 427.024 | 457.149
Populus deltoides | LI1A |  | RACiR 800 to 300 | 326.3509 | 34.2504 | 25.0243 | 999.959 | 0.0085507 | 0.37757 | 425.579 | 455.566
Populus deltoides | LI1A |  | RACiR 800 to 300 | 324.925 | 34.2047 | 25.0225 | 999.956 | 0.0085521 | 0.37771 | 423.961 | 453.871
Populus deltoides | LI1A |  | RACiR 800 to 300 | 324.0395 | 33.9838 | 25.0228 | 1000.31 | 0.008549 | 0.37754 | 422.505 | 452.173
Populus deltoides | LI1A |  | RACiR 800 to 300 | 322.9307 | 33.8049 | 25.0239 | 999.96 | 0.0085483 | 0.37745 | 420.916 | 450.386
Populus deltoides | LI1A |  | RACiR 800 to 300 | 321.6724 | 33.7616 | 25.0238 | 999.959 | 0.0085519 | 0.37759 | 419.478 | 448.878
Populus deltoides | LI1A |  | RACiR 800 to 300 | 320.271 | 33.6596 | 25.0267 | 999.956 | 0.0085504 | 0.3773 | 417.846 | 447.118
Populus deltoides | LI1A |  | RACiR 800 to 300 | 318.8421 | 33.6554 | 25.0265 | 999.958 | 0.0085497 | 0.37728 | 416.377 | 445.631
Populus deltoides | LI1A |  | RACiR 800 to 300 | 316.9672 | 33.7462 | 25.0285 | 999.971 | 0.0085524 | 0.37726 | 414.711 | 444.024
Populus deltoides | LI1A |  | RACiR 800 to 300 | 315.7171 | 33.6493 | 25.0298 | 999.964 | 0.0085493 | 0.37698 | 413.243 | 442.44
Populus deltoides | LI1A |  | RACiR 800 to 300 | 314.7095 | 33.4454 | 25.0276 | 1000.32 | 0.0085478 | 0.37703 | 411.652 | 440.622
Populus deltoides | LI1A |  | RACiR 800 to 300 | 313.6596 | 33.3062 | 25.0254 | 1000.31 | 0.0085471 | 0.3771 | 410.184 | 438.99
Populus deltoides | LI1A |  | RACiR 800 to 300 | 312.7592 | 33.0628 | 25.0206 | 999.969 | 0.0085441 | 0.3772 | 408.581 | 437.116
Populus deltoides | LI1A |  | RACiR 800 to 300 | 311.6713 | 32.9431 | 25.0186 | 999.976 | 0.0085454 | 0.37741 | 407.095 | 435.491
Populus deltoides | LI1A |  | RACiR 800 to 300 | 309.658 | 33.0657 | 25.0203 | 999.97 | 0.008544 | 0.37722 | 405.413 | 433.909
Populus deltoides | LI1A |  | RACiR 800 to 300 | 308.2862 | 33.008 | 25.0244 | 999.963 | 0.0085459 | 0.377 | 403.913 | 432.324
Populus deltoides | LI1A |  | RACiR 800 to 300 | 306.7696 | 32.958 | 25.027 | 999.963 | 0.0085455 | 0.37677 | 402.285 | 430.627
Populus deltoides | LI1A |  | RACiR 800 to 300 | 305.0883 | 33.0306 | 25.0269 | 999.976 | 0.0085428 | 0.37658 | 400.806 | 429.194
Populus deltoides | LI1A |  | RACiR 800 to 300 | 304.4852 | 32.6894 | 25.0233 | 999.972 | 0.0085427 | 0.37682 | 399.212 | 427.235
Populus deltoides | LI1A |  | RACiR 800 to 300 | 303.4008 | 32.5561 | 25.0183 | 999.962 | 0.0085356 | 0.37674 | 397.76 | 425.628
Populus deltoides | LI1A |  | RACiR 800 to 300 | 302.0965 | 32.4697 | 25.0161 | 999.987 | 0.0085366 | 0.37696 | 396.139 | 423.894
Populus deltoides | LI1A |  | RACiR 800 to 300 | 301.2307 | 32.2632 | 25.0144 | 999.968 | 0.0085355 | 0.37701 | 394.685 | 422.212
Populus deltoides | LI1A |  | RACiR 800 to 300 | 299.9348 | 32.1524 | 25.013 | 999.97 | 0.0085348 | 0.37706 | 393.05 | 420.443
Populus deltoides | LI1A |  | RACiR 800 to 300 | 298.2782 | 32.2344 | 25.016 | 999.963 | 0.0085383 | 0.37704 | 391.582 | 419.032
Populus deltoides | LI1A |  | RACiR 800 to 300 | 296.8825 | 32.1347 | 25.0191 | 999.97 | 0.0085379 | 0.37684 | 389.938 | 417.269
Populus deltoides | LI1A |  | RACiR 800 to 300 | 295.7225 | 32.0207 | 25.0229 | 999.954 | 0.0085376 | 0.37659 | 388.507 | 415.702
Populus deltoides | LI1A |  | RACiR 800 to 300 | 294.0212 | 32.0291 | 25.0286 | 999.952 | 0.0085405 | 0.37635 | 386.851 | 414.026
Populus deltoides | LI1A |  | RACiR 800 to 300 | 292.8738 | 31.9115 | 25.0303 | 999.962 | 0.0085394 | 0.37617 | 385.405 | 412.446
Populus deltoides | LI1A |  | RACiR 800 to 300 | 291.5938 | 31.791 | 25.031 | 1000.31 | 0.0085394 | 0.37607 | 383.796 | 410.691
Populus deltoides | LI1A |  | RACiR 800 to 300 | 290.1069 | 31.8179 | 25.03 | 999.963 | 0.0085397 | 0.37617 | 382.324 | 409.223
Populus deltoides | LI1A |  | RACiR 800 to 300 | 289.0608 | 31.6356 | 25.0264 | 999.962 | 0.0085382 | 0.37624 | 380.744 | 407.437
Populus deltoides | LI1A |  | RACiR 800 to 300 | 288.234 | 31.4325 | 25.0231 | 999.962 | 0.0085373 | 0.3764 | 379.313 | 405.781
Populus deltoides | LI1A |  | RACiR 800 to 300 | 286.5444 | 31.4547 | 25.0224 | 999.965 | 0.0085329 | 0.37619 | 377.691 | 404.155
Populus deltoides | LI1A |  | RACiR 800 to 300 | 285.543 | 31.2944 | 25.0229 | 999.97 | 0.0085329 | 0.37611 | 376.255 | 402.534
Populus deltoides | LI1A |  | RACiR 800 to 300 | 284.3833 | 31.1354 | 25.0208 | 999.96 | 0.0085317 | 0.3762 | 374.62 | 400.721
Populus deltoides | LI1A |  | RACiR 800 to 300 | 282.7938 | 31.1908 | 25.0201 | 999.961 | 0.0085315 | 0.37617 | 373.149 | 399.281
Populus deltoides | LI1A |  | RACiR 800 to 300 | 281.735 | 31.0108 | 25.0179 | 1000.31 | 0.0085352 | 0.37647 | 371.513 | 397.44
Populus deltoides | LI1A |  | RACiR 800 to 300 | 280.5368 | 30.9258 | 25.0158 | 999.967 | 0.008535 | 0.37656 | 370.036 | 395.863
Populus deltoides | LI1A |  | RACiR 800 to 300 | 278.6385 | 31.012 | 25.02 | 999.955 | 0.0085345 | 0.37625 | 368.4 | 394.284
Populus deltoides | LI1A |  | RACiR 800 to 300 | 277.7102 | 30.8175 | 25.0223 | 999.961 | 0.0085334 | 0.37601 | 366.984 | 392.651
Populus deltoides | LI1A |  | RACiR 800 to 300 | 276.5919 | 30.6315 | 25.0217 | 999.957 | 0.0085284 | 0.37576 | 365.396 | 390.855
Populus deltoides | LI1A |  | RACiR 800 to 300 | 275.4458 | 30.5324 | 25.0213 | 999.955 | 0.008529 | 0.37576 | 363.957 | 389.291
Populus deltoides | LI1A |  | RACiR 800 to 300 | 274.6338 | 30.2511 | 25.0179 | 999.965 | 0.008527 | 0.3758 | 362.358 | 387.387
Populus deltoides | LI1A |  | RACiR 800 to 300 | 273.6083 | 30.1117 | 25.0142 | 1000.3 | 0.0085275 | 0.37603 | 360.88 | 385.749
Populus deltoides | LI1A |  | RACiR 800 to 300 | 271.999 | 30.1132 | 25.0124 | 999.955 | 0.0085242 | 0.37596 | 359.252 | 384.094
Populus deltoides | LI1A |  | RACiR 800 to 300 | 270.9002 | 29.9557 | 25.011 | 999.954 | 0.0085121 | 0.37544 | 357.816 | 382.476
Populus deltoides | LI1A |  | RACiR 800 to 300 | 269.4641 | 29.8774 | 25.0081 | 999.956 | 0.0085012 | 0.37507 | 356.212 | 380.766
Populus deltoides | LI1A |  | RACiR 800 to 300 | 267.7803 | 29.9397 | 25.0084 | 999.957 | 0.0084907 | 0.37456 | 354.769 | 379.362
Populus deltoides | LI1A |  | RACiR 800 to 300 | 266.4772 | 29.8079 | 25.011 | 999.956 | 0.008487 | 0.3742 | 353.162 | 377.594
Populus deltoides | LI1A |  | RACiR 800 to 300 | 265.3889 | 29.6735 | 25.011 | 1000.31 | 0.0084864 | 0.37418 | 351.689 | 375.968
Populus deltoides | LI1A |  | RACiR 800 to 300 | 263.8809 | 29.6437 | 25.0118 | 999.945 | 0.0084886 | 0.37423 | 350.056 | 374.282
Populus deltoides | LI1A |  | RACiR 800 to 300 | 262.8702 | 29.4951 | 25.0104 | 999.951 | 0.0084903 | 0.3744 | 348.582 | 372.635
Populus deltoides | LI1A |  | RACiR 800 to 300 | 261.6797 | 29.3605 | 25.0042 | 999.957 | 0.0084857 | 0.3746 | 346.951 | 370.849
Populus deltoides | LI1A |  | RACiR 800 to 300 | 260.2695 | 29.3711 | 25.0007 | 999.964 | 0.008487 | 0.37489 | 345.473 | 369.355
Populus deltoides | LI1A |  | RACiR 800 to 300 | 259.1222 | 29.1842 | 25.0008 | 999.95 | 0.0084819 | 0.37467 | 343.84 | 367.507
Populus deltoides | LI1A |  | RACiR 800 to 300 | 257.9163 | 29.05 | 25.0037 | 999.947 | 0.0084713 | 0.37406 | 342.375 | 365.888
Populus deltoides | LI1A |  | RACiR 800 to 300 | 256.3584 | 29.0158 | 25.0055 | 999.965 | 0.0084658 | 0.37376 | 340.752 | 364.2
Populus deltoides | LI1A |  | RACiR 800 to 300 | 255.1624 | 28.9005 | 25.0082 | 999.953 | 0.0084557 | 0.37324 | 339.329 | 362.638
Populus deltoides | LI1A |  | RACiR 800 to 300 | 253.7127 | 28.8397 | 25.0091 | 999.949 | 0.0084463 | 0.37287 | 337.757 | 360.984
Populus deltoides | LI1A |  | RACiR 800 to 300 | 252.335 | 28.8608 | 25.0068 | 999.953 | 0.0084497 | 0.37325 | 336.32 | 359.543
Populus deltoides | LI1A |  | RACiR 800 to 300 | 251.3391 | 28.6502 | 25.0103 | 999.952 | 0.0084519 | 0.37326 | 334.73 | 357.719
Populus deltoides | LI1A |  | RACiR 800 to 300 | 250.5293 | 28.4176 | 25.0149 | 1000.3 | 0.0084535 | 0.37317 | 333.294 | 356.035
Populus deltoides | LI1A |  | RACiR 800 to 300 | 249.2086 | 28.3353 | 25.0172 | 999.953 | 0.0084599 | 0.37336 | 331.682 | 354.31
Populus deltoides | LI1A |  | RACiR 800 to 300 | 248.2389 | 28.158 | 25.0189 | 999.959 | 0.0084591 | 0.37325 | 330.235 | 352.667
Populus deltoides | LI1A |  | RACiR 800 to 300 | 246.7517 | 28.1294 | 25.018 | 999.941 | 0.0084602 | 0.37335 | 328.614 | 350.995
Populus deltoides | LI1A |  | RACiR 800 to 300 | 245.0859 | 28.2589 | 25.0125 | 999.943 | 0.0084622 | 0.3738 | 327.162 | 349.651
Populus deltoides | LI1A |  | RACiR 800 to 300 | 244.1617 | 28.0539 | 25.0087 | 999.95 | 0.0084645 | 0.37422 | 325.571 | 347.834
Populus deltoides | LI1A |  | RACiR 800 to 300 | 243.3579 | 27.8587 | 25.0082 | 999.944 | 0.0084674 | 0.37447 | 324.17 | 346.217
Populus deltoides | LI1A |  | RACiR 800 to 300 | 242.2925 | 27.7047 | 25.0079 | 999.944 | 0.0084706 | 0.37471 | 322.614 | 344.482
Populus deltoides | LI1A |  | RACiR 800 to 300 | 241.4816 | 27.4613 | 25.0102 | 999.948 | 0.0084682 | 0.37442 | 321.187 | 342.789
Populus deltoides | LI1A |  | RACiR 800 to 300 | 240.1332 | 27.3356 | 25.0159 | 999.935 | 0.0084638 | 0.37382 | 319.594 | 341.05
Populus deltoides | LI1A |  | RACiR 800 to 300 | 238.7556 | 27.3041 | 25.0174 | 999.939 | 0.0084607 | 0.37353 | 318.161 | 339.561
Populus deltoides | LI1A |  | RACiR 800 to 300 | 237.8941 | 27.0456 | 25.0153 | 999.936 | 0.0084595 | 0.37357 | 316.569 | 337.686
Populus deltoides | LI1A |  | RACiR 800 to 300 | 236.6413 | 26.9976 | 25.0138 | 999.93 | 0.0084605 | 0.37367 | 315.138 | 336.188
Populus deltoides | LI1A |  | RACiR 800 to 300 | 235.3454 | 26.9316 | 25.0091 | 999.936 | 0.0084603 | 0.37396 | 313.572 | 334.528
Populus deltoides | LI1A |  | RACiR 800 to 300 | 234.434 | 26.7671 | 25.0046 | 1000.29 | 0.0084625 | 0.37433 | 312.118 | 332.89
Populus deltoides | LI1A |  | RACiR 800 to 300 | 233.0638 | 26.6979 | 25.0035 | 999.937 | 0.0084625 | 0.37435 | 310.525 | 331.204
Populus deltoides | LI1A |  | RACiR 800 to 300 | 231.7859 | 26.6375 | 25.0043 | 999.94 | 0.0084614 | 0.37421 | 309.085 | 329.68
Populus deltoides | LI1A |  | RACiR 800 to 300 | 230.7397 | 26.4239 | 25.0054 | 999.925 | 0.0084608 | 0.37404 | 307.472 | 327.831
Populus deltoides | LI1A |  | RACiR 800 to 300 | 229.7605 | 26.2365 | 25.0085 | 999.933 | 0.0084586 | 0.37368 | 306.036 | 326.185
Populus deltoides | LI1A |  | RACiR 800 to 300 | 228.1006 | 26.237 | 25.0117 | 1000.27 | 0.0084564 | 0.37333 | 304.41 | 324.533
Populus deltoides | LI1A |  | RACiR 800 to 300 | 227.2226 | 26.0327 | 25.0119 | 999.929 | 0.0084528 | 0.37313 | 302.996 | 322.896
Populus deltoides | LI1A |  | RACiR 800 to 300 | 225.8969 | 25.9331 | 25.0118 | 999.921 | 0.0084499 | 0.37297 | 301.401 | 321.175
Populus deltoides | LI1A |  | RACiR 800 to 300 | 224.5589 | 25.9068 | 25.0097 | 999.925 | 0.0084494 | 0.37302 | 299.95 | 319.675
Populus deltoides | LI1A |  | RACiR 800 to 300 | 223.5225 | 25.708 | 25.0086 | 999.921 | 0.0084477 | 0.37298 | 298.359 | 317.866
Populus deltoides | LI1A |  | RACiR 800 to 300 | 222.4391 | 25.5936 | 25.0072 | 999.921 | 0.0084477 | 0.37305 | 296.926 | 316.293
Populus deltoides | LI1A |  | RACiR 800 to 300 | 220.9654 | 25.5751 | 25.0037 | 999.926 | 0.008447 | 0.3732 | 295.338 | 314.663
Populus deltoides | LI1A |  | RACiR 800 to 300 | 219.857 | 25.4725 | 25.0019 | 1000.27 | 0.0084445 | 0.37318 | 293.929 | 313.131
Populus deltoides | LI1A |  | RACiR 800 to 300 | 218.6876 | 25.3506 | 24.9996 | 999.93 | 0.0084468 | 0.37341 | 292.358 | 311.413
Populus deltoides | LI1A |  | RACiR 800 to 300 | 217.3002 | 25.342 | 24.9996 | 999.925 | 0.0084465 | 0.37333 | 290.93 | 309.955
Populus deltoides | LI1A |  | RACiR 800 to 300 | 215.666 | 25.3716 | 25.0012 | 999.927 | 0.0084451 | 0.37317 | 289.371 | 308.405
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 348.3693 | 30.9649 | 25.1674 | 1000 | 0.0058634 | 0.36838 | 438.689 | 463.226
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 346.013 | 30.7753 | 25.1497 | 999.984 | 0.0058511 | 0.36888 | 435.641 | 459.95
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 343.7376 | 30.5618 | 25.133 | 999.992 | 0.0058397 | 0.36941 | 432.605 | 456.66
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 341.2813 | 30.3644 | 25.1245 | 999.983 | 0.0058313 | 0.36945 | 429.553 | 453.364
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 339.1625 | 30.0217 | 25.1134 | 999.984 | 0.0058216 | 0.36957 | 426.426 | 449.852
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 336.622 | 29.8246 | 25.1021 | 999.988 | 0.0058098 | 0.36961 | 423.287 | 446.476
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 333.6137 | 29.8706 | 25.082 | 1000.02 | 0.005798 | 0.37038 | 420.162 | 443.353
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 331.2155 | 29.6265 | 25.0702 | 999.979 | 0.0057866 | 0.37046 | 417.032 | 439.938
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 328.7757 | 29.3769 | 25.0619 | 1000.01 | 0.0057755 | 0.37021 | 413.92 | 436.535
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 325.6832 | 29.3844 | 25.0493 | 1000.01 | 0.0057621 | 0.3702 | 410.789 | 433.368
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 322.8611 | 29.2736 | 25.0358 | 1000.01 | 0.0057472 | 0.3701 | 407.63 | 430.055
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 320.2896 | 29.1139 | 25.0229 | 999.974 | 0.0057349 | 0.37024 | 404.539 | 426.759
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 317.4672 | 29.0379 | 25.0142 | 999.985 | 0.0057257 | 0.37025 | 401.453 | 423.558
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 315.2672 | 28.7532 | 25.0044 | 999.988 | 0.0057194 | 0.37059 | 398.361 | 420.143
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 313.0289 | 28.4812 | 24.9923 | 999.991 | 0.0057108 | 0.37093 | 395.264 | 416.732
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 310.0773 | 28.4437 | 24.9835 | 999.99 | 0.0057038 | 0.37114 | 392.112 | 413.501
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 307.9865 | 28.1236 | 24.9691 | 999.96 | 0.0056965 | 0.37185 | 388.956 | 409.985
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 305.8353 | 27.8493 | 24.9596 | 999.947 | 0.0056969 | 0.37297 | 385.786 | 406.503
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 303.3026 | 27.8789 | 24.9566 | 999.974 | 0.0057231 | 0.37628 | 382.61 | 403.325
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 301.7463 | 27.7199 | 24.9573 | 999.983 | 0.0057725 | 0.38186 | 379.487 | 400.031
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 300.3434 | 27.5385 | 24.9597 | 999.974 | 0.0058298 | 0.38804 | 376.394 | 396.743
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 298.3054 | 27.4464 | 24.9612 | 1000.29 | 0.0058686 | 0.39228 | 373.295 | 393.522
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 295.9967 | 27.1486 | 24.9784 | 999.988 | 0.0058809 | 0.39219 | 370.214 | 390.115
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 293.1624 | 26.9601 | 24.9939 | 999.973 | 0.0058811 | 0.3909 | 367.098 | 386.77
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 289.5417 | 26.9957 | 25.011 | 999.965 | 0.0058727 | 0.3886 | 363.949 | 383.616
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 286.4937 | 26.7905 | 25.0314 | 999.979 | 0.0058644 | 0.38597 | 360.823 | 380.246
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 283.62 | 26.6034 | 25.0423 | 999.971 | 0.0058553 | 0.38408 | 357.775 | 376.968
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 280.5363 | 26.485 | 25.054 | 999.976 | 0.0058429 | 0.38188 | 354.751 | 373.783
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 278.0592 | 26.1563 | 25.062 | 999.975 | 0.0058323 | 0.3803 | 351.666 | 370.331
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 275.437 | 25.9128 | 25.0685 | 999.973 | 0.0058269 | 0.37933 | 348.541 | 366.923
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 272.443 | 25.8469 | 25.0723 | 1000.27 | 0.0058208 | 0.3786 | 345.462 | 363.736
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 269.9221 | 25.6242 | 25.0755 | 999.965 | 0.0058161 | 0.37799 | 342.423 | 360.438
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 267.424 | 25.3953 | 25.0798 | 999.964 | 0.0058146 | 0.37752 | 339.365 | 357.115
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 264.6214 | 25.2825 | 25.082 | 1000.26 | 0.0058118 | 0.3772 | 336.277 | 353.872
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 262.1338 | 25.0081 | 25.0903 | 999.958 | 0.005806 | 0.37612 | 333.218 | 350.5
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 258.8154 | 24.7163 | 25.1028 | 999.948 | 0.0058004 | 0.37472 | 329.327 | 346.269
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 255.9506 | 24.8738 | 25.1144 | 1000.26 | 0.0057986 | 0.37364 | 327.047 | 344.121
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 253.7699 | 24.5643 | 25.1189 | 999.958 | 0.005796 | 0.37314 | 324.089 | 340.817
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 252.2868 | 24.0187 | 25.118 | 1000.26 | 0.0057966 | 0.37334 | 321.072 | 337.215
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 249.8124 | 23.7764 | 25.1195 | 999.927 | 0.0057939 | 0.37314 | 317.939 | 333.799
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 247.0472 | 23.6381 | 25.1201 | 1000.25 | 0.0057881 | 0.37275 | 314.825 | 330.509
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 244.1888 | 23.5367 | 25.1209 | 999.961 | 0.0057831 | 0.37233 | 311.72 | 327.262
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 241.1891 | 23.4876 | 25.125 | 1000.26 | 0.0057783 | 0.37172 | 308.649 | 324.106
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 239.1024 | 23.1149 | 25.1282 | 999.965 | 0.0057719 | 0.37095 | 305.653 | 320.698
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 236.7687 | 22.8452 | 25.1269 | 999.957 | 0.0057651 | 0.37055 | 302.616 | 317.352
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 234.0547 | 22.7388 | 25.126 | 999.909 | 0.0057631 | 0.37061 | 299.558 | 314.15
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 231.7931 | 22.4048 | 25.1324 | 999.953 | 0.0057586 | 0.36981 | 296.491 | 310.711
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 229.3724 | 22.1314 | 25.1354 | 1000.24 | 0.0057534 | 0.36926 | 293.379 | 307.285
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 226.447 | 22.0612 | 25.1365 | 999.939 | 0.0057493 | 0.36895 | 290.267 | 304.064
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 224.1779 | 21.8151 | 25.1279 | 999.935 | 0.0057452 | 0.36946 | 287.2 | 300.714
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 221.7481 | 21.5985 | 25.1245 | 1000.25 | 0.005741 | 0.36944 | 284.142 | 297.399
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 218.9545 | 21.4705 | 25.1294 | 999.933 | 0.0057326 | 0.36837 | 281.131 | 294.222
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 216.9477 | 21.0978 | 25.128 | 999.907 | 0.0057244 | 0.36791 | 278.144 | 290.826
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 214.6539 | 20.8314 | 25.127 | 999.935 | 0.0057216 | 0.36786 | 275.091 | 287.467
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 211.9331 | 20.7111 | 25.125 | 999.897 | 0.005715 | 0.3676 | 272.037 | 284.254
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 209.7138 | 20.4371 | 25.1166 | 1000.27 | 0.0057067 | 0.3677 | 269.009 | 280.913
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 207.2501 | 20.231 | 25.1167 | 999.942 | 0.0057006 | 0.36724 | 266.013 | 277.674
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 204.7864 | 20.024 | 25.1201 | 1000.25 | 0.0056958 | 0.36657 | 263.047 | 274.464
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 202.8026 | 19.6808 | 25.1141 | 999.946 | 0.0056913 | 0.36676 | 260.056 | 271.091
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 200.5314 | 19.4179 | 25.1051 | 999.957 | 0.0056803 | 0.3667 | 257.029 | 267.762
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 197.9702 | 19.2686 | 25.096 | 999.948 | 0.0056711 | 0.36681 | 253.995 | 264.539
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 195.6191 | 19.0102 | 25.0945 | 1000.25 | 0.0056638 | 0.36634 | 250.964 | 261.212
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 193.3845 | 18.7082 | 25.0967 | 999.937 | 0.0056571 | 0.36558 | 247.973 | 257.882
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 190.9692 | 18.453 | 25.1 | 1000.25 | 0.0056505 | 0.36479 | 244.931 | 254.545
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 188.8034 | 18.1463 | 25.0924 | 999.953 | 0.0056419 | 0.36477 | 241.882 | 251.149
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 186.3343 | 17.9421 | 25.0841 | 1000.26 | 0.0056299 | 0.36455 | 238.834 | 247.858
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 183.626 | 17.8267 | 25.08 | 999.94 | 0.0056218 | 0.36425 | 235.803 | 244.675
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 181.7254 | 17.4616 | 25.0717 | 1000.25 | 0.0056159 | 0.36455 | 232.817 | 241.286
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 179.7438 | 17.0973 | 25.0644 | 999.942 | 0.0056079 | 0.36456 | 229.792 | 237.857
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 177.0094 | 16.9808 | 25.0628 | 1000.25 | 0.0056022 | 0.36428 | 226.729 | 234.64
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 174.7355 | 16.7195 | 25.0583 | 999.932 | 0.005596 | 0.36418 | 223.707 | 231.319
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 172.5003 | 16.4642 | 25.0474 | 1000.24 | 0.0055873 | 0.36445 | 220.689 | 228.008
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 169.9066 | 16.3271 | 25.0404 | 999.931 | 0.0055786 | 0.36437 | 217.683 | 224.827
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 167.8954 | 15.9723 | 25.0381 | 999.926 | 0.0055722 | 0.36402 | 214.701 | 221.452
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 165.9456 | 15.6093 | 25.0318 | 999.923 | 0.0055662 | 0.36411 | 211.702 | 218.053
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 163.1783 | 15.5274 | 25.0294 | 999.931 | 0.0055612 | 0.36392 | 208.687 | 214.919
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 161.0924 | 15.2105 | 25.0253 | 1000.23 | 0.0055539 | 0.36372 | 205.712 | 211.59
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 159.1554 | 14.8336 | 25.0223 | 999.927 | 0.0055458 | 0.36334 | 202.743 | 208.206
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 156.7568 | 14.6282 | 25.0152 | 999.885 | 0.0055375 | 0.36333 | 199.735 | 204.955
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 154.5289 | 14.338 | 25.0128 | 1000.23 | 0.0055311 | 0.36303 | 196.7 | 201.592
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 152.2392 | 14.0812 | 25.0075 | 999.924 | 0.0055248 | 0.36299 | 193.664 | 198.262
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 149.7885 | 13.905 | 25.0004 | 999.918 | 0.0055193 | 0.3632 | 190.66 | 195.045
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 147.8566 | 13.5017 | 25.0043 | 1000.23 | 0.005513 | 0.36241 | 187.664 | 191.608
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 145.7356 | 13.176 | 25.0041 | 999.918 | 0.0055088 | 0.36212 | 184.634 | 188.215
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 143.0877 | 13.0579 | 24.9964 | 999.913 | 0.005504 | 0.36242 | 181.583 | 185.009
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 140.9452 | 12.7372 | 24.9979 | 999.878 | 0.0054984 | 0.36185 | 178.573 | 181.641
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 138.4886 | 12.5393 | 24.9932 | 999.902 | 0.0054891 | 0.36154 | 175.556 | 178.389
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 135.9392 | 12.4157 | 24.9872 | 999.904 | 0.0054871 | 0.36203 | 172.572 | 175.246
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 134.104 | 12.0755 | 24.9854 | 999.912 | 0.0055002 | 0.36363 | 169.609 | 171.909
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 132.458 | 11.6778 | 24.99 | 1000.22 | 0.00552 | 0.36554 | 166.663 | 168.533
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 130.0994 | 11.4854 | 25.0048 | 1000.22 | 0.0055384 | 0.36629 | 163.676 | 165.321
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 127.9083 | 11.2279 | 25.0087 | 1000.23 | 0.0055473 | 0.36705 | 160.679 | 162.032
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 125.8162 | 10.9374 | 25.0072 | 999.921 | 0.0055482 | 0.3675 | 157.718 | 158.745
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 123.2718 | 10.7816 | 25.0119 | 999.909 | 0.0055472 | 0.36708 | 154.744 | 155.58
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 121.3139 | 10.3927 | 25.0207 | 1000.22 | 0.0055422 | 0.36593 | 151.783 | 152.193
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 119.3491 | 9.9954 | 25.0282 | 999.913 | 0.0055356 | 0.36477 | 148.786 | 148.761
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 116.9426 | 9.7818 | 25.0289 | 999.901 | 0.0055318 | 0.36447 | 145.775 | 145.5
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 114.9995 | 9.3988 | 25.032 | 999.91 | 0.0055299 | 0.36407 | 142.773 | 142.078
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 112.8788 | 9.0499 | 25.0364 | 1000.22 | 0.0055224 | 0.36308 | 139.725 | 138.643
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 110.3274 | 8.8795 | 25.0357 | 1000.22 | 0.0055166 | 0.36268 | 136.689 | 135.4
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 108.177 | 8.5683 | 25.0385 | 999.909 | 0.0055134 | 0.36222 | 133.673 | 132.036
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 105.8223 | 8.3359 | 25.0382 | 999.86 | 0.0055076 | 0.36181 | 130.662 | 128.756
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 103.2319 | 8.1898 | 25.0393 | 999.868 | 0.0055041 | 0.36146 | 127.647 | 125.559
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 101.2858 | 7.8212 | 25.0416 | 999.91 | 0.0055008 | 0.36101 | 124.672 | 122.179
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 99.2599 | 7.4652 | 25.0499 | 1000.22 | 0.005496 | 0.35999 | 121.684 | 118.798
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 96.817 | 7.2549 | 25.0603 | 1000.22 | 0.0054937 | 0.35896 | 118.677 | 115.545
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 94.934 | 6.8584 | 25.0591 | 999.896 | 0.0054901 | 0.35878 | 115.661 | 112.095
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 92.8144 | 6.5277 | 25.0605 | 999.913 | 0.0054864 | 0.35844 | 112.598 | 108.664
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 90.1496 | 6.3924 | 25.0617 | 999.869 | 0.0054821 | 0.35808 | 109.53 | 105.425
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 87.9829 | 6.0978 | 25.0594 | 999.911 | 0.0054772 | 0.35792 | 106.508 | 102.071
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 85.6387 | 5.8537 | 25.0655 | 1000.21 | 0.0054732 | 0.35707 | 103.482 | 98.7642
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 83.1631 | 5.6714 | 25.0671 | 1000.22 | 0.0054671 | 0.35651 | 100.479 | 95.543
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 81.1576 | 5.3276 | 25.0661 | 999.901 | 0.0054608 | 0.35613 | 97.4868 | 92.1703
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 79.1344 | 4.9912 | 25.0633 | 999.868 | 0.0054569 | 0.35607 | 94.4822 | 88.793
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 76.7682 | 4.776 | 25.0607 | 999.892 | 0.0054539 | 0.35614 | 91.4669 | 85.5265
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 74.7258 | 4.4478 | 25.0602 | 999.892 | 0.0054485 | 0.35582 | 88.4755 | 82.1699
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 72.44 | 4.2086 | 25.0642 | 1000.21 | 0.005444 | 0.35512 | 85.5018 | 78.9216
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 70.1066 | 3.9934 | 25.0663 | 1000.21 | 0.0054399 | 0.35462 | 82.5386 | 75.7085
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 68.1269 | 3.6457 | 25.0687 | 999.898 | 0.0054338 | 0.35395 | 79.5607 | 72.346
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 66.0587 | 3.3367 | 25.0694 | 999.865 | 0.005431 | 0.35372 | 76.587 | 69.0279
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 63.7358 | 3.1219 | 25.0602 | 999.9 | 0.0054259 | 0.35417 | 73.6025 | 65.793
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 61.7135 | 2.7949 | 25.0598 | 999.91 | 0.0054219 | 0.35395 | 70.6234 | 62.4497
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 59.4804 | 2.5472 | 25.0619 | 999.874 | 0.0054186 | 0.35356 | 67.6591 | 59.2022
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 57.1982 | 2.316 | 25.0574 | 999.907 | 0.0054125 | 0.35343 | 64.6843 | 55.9615
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 55.3622 | 1.9242 | 25.0512 | 999.905 | 0.0054086 | 0.35364 | 61.6986 | 52.5472
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 53.3902 | 1.5708 | 25.0483 | 1000.21 | 0.0054042 | 0.35356 | 58.6895 | 49.1486
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 51.0186 | 1.3548 | 25.0479 | 1000.22 | 0.0053985 | 0.35319 | 55.6699 | 45.8781
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 49.2514 | 0.92474 | 25.0454 | 1000.23 | 0.0053925 | 0.35291 | 52.6503 | 42.39
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 47.2751 | 0.55956 | 25.0428 | 999.906 | 0.0053851 | 0.35252 | 49.6024 | 38.9386
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 44.7233 | 0.38726 | 25.0414 | 999.909 | 0.0053828 | 0.35249 | 46.5186 | 35.6487
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 42.6235 | 0.067136 | 25.036 | 999.877 | 0.0053798 | 0.35273 | 43.4704 | 32.2437
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 40.2757 | -0.15519 | 25.0293 | 999.902 | 0.0053712 | 0.35259 | 40.4504 | 28.9645
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 37.7605 | -0.31149 | 25.0267 | 999.925 | 0.0053672 | 0.35247 | 37.4494 | 25.7732
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 35.9329 | -0.69989 | 25.0258 | 999.926 | 0.0053619 | 0.35217 | 34.4816 | 22.3804
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 33.9711 | -1.0506 | 25.0219 | 1000.23 | 0.0053538 | 0.35185 | 31.4832 | 18.9941
Populus deltoides | LI1A |  | RACiR 500 to 0 no VPD control | 31.6441 | -1.2866 | 25.0135 | 999.918 | 0.0053511 | 0.35238 | 28.4491 | 15.6889
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 296.3724 | 29.2092 | 25.2587 | 1000.01 | 0.0041829 | 0.44869 | 417.436 | 431.234
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 291.1973 | 29.0075 | 25.2417 | 1000.01 | 0.004169 | 0.44833 | 411.442 | 424.96
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 287.7371 | 28.3452 | 25.2292 | 1000.01 | 0.0041568 | 0.44766 | 405.435 | 418.212
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 283.908 | 27.7173 | 25.223 | 999.995 | 0.0041448 | 0.44634 | 399.361 | 411.427
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 278.3786 | 27.5831 | 25.2061 | 999.991 | 0.0041336 | 0.44619 | 393.212 | 405.066
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 274.18 | 27.1112 | 25.1886 | 999.988 | 0.0041206 | 0.44584 | 387.095 | 398.399
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 269.3555 | 26.7582 | 25.1783 | 999.987 | 0.0041075 | 0.44473 | 381.01 | 391.879
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 264.0766 | 26.5533 | 25.1664 | 999.987 | 0.0040963 | 0.44418 | 374.952 | 385.538
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 260.6375 | 25.9011 | 25.1455 | 999.988 | 0.0040803 | 0.444 | 368.876 | 378.732
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 256.4734 | 25.4498 | 25.1237 | 999.987 | 0.0040681 | 0.44443 | 362.737 | 372.059
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 250.968 | 25.3138 | 25.1092 | 1000.02 | 0.0040595 | 0.4443 | 356.58 | 365.689
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 246.8164 | 24.8444 | 25.0943 | 1000.02 | 0.004049 | 0.44397 | 350.499 | 359.062
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 242.0704 | 24.5659 | 25.0779 | 999.977 | 0.0040407 | 0.44439 | 344.47 | 352.678
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 237.1511 | 24.3471 | 25.0599 | 1000.01 | 0.0040323 | 0.44517 | 338.458 | 346.372
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 233.8177 | 23.6963 | 25.0465 | 1000 | 0.0040231 | 0.44522 | 332.451 | 339.639
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 230.2312 | 23.1255 | 25.0339 | 1000.01 | 0.0040207 | 0.44587 | 326.372 | 332.91
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 225.347 | 22.9043 | 25.0118 | 999.978 | 0.0040139 | 0.44728 | 320.253 | 326.495
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 222.2041 | 22.2421 | 24.9923 | 1000 | 0.0040082 | 0.44853 | 314.16 | 319.663
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 218.6977 | 21.6855 | 24.9775 | 999.974 | 0.0040112 | 0.45057 | 308.016 | 312.886
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 213.6048 | 21.555 | 24.9657 | 999.973 | 0.0040195 | 0.45334 | 301.843 | 306.513
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 211.649 | 21.0544 | 24.952 | 999.981 | 0.004085 | 0.46673 | 295.679 | 299.793
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 210.7206 | 20.6949 | 24.9483 | 999.986 | 0.0042223 | 0.49261 | 289.537 | 293.255
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 208.7589 | 20.5028 | 24.9619 | 999.944 | 0.004364 | 0.51849 | 283.486 | 286.98
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 207.3236 | 19.8424 | 24.9784 | 999.982 | 0.0044739 | 0.53792 | 277.464 | 280.251
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 204.5457 | 19.2307 | 24.9919 | 999.979 | 0.0045391 | 0.54879 | 271.409 | 273.524
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 199.6773 | 18.9439 | 25.0083 | 999.981 | 0.0045615 | 0.5508 | 265.328 | 267.086
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 194.902 | 18.4611 | 25.0238 | 999.974 | 0.0045562 | 0.54691 | 259.276 | 260.472
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 189.3444 | 18.1221 | 25.0393 | 999.941 | 0.0045335 | 0.53949 | 253.249 | 254.024
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 183.3455 | 17.8662 | 25.0525 | 999.975 | 0.0044957 | 0.53029 | 247.267 | 247.702
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 178.8583 | 17.1918 | 25.0629 | 999.967 | 0.0044574 | 0.52134 | 241.335 | 241.009
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 174.2684 | 16.5798 | 25.0743 | 999.934 | 0.0044274 | 0.51344 | 235.34 | 234.318
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 168.7022 | 16.2647 | 25.0851 | 999.976 | 0.0044009 | 0.50686 | 229.263 | 227.844
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 164.3759 | 15.6177 | 25.0961 | 999.971 | 0.004377 | 0.50074 | 223.188 | 221.038
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 160.2286 | 14.9832 | 25.1045 | 999.976 | 0.0043618 | 0.49657 | 217.116 | 214.251
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 154.9558 | 14.688 | 25.1055 | 999.93 | 0.0043444 | 0.49315 | 211.021 | 207.781
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 150.4338 | 14.1992 | 25.1062 | 999.977 | 0.0043275 | 0.48983 | 204.956 | 201.147
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 145.2735 | 13.8946 | 25.1101 | 999.967 | 0.0043134 | 0.48682 | 198.898 | 194.707
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 139.9613 | 13.6648 | 25.115 | 999.925 | 0.0043033 | 0.48429 | 192.903 | 188.408
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 136.5482 | 12.9422 | 25.1183 | 999.961 | 0.0042942 | 0.48191 | 186.956 | 181.659
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 132.6474 | 12.36 | 25.1178 | 1000.26 | 0.004285 | 0.48031 | 180.965 | 175.008
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 127.673 | 12.0505 | 25.128 | 999.941 | 0.0042808 | 0.47819 | 174.951 | 168.611
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 123.6831 | 11.497 | 25.1355 | 1000.25 | 0.0042772 | 0.47633 | 168.967 | 161.995
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 119.1043 | 11.1351 | 25.1327 | 999.946 | 0.0042737 | 0.47581 | 162.991 | 155.585
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 114.4311 | 10.7986 | 25.1295 | 1000.25 | 0.0042673 | 0.47506 | 157.04 | 149.221
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 111.2722 | 10.0516 | 25.1296 | 999.943 | 0.0042592 | 0.4736 | 151.119 | 142.474
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 107.8251 | 9.3748 | 25.1304 | 999.898 | 0.0042547 | 0.47245 | 145.132 | 135.738
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 103.2577 | 8.9979 | 25.1258 | 999.931 | 0.0042498 | 0.47208 | 139.087 | 129.238
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 100.0132 | 8.2848 | 25.1153 | 999.915 | 0.0042443 | 0.47234 | 133.063 | 122.421
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 96.369 | 7.6491 | 25.1138 | 1000.25 | 0.004241 | 0.47161 | 126.984 | 115.631
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 91.3829 | 7.3596 | 25.1124 | 999.948 | 0.0042366 | 0.47075 | 120.861 | 109.144
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 87.4137 | 6.8132 | 25.1061 | 999.949 | 0.0042288 | 0.47022 | 114.775 | 102.433
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 82.6087 | 6.4863 | 25.1055 | 999.941 | 0.0042246 | 0.46953 | 108.686 | 95.9411
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 77.6074 | 6.2318 | 25.1046 | 1000.26 | 0.0042209 | 0.46874 | 102.669 | 89.5977
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 74.4317 | 5.515 | 25.0992 | 999.933 | 0.0042173 | 0.46875 | 96.7023 | 82.8355
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 70.772 | 4.919 | 25.0969 | 999.899 | 0.0042153 | 0.46864 | 90.7141 | 76.1748
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 66.1889 | 4.5624 | 25.0936 | 999.937 | 0.0042137 | 0.46838 | 84.6973 | 69.7317
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 62.6352 | 3.9385 | 25.0841 | 1000.25 | 0.0042093 | 0.46878 | 78.689 | 63.018
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 58.3277 | 3.5102 | 25.0807 | 999.935 | 0.0042095 | 0.46909 | 72.6696 | 56.4982
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 53.7366 | 3.1582 | 25.0781 | 999.925 | 0.0042062 | 0.46873 | 66.6684 | 50.0723
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 50.6434 | 2.4147 | 25.0717 | 1000.24 | 0.0042016 | 0.46857 | 60.686 | 43.2639
Populus deltoides | LI1A |  | RACiR 500 to 0 2.5 min | 47.2113 | 1.7415 | 25.0723 | 999.93 | 0.0042013 | 0.46807 | 54.6311 | 36.4647
Populus deltoides | LI1A |  | RACiR 300 to 800 | 222.4728 | 22.5424 | 24.9578 | 1000.23 | 0.0088715 | 0.39467 | 285.297 | 319.345
Populus deltoides | LI1A |  | RACiR 300 to 800 | 223.5837 | 22.6459 | 24.9561 | 999.879 | 0.0088738 | 0.39488 | 286.665 | 320.833
Populus deltoides | LI1A |  | RACiR 300 to 800 | 224.0993 | 22.8952 | 24.9637 | 1000.24 | 0.0088733 | 0.39437 | 287.914 | 322.354
Populus deltoides | LI1A |  | RACiR 300 to 800 | 224.9716 | 23.0931 | 24.9625 | 999.885 | 0.0088724 | 0.39442 | 289.306 | 323.969
Populus deltoides | LI1A |  | RACiR 300 to 800 | 225.7761 | 23.2558 | 24.9661 | 999.887 | 0.0088721 | 0.39415 | 290.59 | 325.423
Populus deltoides | LI1A |  | RACiR 300 to 800 | 226.1708 | 23.6438 | 24.968 | 999.872 | 0.0088696 | 0.3939 | 292.03 | 327.276
Populus deltoides | LI1A |  | RACiR 300 to 800 | 226.7648 | 23.9318 | 24.9643 | 999.878 | 0.0088681 | 0.39405 | 293.351 | 328.914
Populus deltoides | LI1A |  | RACiR 300 to 800 | 227.9246 | 24.0553 | 24.9631 | 1000.22 | 0.008868 | 0.39407 | 294.85 | 330.554
Populus deltoides | LI1A |  | RACiR 300 to 800 | 228.5537 | 24.3399 | 24.9624 | 999.883 | 0.0088679 | 0.39413 | 296.213 | 332.223
Populus deltoides | LI1A |  | RACiR 300 to 800 | 229.5335 | 24.5628 | 24.96 | 1000.23 | 0.0088676 | 0.39427 | 297.763 | 334.031
Populus deltoides | LI1A |  | RACiR 300 to 800 | 230.8456 | 24.6005 | 24.9578 | 999.88 | 0.0088691 | 0.39448 | 299.166 | 335.485
Populus deltoides | LI1A |  | RACiR 300 to 800 | 231.852 | 24.8077 | 24.9584 | 999.891 | 0.0088691 | 0.39446 | 300.73 | 337.279
Populus deltoides | LI1A |  | RACiR 300 to 800 | 232.9371 | 24.9235 | 24.9574 | 1000.24 | 0.0088667 | 0.39439 | 302.147 | 338.848
Populus deltoides | LI1A |  | RACiR 300 to 800 | 234.4705 | 24.9207 | 24.9572 | 999.883 | 0.0088666 | 0.3944 | 303.706 | 340.416
Populus deltoides | LI1A |  | RACiR 300 to 800 | 235.2798 | 25.1285 | 24.9615 | 1000.23 | 0.0088672 | 0.39416 | 305.11 | 342.05
Populus deltoides | LI1A |  | RACiR 300 to 800 | 236.3791 | 25.2915 | 24.9644 | 999.882 | 0.0088664 | 0.39397 | 306.687 | 343.82
Populus deltoides | LI1A |  | RACiR 300 to 800 | 237.6086 | 25.3532 | 24.966 | 999.883 | 0.0088696 | 0.39401 | 308.098 | 345.301
Populus deltoides | LI1A |  | RACiR 300 to 800 | 238.4437 | 25.6145 | 24.9694 | 999.883 | 0.0088699 | 0.39381 | 309.659 | 347.15
Populus deltoides | LI1A |  | RACiR 300 to 800 | 239.4565 | 25.7383 | 24.9709 | 999.879 | 0.0088652 | 0.39346 | 311.074 | 348.725
Populus deltoides | LI1A |  | RACiR 300 to 800 | 240.8902 | 25.8002 | 24.9702 | 1000.23 | 0.0088686 | 0.39369 | 312.66 | 350.385
Populus deltoides | LI1A |  | RACiR 300 to 800 | 241.8319 | 25.984 | 24.9686 | 999.888 | 0.0088691 | 0.39381 | 314.072 | 352.007
Populus deltoides | LI1A |  | RACiR 300 to 800 | 242.9349 | 26.1672 | 24.9655 | 1000.24 | 0.0088657 | 0.39388 | 315.656 | 353.812
Populus deltoides | LI1A |  | RACiR 300 to 800 | 244.1514 | 26.2754 | 24.9624 | 999.881 | 0.0088727 | 0.39441 | 317.083 | 355.349
Populus deltoides | LI1A |  | RACiR 300 to 800 | 245.1487 | 26.4978 | 24.9623 | 999.89 | 0.0088726 | 0.39445 | 318.666 | 357.184
Populus deltoides | LI1A |  | RACiR 300 to 800 | 246.1675 | 26.6299 | 24.963 | 1000.24 | 0.008867 | 0.39415 | 320.099 | 358.784
Populus deltoides | LI1A |  | RACiR 300 to 800 | 247.6311 | 26.662 | 24.968 | 999.889 | 0.0088705 | 0.39401 | 321.709 | 360.436
Populus deltoides | LI1A |  | RACiR 300 to 800 | 248.7406 | 26.7994 | 24.9693 | 1000.23 | 0.0088768 | 0.39421 | 323.162 | 362.056
Populus deltoides | LI1A |  | RACiR 300 to 800 | 249.9263 | 26.9883 | 24.9678 | 999.897 | 0.0088875 | 0.39475 | 324.763 | 363.885
Populus deltoides | LI1A |  | RACiR 300 to 800 | 251.2178 | 27.0725 | 24.9703 | 999.892 | 0.0089015 | 0.39514 | 326.235 | 365.458
Populus deltoides | LI1A |  | RACiR 300 to 800 | 252.4047 | 27.2763 | 24.968 | 999.883 | 0.0089168 | 0.39582 | 327.851 | 367.302
Populus deltoides | LI1A |  | RACiR 300 to 800 | 253.409 | 27.421 | 24.9718 | 999.887 | 0.0089191 | 0.39552 | 329.308 | 368.94
Populus deltoides | LI1A |  | RACiR 300 to 800 | 254.9938 | 27.3953 | 24.9737 | 999.892 | 0.0089144 | 0.39499 | 330.962 | 370.587
Populus deltoides | LI1A |  | RACiR 300 to 800 | 255.9998 | 27.5636 | 24.9716 | 999.896 | 0.0089117 | 0.39489 | 332.44 | 372.251
Populus deltoides | LI1A |  | RACiR 300 to 800 | 257.3131 | 27.6615 | 24.9707 | 999.897 | 0.0089087 | 0.39469 | 334.071 | 374.011
Populus deltoides | LI1A |  | RACiR 300 to 800 | 258.7941 | 27.6356 | 24.9673 | 1000.24 | 0.0089036 | 0.39458 | 335.537 | 375.472
Populus deltoides | LI1A |  | RACiR 300 to 800 | 259.6841 | 27.9008 | 24.9658 | 999.886 | 0.0089015 | 0.39448 | 337.145 | 377.363
Populus deltoides | LI1A |  | RACiR 300 to 800 | 260.7854 | 28.0268 | 24.9646 | 999.893 | 0.0089005 | 0.39445 | 338.6 | 378.969
Populus deltoides | LI1A |  | RACiR 300 to 800 | 262.167 | 28.0978 | 24.9645 | 999.912 | 0.008896 | 0.39421 | 340.24 | 380.707
Populus deltoides | LI1A |  | RACiR 300 to 800 | 263.1623 | 28.2763 | 24.964 | 999.897 | 0.0088975 | 0.39425 | 341.71 | 382.369
Populus deltoides | LI1A |  | RACiR 300 to 800 | 264.3468 | 28.4298 | 24.9657 | 1000.25 | 0.0088959 | 0.39406 | 343.352 | 384.195
Populus deltoides | LI1A |  | RACiR 300 to 800 | 265.801 | 28.4096 | 24.9687 | 1000.25 | 0.008897 | 0.3939 | 344.822 | 385.668
Populus deltoides | LI1A |  | RACiR 300 to 800 | 267.0065 | 28.5631 | 24.9692 | 999.905 | 0.0088973 | 0.39388 | 346.452 | 387.474
Populus deltoides | LI1A |  | RACiR 300 to 800 | 267.9423 | 28.749 | 24.9709 | 1000.24 | 0.0088971 | 0.3938 | 347.903 | 389.14
Populus deltoides | LI1A |  | RACiR 300 to 800 | 269.6056 | 28.7215 | 24.9715 | 999.902 | 0.0088974 | 0.39378 | 349.539 | 390.771
Populus deltoides | LI1A |  | RACiR 300 to 800 | 270.8026 | 28.8251 | 24.9697 | 999.906 | 0.0088978 | 0.39392 | 351.001 | 392.358
Populus deltoides | LI1A |  | RACiR 300 to 800 | 272.2152 | 28.9075 | 24.9664 | 1000.26 | 0.0088984 | 0.39418 | 352.603 | 394.066
Populus deltoides | LI1A |  | RACiR 300 to 800 | 273.564 | 28.9256 | 24.9672 | 999.912 | 0.0088976 | 0.39412 | 354.042 | 395.545
Populus deltoides | LI1A |  | RACiR 300 to 800 | 274.7272 | 29.097 | 24.9665 | 999.904 | 0.0089009 | 0.39432 | 355.631 | 397.332
Populus deltoides | LI1A |  | RACiR 300 to 800 | 275.5991 | 29.3031 | 24.9673 | 999.904 | 0.0088999 | 0.39425 | 357.067 | 398.994
Populus deltoides | LI1A |  | RACiR 300 to 800 | 276.9499 | 29.3723 | 24.974 | 999.922 | 0.0089016 | 0.39391 | 358.698 | 400.719
Populus deltoides | LI1A |  | RACiR 300 to 800 | 277.97 | 29.5285 | 24.9769 | 999.903 | 0.0089015 | 0.39375 | 360.176 | 402.38
Populus deltoides | LI1A |  | RACiR 300 to 800 | 279.2323 | 29.6451 | 24.9812 | 999.899 | 0.0088996 | 0.3934 | 361.839 | 404.188
Populus deltoides | LI1A |  | RACiR 300 to 800 | 280.5669 | 29.6645 | 24.9867 | 999.9 | 0.0089008 | 0.39309 | 363.322 | 405.706
Populus deltoides | LI1A |  | RACiR 300 to 800 | 281.7928 | 29.7979 | 24.9877 | 1000.24 | 0.0088986 | 0.39294 | 364.95 | 407.497
Populus deltoides | LI1A |  | RACiR 300 to 800 | 282.9644 | 29.8942 | 24.9865 | 999.899 | 0.0088964 | 0.39291 | 366.4 | 409.075
Populus deltoides | LI1A |  | RACiR 300 to 800 | 284.6538 | 29.8622 | 24.9879 | 999.907 | 0.0089003 | 0.393 | 368.029 | 410.681
Populus deltoides | LI1A |  | RACiR 300 to 800 | 285.5984 | 30.058 | 24.9858 | 1000.26 | 0.0089006 | 0.3932 | 369.457 | 412.331
Populus deltoides | LI1A |  | RACiR 300 to 800 | 286.7353 | 30.234 | 24.9865 | 999.907 | 0.0089001 | 0.39316 | 371.081 | 414.165
Populus deltoides | LI1A |  | RACiR 300 to 800 | 287.8885 | 30.353 | 24.9864 | 999.92 | 0.0089038 | 0.39331 | 372.537 | 415.749
Populus deltoides | LI1A |  | RACiR 300 to 800 | 288.7702 | 30.6219 | 24.9871 | 999.912 | 0.0089015 | 0.39314 | 374.165 | 417.679
Populus deltoides | LI1A |  | RACiR 300 to 800 | 290.0615 | 30.7028 | 24.9843 | 1000.25 | 0.0089005 | 0.39325 | 375.669 | 419.289
Populus deltoides | LI1A |  | RACiR 300 to 800 | 291.8667 | 30.6403 | 24.98 | 999.91 | 0.0088991 | 0.39344 | 377.312 | 420.89
Populus deltoides | LI1A |  | RACiR 300 to 800 | 293.079 | 30.7448 | 24.9771 | 999.912 | 0.0088985 | 0.39358 | 378.789 | 422.494
Populus deltoides | LI1A |  | RACiR 300 to 800 | 294.3402 | 30.8752 | 24.977 | 999.915 | 0.0088993 | 0.39362 | 380.406 | 424.266
Populus deltoides | LI1A |  | RACiR 300 to 800 | 295.76 | 30.8638 | 24.9811 | 999.905 | 0.0088991 | 0.39337 | 381.884 | 425.754
Populus deltoides | LI1A |  | RACiR 300 to 800 | 296.5586 | 31.1598 | 24.9859 | 999.948 | 0.0089014 | 0.3932 | 383.503 | 427.695
Populus deltoides | LI1A |  | RACiR 300 to 800 | 297.5886 | 31.2977 | 24.991 | 1000.25 | 0.0089013 | 0.39289 | 384.98 | 429.342
Populus deltoides | LI1A |  | RACiR 300 to 800 | 299.2404 | 31.2999 | 24.9893 | 999.898 | 0.008901 | 0.393 | 386.652 | 431.036
Populus deltoides | LI1A |  | RACiR 300 to 800 | 300.1574 | 31.5034 | 24.9904 | 999.897 | 0.0089009 | 0.39292 | 388.132 | 432.737
Populus deltoides | LI1A |  | RACiR 300 to 800 | 301.5196 | 31.6009 | 24.9898 | 1000.25 | 0.0088995 | 0.39289 | 389.781 | 434.516
Populus deltoides | LI1A |  | RACiR 300 to 800 | 303.168 | 31.5189 | 24.9898 | 999.903 | 0.0088978 | 0.39284 | 391.265 | 435.938
Populus deltoides | LI1A |  | RACiR 300 to 800 | 304.201 | 31.7269 | 24.9936 | 999.903 | 0.0089013 | 0.39277 | 392.878 | 437.779
Populus deltoides | LI1A |  | RACiR 300 to 800 | 305.3905 | 31.8151 | 24.9954 | 999.912 | 0.0089006 | 0.39265 | 394.349 | 439.375
Populus deltoides | LI1A |  | RACiR 300 to 800 | 306.9994 | 31.8117 | 24.9972 | 999.905 | 0.0089034 | 0.39269 | 395.979 | 441.016
Populus deltoides | LI1A |  | RACiR 300 to 800 | 308.169 | 31.9246 | 24.9977 | 999.913 | 0.0089048 | 0.39271 | 397.462 | 442.634
Populus deltoides | LI1A |  | RACiR 300 to 800 | 309.584 | 32.0007 | 24.9945 | 999.918 | 0.0089015 | 0.39277 | 399.088 | 444.372
Populus deltoides | LI1A |  | RACiR 300 to 800 | 310.9357 | 32.0517 | 24.9931 | 999.917 | 0.0089041 | 0.39303 | 400.546 | 445.89
Populus deltoides | LI1A |  | RACiR 300 to 800 | 312.0552 | 32.2076 | 24.9956 | 999.911 | 0.0089009 | 0.39271 | 402.163 | 447.693
Populus deltoides | LI1A |  | RACiR 300 to 800 | 313.0168 | 32.3851 | 24.9964 | 1000.26 | 0.0088998 | 0.39262 | 403.624 | 449.356
Populus deltoides | LI1A |  | RACiR 300 to 800 | 314.3981 | 32.4689 | 24.9973 | 999.914 | 0.0089019 | 0.39264 | 405.249 | 451.082
Populus deltoides | LI1A |  | RACiR 300 to 800 | 315.5265 | 32.5798 | 24.9986 | 999.914 | 0.0089007 | 0.39249 | 406.723 | 452.699
Populus deltoides | LI1A |  | RACiR 300 to 800 | 317.0072 | 32.6451 | 24.9981 | 1000.26 | 0.0089032 | 0.39271 | 408.358 | 454.427
Populus deltoides | LI1A |  | RACiR 300 to 800 | 318.5775 | 32.6131 | 24.9976 | 999.924 | 0.0089055 | 0.39286 | 409.848 | 455.9
Populus deltoides | LI1A |  | RACiR 300 to 800 | 319.6535 | 32.7959 | 25.0018 | 999.921 | 0.0089049 | 0.39259 | 411.482 | 457.75
Populus deltoides | LI1A |  | RACiR 300 to 800 | 320.871 | 32.8804 | 25.003 | 999.925 | 0.0089058 | 0.39259 | 412.946 | 459.315
Populus deltoides | LI1A |  | RACiR 300 to 800 | 322.5184 | 32.8795 | 25.0004 | 999.939 | 0.0089048 | 0.39271 | 414.599 | 460.987
Populus deltoides | LI1A |  | RACiR 300 to 800 | 323.4146 | 33.0791 | 25.0019 | 999.914 | 0.0089055 | 0.39268 | 416.039 | 462.656
Populus deltoides | LI1A |  | RACiR 300 to 800 | 324.6264 | 33.25 | 24.9998 | 999.921 | 0.0089077 | 0.39296 | 417.654 | 464.466
Populus deltoides | LI1A |  | RACiR 300 to 800 | 325.808 | 33.375 | 24.9981 | 999.921 | 0.0089103 | 0.39321 | 419.127 | 466.081
Populus deltoides | LI1A |  | RACiR 300 to 800 | 326.7826 | 33.5988 | 25.0028 | 999.926 | 0.0089115 | 0.39296 | 420.761 | 467.967
Populus deltoides | LI1A |  | RACiR 300 to 800 | 328.0596 | 33.6593 | 25.0021 | 999.92 | 0.0089049 | 0.39268 | 422.285 | 469.588
Populus deltoides | LI1A |  | RACiR 300 to 800 | 329.5979 | 33.6723 | 25.005 | 999.918 | 0.0089052 | 0.39252 | 423.931 | 471.263
Populus deltoides | LI1A |  | RACiR 300 to 800 | 330.5984 | 33.8306 | 25.008 | 999.936 | 0.0089062 | 0.39233 | 425.409 | 472.915
Populus deltoides | LI1A |  | RACiR 300 to 800 | 332.0989 | 33.8723 | 25.0049 | 999.916 | 0.008902 | 0.39234 | 427.046 | 474.637
Populus deltoides | LI1A |  | RACiR 300 to 800 | 333.5536 | 33.8614 | 25.0069 | 999.913 | 0.0089045 | 0.39233 | 428.511 | 476.101
Populus deltoides | LI1A |  | RACiR 300 to 800 | 334.5946 | 34.1015 | 25.0046 | 1000.26 | 0.0089083 | 0.39263 | 430.123 | 477.968
Populus deltoides | LI1A |  | RACiR 300 to 800 | 335.7918 | 34.1869 | 25.0016 | 999.911 | 0.008903 | 0.39256 | 431.58 | 479.554
Populus deltoides | LI1A |  | RACiR 300 to 800 | 337.4413 | 34.1782 | 25.001 | 999.924 | 0.0089022 | 0.39254 | 433.249 | 481.234
Populus deltoides | LI1A |  | RACiR 300 to 800 | 338.3593 | 34.4251 | 24.9972 | 999.915 | 0.0089061 | 0.39296 | 434.725 | 482.967
Populus deltoides | LI1A |  | RACiR 300 to 800 | 339.7064 | 34.515 | 24.9965 | 999.93 | 0.0089004 | 0.39268 | 436.399 | 484.773
Populus deltoides | LI1A |  | RACiR 300 to 800 | 341.2705 | 34.4663 | 24.9984 | 999.924 | 0.0089008 | 0.39255 | 437.906 | 486.244
Populus deltoides | LI1A |  | RACiR 300 to 800 | 342.4665 | 34.6496 | 24.9966 | 999.935 | 0.0089024 | 0.39272 | 439.559 | 488.1
Populus deltoides | LI1A |  | RACiR 300 to 800 | 343.8581 | 34.7105 | 24.995 | 1000.27 | 0.008909 | 0.39311 | 441.05 | 489.684
Populus deltoides | LI1A |  | RACiR 300 to 800 | 345.8211 | 34.64 | 24.9931 | 999.933 | 0.0089254 | 0.39392 | 442.692 | 491.284
Populus deltoides | LI1A |  | RACiR 300 to 800 | 347.168 | 34.7457 | 24.9888 | 999.922 | 0.0089381 | 0.39468 | 444.164 | 492.892
Populus deltoides | LI1A |  | RACiR 300 to 800 | 348.6203 | 34.8414 | 24.9852 | 999.932 | 0.0089479 | 0.39519 | 445.778 | 494.637
Populus deltoides | LI1A |  | RACiR 300 to 800 | 350.1232 | 34.8292 | 24.9818 | 1000.27 | 0.0089527 | 0.39544 | 447.227 | 496.091
Populus deltoides | LI1A |  | RACiR 300 to 800 | 351.2925 | 34.9652 | 24.9803 | 999.931 | 0.0089471 | 0.39507 | 448.857 | 497.885
Populus deltoides | LI1A |  | RACiR 300 to 800 | 352.2035 | 35.1395 | 24.9806 | 999.935 | 0.0089427 | 0.39472 | 450.317 | 499.541
Populus deltoides | LI1A |  | RACiR 300 to 800 | 353.4199 | 35.2519 | 24.9856 | 999.937 | 0.0089423 | 0.39428 | 451.962 | 501.316
Populus deltoides | LI1A |  | RACiR 300 to 800 | 354.4399 | 35.369 | 24.9902 | 1000.29 | 0.0089383 | 0.39369 | 453.45 | 502.944
Populus deltoides | LI1A |  | RACiR 300 to 800 | 355.6189 | 35.4939 | 24.9917 | 999.935 | 0.0089321 | 0.39323 | 455.09 | 504.736
Populus deltoides | LI1A |  | RACiR 300 to 800 | 357.0201 | 35.4922 | 24.9916 | 999.93 | 0.0089272 | 0.39294 | 456.587 | 506.245
Populus deltoides | LI1A |  | RACiR 300 to 800 | 358.2618 | 35.6083 | 24.99 | 999.945 | 0.0089195 | 0.39262 | 458.231 | 508.028
Populus deltoides | LI1A |  | RACiR 300 to 800 | 359.5466 | 35.6459 | 24.9896 | 999.936 | 0.0089127 | 0.39233 | 459.708 | 509.573
Populus deltoides | LI1A |  | RACiR 300 to 800 | 361.4796 | 35.5579 | 24.9841 | 1000.29 | 0.0089118 | 0.39269 | 461.367 | 511.158
Populus deltoides | LI1A |  | RACiR 300 to 800 | 362.5423 | 35.7072 | 24.9829 | 999.939 | 0.0089097 | 0.39267 | 462.84 | 512.803
Populus deltoides | LI1A |  | RACiR 300 to 800 | 363.5103 | 35.9577 | 24.9809 | 999.93 | 0.0089083 | 0.39272 | 464.462 | 514.706
Populus deltoides | LI1A |  | RACiR 300 to 800 | 364.821 | 36.0298 | 24.9796 | 999.931 | 0.0089093 | 0.39282 | 465.963 | 516.29
Populus deltoides | LI1A |  | RACiR 300 to 800 | 365.8615 | 36.2456 | 24.9785 | 1000.27 | 0.0089047 | 0.39264 | 467.624 | 518.207
Populus deltoides | LI1A |  | RACiR 300 to 800 | 367.2812 | 36.3038 | 24.9736 | 999.934 | 0.0089063 | 0.393 | 469.135 | 519.796
Populus deltoides | LI1A |  | RACiR 300 to 800 | 369.0997 | 36.2706 | 24.9701 | 999.935 | 0.0089083 | 0.39333 | 470.826 | 521.473
Populus deltoides | LI1A |  | RACiR 300 to 800 | 370.3322 | 36.3243 | 24.9713 | 999.934 | 0.0089007 | 0.3929 | 472.33 | 523.061
Populus deltoides | LI1A |  | RACiR 300 to 800 | 371.5697 | 36.3972 | 24.9731 | 999.946 | 0.008885 | 0.39205 | 473.988 | 524.805
Populus deltoides | LI1A |  | RACiR 300 to 800 | 372.8686 | 36.402 | 24.9755 | 999.933 | 0.0088733 | 0.39144 | 475.48 | 526.31
Populus deltoides | LI1A |  | RACiR 300 to 800 | 373.9432 | 36.5746 | 24.974 | 999.933 | 0.0088592 | 0.391 | 477.125 | 528.148
Populus deltoides | LI1A |  | RACiR 300 to 800 | 375.1915 | 36.6677 | 24.9682 | 999.925 | 0.0088506 | 0.39109 | 478.61 | 529.753
Populus deltoides | LI1A |  | RACiR 300 to 800 | 376.7213 | 36.7363 | 24.97 | 1000.28 | 0.0088561 | 0.39141 | 480.272 | 531.503
Populus deltoides | LI1A |  | RACiR 300 to 800 | 377.6823 | 36.9577 | 24.9703 | 999.929 | 0.0088576 | 0.39164 | 481.769 | 533.256
Populus deltoides | LI1A |  | RACiR 300 to 800 | 379.0929 | 37.0557 | 24.9735 | 999.928 | 0.0088616 | 0.39173 | 483.445 | 535.061
Populus deltoides | LI1A |  | RACiR 300 to 800 | 380.6271 | 37.0216 | 24.979 | 999.924 | 0.0088646 | 0.39159 | 484.965 | 536.557
Populus deltoides | LI1A |  | RACiR 300 to 800 | 382.078 | 37.1246 | 24.9794 | 1000.28 | 0.0088668 | 0.39181 | 486.659 | 538.387
Populus deltoides | LI1A |  | RACiR 300 to 800 | 383.5163 | 37.1504 | 24.98 | 999.935 | 0.0088694 | 0.39193 | 488.168 | 539.947
Populus deltoides | LI1A |  | RACiR 300 to 800 | 385.3479 | 37.0835 | 24.9841 | 999.928 | 0.0088759 | 0.39204 | 489.846 | 541.576
Populus deltoides | LI1A |  | RACiR 300 to 800 | 386.5536 | 37.1695 | 24.9863 | 999.927 | 0.008876 | 0.39193 | 491.33 | 543.173
Populus deltoides | LI1A |  | RACiR 300 to 800 | 388.0494 | 37.1965 | 24.9892 | 999.939 | 0.0088759 | 0.39177 | 492.975 | 544.875
Populus deltoides | LI1A |  | RACiR 300 to 800 | 389.5031 | 37.1754 | 24.9923 | 999.934 | 0.0088757 | 0.39157 | 494.461 | 546.356
Populus deltoides | LI1A |  | RACiR 300 to 800 | 390.5092 | 37.3954 | 24.991 | 999.93 | 0.0088713 | 0.39146 | 496.082 | 548.231
Populus deltoides | LI1A |  | RACiR 300 to 800 | 391.7891 | 37.4785 | 24.987 | 999.931 | 0.0088692 | 0.3916 | 497.567 | 549.826
Populus deltoides | LI1A |  | RACiR 300 to 800 | 393.4451 | 37.4966 | 24.9839 | 1000.28 | 0.0088713 | 0.39189 | 499.229 | 551.516
Populus deltoides | LI1A |  | RACiR 300 to 800 | 394.467 | 37.682 | 24.9813 | 999.928 | 0.00887 | 0.39199 | 500.726 | 553.224
Populus deltoides | LI1A |  | RACiR 300 to 800 | 395.9664 | 37.7478 | 24.9793 | 999.929 | 0.0088689 | 0.39207 | 502.405 | 555
Populus deltoides | LI1A |  | RACiR 300 to 800 | 397.4533 | 37.7563 | 24.9811 | 999.928 | 0.0088722 | 0.39216 | 503.927 | 556.544
Populus deltoides | LI1A |  | RACiR 300 to 800 | 398.5658 | 37.9254 | 24.9865 | 1000.29 | 0.0088714 | 0.39187 | 505.58 | 558.402
Populus deltoides | LI1A |  | RACiR 300 to 800 | 399.8045 | 38.0194 | 24.9885 | 999.936 | 0.0088735 | 0.39193 | 507.077 | 560.021
Populus deltoides | LI1A |  | RACiR 300 to 800 | 401.5872 | 37.9949 | 24.9883 | 999.939 | 0.0088818 | 0.39235 | 508.729 | 561.663
Populus deltoides | LI1A |  | RACiR 300 to 800 | 402.911 | 38.0369 | 24.9872 | 999.942 | 0.008879 | 0.39234 | 510.193 | 563.195
Populus deltoides | LI1A |  | RACiR 300 to 800 | 404.2076 | 38.1806 | 24.9863 | 1000.29 | 0.0088839 | 0.39261 | 511.819 | 564.999
Populus deltoides | LI1A |  | RACiR 300 to 800 | 405.7153 | 38.1681 | 24.9865 | 1000.28 | 0.008886 | 0.39268 | 513.313 | 566.49
Populus deltoides | LI1A |  | RACiR 300 to 800 | 406.9194 | 38.3222 | 24.9866 | 999.939 | 0.0088829 | 0.39252 | 514.981 | 568.344
Populus deltoides | LI1A |  | RACiR 300 to 800 | 408.0209 | 38.4646 | 24.9871 | 999.937 | 0.008881 | 0.3924 | 516.508 | 570.046
Populus deltoides | LI1A |  | RACiR 300 to 800 | 409.9958 | 38.3481 | 24.9868 | 999.938 | 0.0088818 | 0.39245 | 518.214 | 571.651
Populus deltoides | LI1A |  | RACiR 300 to 800 | 411.1357 | 38.4863 | 24.9871 | 1000.29 | 0.0088831 | 0.39252 | 519.718 | 573.315
Populus deltoides | LI1A |  | RACiR 300 to 800 | 412.5604 | 38.5306 | 24.991 | 999.938 | 0.00888 | 0.39216 | 521.387 | 575.064
Populus deltoides | LI1A |  | RACiR 300 to 800 | 414.0057 | 38.5234 | 24.9941 | 999.937 | 0.0088838 | 0.39214 | 522.856 | 576.536
Populus deltoides | LI1A |  | RACiR 300 to 800 | 415.1527 | 38.7397 | 24.992 | 999.946 | 0.0088854 | 0.39239 | 524.52 | 578.439
Populus deltoides | LI1A |  | RACiR 300 to 800 | 416.3702 | 38.8171 | 24.9929 | 999.949 | 0.0088817 | 0.39215 | 526.028 | 580.064
Populus deltoides | LI1A |  | RACiR 300 to 800 | 418.1733 | 38.7743 | 24.9927 | 1000.29 | 0.0088831 | 0.39224 | 527.738 | 581.75
Populus deltoides | LI1A |  | RACiR 300 to 800 | 419.3344 | 38.9246 | 24.9902 | 999.953 | 0.0088821 | 0.39237 | 529.276 | 583.454
Populus deltoides | LI1A |  | RACiR 300 to 800 | 420.6575 | 39.0449 | 24.9896 | 999.954 | 0.0088803 | 0.39232 | 530.952 | 585.292
Populus deltoides | LI1A |  | RACiR 300 to 800 | 422.4114 | 38.9654 | 24.9888 | 999.945 | 0.0088851 | 0.39261 | 532.463 | 586.731
Populus deltoides | LI1A |  | RACiR 300 to 800 | 423.6512 | 39.1175 | 24.9898 | 999.935 | 0.0088857 | 0.39261 | 534.124 | 588.573
Populus deltoides | LI1A |  | RACiR 300 to 800 | 424.8575 | 39.2039 | 24.9918 | 1000.29 | 0.0088852 | 0.39249 | 535.614 | 590.187
Populus deltoides | LI1A |  | RACiR 300 to 800 | 426.8538 | 39.0607 | 24.994 | 999.941 | 0.0088861 | 0.39238 | 537.319 | 591.76
Populus deltoides | LI1A |  | RACiR 300 to 800 | 427.9549 | 39.1829 | 24.9947 | 999.943 | 0.0088835 | 0.39219 | 538.812 | 593.403
Populus deltoides | LI1A |  | RACiR 300 to 800 | 429.5038 | 39.2256 | 24.9918 | 999.937 | 0.0088842 | 0.39238 | 540.452 | 595.111
Populus deltoides | LI1A |  | RACiR 300 to 800 | 431.0935 | 39.1636 | 24.9923 | 999.945 | 0.0088832 | 0.39227 | 541.949 | 596.562
Populus deltoides | LI1A |  | RACiR 300 to 800 | 432.2148 | 39.3349 | 24.9962 | 1000.31 | 0.0088843 | 0.3921 | 543.588 | 598.396
Populus deltoides | LI1A |  | RACiR 300 to 800 | 433.5357 | 39.3697 | 24.9982 | 999.947 | 0.0088827 | 0.39192 | 545.08 | 599.955
Populus deltoides | LI1A |  | RACiR 300 to 800 | 435.0523 | 39.4015 | 25.0007 | 999.949 | 0.0088826 | 0.39177 | 546.757 | 601.688
Populus deltoides | LI1A |  | RACiR 300 to 800 | 436.2887 | 39.5189 | 24.9998 | 999.95 | 0.0088862 | 0.392 | 548.263 | 603.326
Populus deltoides | LI1A |  | RACiR 300 to 800 | 437.6389 | 39.6052 | 25.0001 | 999.944 | 0.0088827 | 0.3918 | 549.918 | 605.105
Populus deltoides | LI1A |  | RACiR 300 to 800 | 438.9935 | 39.6644 | 24.9994 | 999.941 | 0.008883 | 0.39186 | 551.44 | 606.701
Populus deltoides | LI1A |  | RACiR 300 to 800 | 440.4174 | 39.7471 | 24.9995 | 999.943 | 0.0088823 | 0.39179 | 553.129 | 608.493
Populus deltoides | LI1A |  | RACiR 300 to 800 | 441.7969 | 39.7633 | 24.9988 | 999.942 | 0.0088757 | 0.39151 | 554.656 | 610.071
Populus deltoides | LI1A |  | RACiR 300 to 800 | 443.7846 | 39.665 | 24.9954 | 999.941 | 0.0088762 | 0.39174 | 556.368 | 611.698
Populus deltoides | LI1A |  | RACiR 300 to 800 | 445.1184 | 39.7273 | 24.9954 | 999.942 | 0.0088782 | 0.39181 | 557.876 | 613.285
Populus deltoides | LI1A |  | RACiR 300 to 800 | 446.5549 | 39.8603 | 24.9864 | 1000.29 | 0.0088745 | 0.39225 | 559.555 | 615.143
Populus deltoides | LI1A |  | RACiR 300 to 800 | 448.0039 | 39.9423 | 24.9814 | 999.939 | 0.0088788 | 0.3928 | 561.094 | 616.769
Populus deltoides | LI1A |  | RACiR 300 to 800 | 449.0431 | 40.1764 | 24.9819 | 999.939 | 0.0088783 | 0.39279 | 562.763 | 618.711
Populus deltoides | LI1A |  | RACiR 300 to 800 | 450.3437 | 40.2515 | 24.9827 | 999.941 | 0.0088786 | 0.39278 | 564.29 | 620.339
Populus deltoides | LI1A |  | RACiR 300 to 800 | 452.0788 | 40.1801 | 24.9897 | 999.938 | 0.0088804 | 0.39237 | 566.004 | 621.988
Populus deltoides | LI1A |  | RACiR 300 to 800 | 453.1985 | 40.2846 | 24.9931 | 1000.28 | 0.0088781 | 0.39202 | 567.518 | 623.639
Populus deltoides | LI1A |  | RACiR 300 to 800 | 454.9615 | 40.2593 | 24.9922 | 999.946 | 0.0088795 | 0.39214 | 569.224 | 625.354
Populus deltoides | LI1A |  | RACiR 300 to 800 | 456.6568 | 40.2092 | 24.9896 | 999.937 | 0.0088805 | 0.39237 | 570.76 | 626.855
Populus deltoides | LI1A |  | RACiR 300 to 800 | 457.7659 | 40.4134 | 24.992 | 999.942 | 0.0088823 | 0.39236 | 572.427 | 628.753
Populus deltoides | LI1A |  | RACiR 300 to 800 | 459.0555 | 40.457 | 24.9958 | 999.975 | 0.0088799 | 0.39201 | 573.961 | 630.362
Populus deltoides | LI1A |  | RACiR 300 to 800 | 460.6263 | 40.4531 | 24.9991 | 1000.28 | 0.0088803 | 0.39177 | 575.629 | 632.04
Populus deltoides | LI1A |  | RACiR 300 to 800 | 461.6166 | 40.612 | 25.0031 | 999.94 | 0.0088793 | 0.39142 | 577.155 | 633.748
Populus deltoides | LI1A |  | RACiR 300 to 800 | 462.9643 | 40.7068 | 25.0055 | 999.939 | 0.0088767 | 0.39114 | 578.859 | 635.584
Populus deltoides | LI1A |  | RACiR 300 to 800 | 464.5906 | 40.6495 | 25.0075 | 999.938 | 0.0088782 | 0.39107 | 580.398 | 637.077
Populus deltoides | LI1A |  | RACiR 300 to 800 | 465.9749 | 40.7646 | 25.0057 | 999.945 | 0.0088787 | 0.39119 | 582.075 | 638.892
Populus deltoides | LI1A |  | RACiR 300 to 800 | 467.386 | 40.7993 | 25.0034 | 999.941 | 0.0088771 | 0.39132 | 583.567 | 640.461
Populus deltoides | LI1A |  | RACiR 300 to 800 | 469.3408 | 40.7005 | 25.0012 | 999.943 | 0.0088786 | 0.39151 | 585.256 | 642.068
Populus deltoides | LI1A |  | RACiR 300 to 800 | 470.9887 | 40.6536 | 24.997 | 999.94 | 0.0088793 | 0.39181 | 586.731 | 643.519
Populus deltoides | LI1A |  | RACiR 300 to 800 | 473.7089 | 40.5603 | 24.9968 | 999.931 | 0.0088932 | 0.39247 | 589.095 | 645.837
Populus deltoides | LI1A |  | RACiR 300 to 800 | 473.9424 | 40.5158 | 24.9968 | 999.93 | 0.0089038 | 0.39294 | 589.095 | 645.837
Populus deltoides | LI1A |  | RACiR 300 to 800 | 476.8021 | 40.3978 | 24.995 | 999.945 | 0.0089182 | 0.39353 | 591.564 | 648.177
Populus deltoides | LI1A |  | RACiR 300 to 800 | 477.7293 | 40.6335 | 24.9937 | 999.938 | 0.0089288 | 0.39398 | 592.997 | 649.887
Populus deltoides | LI1A |  | RACiR 300 to 800 | 478.8178 | 40.8606 | 24.9929 | 999.942 | 0.0089379 | 0.39425 | 594.626 | 651.758
Populus deltoides | LI1A |  | RACiR 300 to 800 | 479.5066 | 41.1123 | 24.9928 | 1000.29 | 0.0089334 | 0.39383 | 596.097 | 653.507
Populus deltoides | LI1A |  | RACiR 300 to 800 | 480.2683 | 41.4105 | 24.9925 | 999.939 | 0.0089246 | 0.3933 | 597.79 | 655.535
Populus deltoides | LI1A |  | RACiR 300 to 800 | 481.8848 | 41.3443 | 24.9943 | 999.935 | 0.0089229 | 0.39294 | 599.378 | 657.06
Populus deltoides | LI1A |  | RACiR 300 to 800 | 483.4999 | 41.3242 | 24.9954 | 999.946 | 0.0089175 | 0.39247 | 601.114 | 658.808
Populus deltoides | LI1A |  | RACiR 300 to 800 | 485.1569 | 41.241 | 24.9936 | 999.941 | 0.0089115 | 0.39221 | 602.667 | 660.307
Populus deltoides | LI1A |  | RACiR 300 to 800 | 487.0618 | 41.1526 | 24.992 | 1000.27 | 0.0089117 | 0.39223 | 604.382 | 661.947
Populus deltoides | LI1A |  | RACiR 300 to 800 | 488.1781 | 41.2885 | 24.9884 | 999.935 | 0.008907 | 0.39214 | 605.895 | 663.622
Populus deltoides | LI1A |  | RACiR 300 to 800 | 489.5039 | 41.405 | 24.9853 | 999.934 | 0.0089006 | 0.39194 | 607.602 | 665.479
Populus deltoides | LI1A |  | RACiR 300 to 800 | 490.9694 | 41.4026 | 24.9866 | 999.94 | 0.0089007 | 0.39177 | 609.148 | 667.032
Populus deltoides | LI1A |  | RACiR 300 to 800 | 492.1619 | 41.5415 | 24.9879 | 999.953 | 0.0088939 | 0.39128 | 610.865 | 668.909
Populus deltoides | LI1A |  | RACiR 300 to 800 | 493.4781 | 41.5682 | 24.988 | 1000.29 | 0.0088829 | 0.39073 | 612.434 | 670.537
Populus deltoides | LI1A |  | RACiR 300 to 800 | 495.5281 | 41.4115 | 24.9859 | 999.941 | 0.0088772 | 0.39059 | 614.162 | 672.112
Populus deltoides | LI1A |  | RACiR 300 to 800 | 496.9358 | 41.4537 | 24.9822 | 999.946 | 0.0088703 | 0.39058 | 615.703 | 673.721
Populus deltoides | LI1A |  | RACiR 300 to 800 | 498.5691 | 41.4546 | 24.978 | 999.946 | 0.0088624 | 0.39049 | 617.393 | 675.437
Populus deltoides | LI1A |  | RACiR 300 to 800 | 500.0359 | 41.4832 | 24.9772 | 999.947 | 0.0088648 | 0.3907 | 618.906 | 676.989
Populus deltoides | LI1A |  | RACiR 300 to 800 | 501.2687 | 41.6424 | 24.9794 | 999.98 | 0.0088658 | 0.39068 | 620.588 | 678.857
Populus deltoides | LI1A |  | RACiR 300 to 800 | 502.461 | 41.7519 | 24.9804 | 999.977 | 0.0088651 | 0.39061 | 622.111 | 680.53
Populus deltoides | LI1A |  | RACiR 300 to 800 | 504.3401 | 41.7048 | 24.9811 | 999.957 | 0.0088716 | 0.39098 | 623.808 | 682.196
Populus deltoides | LI1A |  | RACiR 300 to 800 | 505.6498 | 41.7965 | 24.9815 | 999.938 | 0.0088732 | 0.39108 | 625.356 | 683.855
Populus deltoides | LI1A |  | RACiR 300 to 800 | 507.2173 | 41.8574 | 24.9788 | 999.949 | 0.0088715 | 0.39123 | 627.068 | 685.67
Populus deltoides | LI1A |  | RACiR 300 to 800 | 508.8512 | 41.8417 | 24.9769 | 999.929 | 0.0088735 | 0.39148 | 628.623 | 687.226
Populus deltoides | LI1A |  | RACiR 300 to 800 | 510.2392 | 41.9674 | 24.9786 | 1000.28 | 0.0088783 | 0.39164 | 630.326 | 689.075
Populus deltoides | LI1A |  | RACiR 300 to 800 | 511.639 | 41.9883 | 24.98 | 999.947 | 0.0088759 | 0.39149 | 631.853 | 690.667
Populus deltoides | LI1A |  | RACiR 300 to 800 | 513.5617 | 41.8863 | 24.9838 | 999.926 | 0.0088846 | 0.39168 | 633.51 | 692.235
Populus deltoides | LI1A |  | RACiR 300 to 800 | 515.1503 | 41.8189 | 24.988 | 999.927 | 0.0088851 | 0.39148 | 635.026 | 693.702
Populus deltoides | LI1A |  | RACiR 300 to 800 | 516.812 | 41.8011 | 24.9879 | 1000.28 | 0.0088828 | 0.39138 | 636.705 | 695.402
Populus deltoides | LI1A |  | RACiR 300 to 800 | 518.4943 | 41.7386 | 24.9873 | 999.937 | 0.0088851 | 0.39155 | 638.218 | 696.868
Populus deltoides | LI1A |  | RACiR 300 to 800 | 519.8928 | 41.8478 | 24.9856 | 999.934 | 0.0088829 | 0.39152 | 639.936 | 698.718
Populus deltoides | LI1A |  | RACiR 300 to 800 | 521.0098 | 42.0023 | 24.9823 | 999.937 | 0.0088802 | 0.39156 | 641.461 | 700.433
Populus deltoides | LI1A |  | RACiR 300 to 800 | 522.8107 | 41.9624 | 24.9824 | 1000.28 | 0.0088822 | 0.39162 | 643.184 | 702.126
Populus deltoides | LI1A |  | RACiR 300 to 800 | 523.9953 | 42.0452 | 24.9855 | 999.943 | 0.0088785 | 0.39119 | 644.737 | 703.785
Populus deltoides | LI1A |  | RACiR 300 to 800 | 525.461 | 42.0761 | 24.9887 | 999.948 | 0.0088739 | 0.39075 | 646.449 | 705.56
Populus deltoides | LI1A |  | RACiR 300 to 800 | 527.1903 | 41.9777 | 24.9909 | 999.932 | 0.0088746 | 0.39066 | 647.991 | 707.023
Populus deltoides | LI1A |  | RACiR 300 to 800 | 528.5765 | 42.0818 | 24.9913 | 999.926 | 0.0088734 | 0.39056 | 649.708 | 708.866
Populus deltoides | LI1A |  | RACiR 300 to 800 | 530.2486 | 42.0481 | 24.9878 | 1000.28 | 0.0088724 | 0.39078 | 651.263 | 710.424
Populus deltoides | LI1A |  | RACiR 300 to 800 | 532.1428 | 41.9948 | 24.9861 | 999.935 | 0.0088757 | 0.39103 | 652.987 | 712.1
Populus deltoides | LI1A |  | RACiR 300 to 800 | 533.3009 | 42.0981 | 24.9868 | 999.933 | 0.0088733 | 0.39083 | 654.498 | 713.74
Populus deltoides | LI1A |  | RACiR 300 to 800 | 534.6245 | 42.2004 | 24.9864 | 999.929 | 0.0088706 | 0.3907 | 656.155 | 715.534
Populus deltoides | LI1A |  | RACiR 300 to 800 | 535.8775 | 42.2718 | 24.9882 | 1000.29 | 0.0088724 | 0.39065 | 657.641 | 717.102
Populus deltoides | LI1A |  | RACiR 300 to 800 | 537.1954 | 42.3943 | 24.991 | 999.939 | 0.0088742 | 0.39057 | 659.334 | 718.949
Populus deltoides | LI1A |  | RACiR 300 to 800 | 538.606 | 42.4272 | 24.9898 | 999.933 | 0.0088703 | 0.39048 | 660.885 | 720.573
Populus deltoides | LI1A |  | RACiR 300 to 800 | 540.5811 | 42.3529 | 24.9878 | 999.929 | 0.0088751 | 0.39081 | 662.614 | 722.233
Populus deltoides | LI1A |  | RACiR 300 to 800 | 542.1919 | 42.3362 | 24.9871 | 999.932 | 0.008874 | 0.39085 | 664.207 | 723.84
Populus deltoides | LI1A |  | RACiR 300 to 800 | 543.9606 | 42.325 | 24.9839 | 999.954 | 0.0088707 | 0.39086 | 665.978 | 725.635
Populus deltoides | LI1A |  | RACiR 300 to 800 | 545.7254 | 42.2254 | 24.9856 | 999.948 | 0.0088746 | 0.39093 | 667.509 | 727.082
Populus deltoides | LI1A |  | RACiR 300 to 800 | 547.3317 | 42.2714 | 24.9885 | 999.936 | 0.0088791 | 0.39104 | 669.244 | 728.881
Populus deltoides | LI1A |  | RACiR 300 to 800 | 548.5933 | 42.3588 | 24.9883 | 999.94 | 0.0088752 | 0.39086 | 670.812 | 730.57
Populus deltoides | LI1A |  | RACiR 300 to 800 | 550.2976 | 42.34 | 24.9876 | 999.975 | 0.0088766 | 0.39096 | 672.477 | 732.245
Populus deltoides | LI1A |  | RACiR 300 to 800 | 551.6412 | 42.4385 | 24.9866 | 999.945 | 0.0088799 | 0.39122 | 674.028 | 733.909
Populus deltoides | LI1A |  | RACiR 300 to 800 | 553.3557 | 42.4113 | 24.9882 | 999.939 | 0.0088751 | 0.3909 | 675.807 | 735.692
Populus deltoides | LI1A |  | RACiR 300 to 800 | 554.9948 | 42.3603 | 24.9879 | 999.945 | 0.0088733 | 0.39079 | 677.384 | 737.24
Populus deltoides | LI1A |  | RACiR 300 to 800 | 556.2719 | 42.4999 | 24.9882 | 999.934 | 0.00887 | 0.39054 | 679.123 | 739.142
Populus deltoides | LI1A |  | RACiR 300 to 800 | 557.7526 | 42.5131 | 24.9872 | 999.976 | 0.0088688 | 0.39052 | 680.678 | 740.735
Populus deltoides | LI1A |  | RACiR 300 to 800 | 559.4903 | 42.5306 | 24.9809 | 999.945 | 0.0088695 | 0.39088 | 682.386 | 742.48
Populus deltoides | LI1A |  | RACiR 300 to 800 | 560.8078 | 42.5779 | 24.9828 | 999.943 | 0.0088715 | 0.39088 | 683.862 | 744.035
Populus deltoides | LI1A |  | RACiR 300 to 800 | 562.6613 | 42.4849 | 24.9861 | 999.935 | 0.0088732 | 0.39078 | 685.551 | 745.656
Populus deltoides | LI1A |  | RACiR 300 to 800 | 564.4599 | 42.3539 | 24.9872 | 999.925 | 0.0088725 | 0.39066 | 687.089 | 747.075
Populus deltoides | LI1A |  | RACiR 300 to 800 | 565.6929 | 42.4479 | 24.9928 | 1000.28 | 0.0088682 | 0.39006 | 688.777 | 748.894
Populus deltoides | LI1A |  | RACiR 300 to 800 | 567.1716 | 42.4602 | 24.9917 | 999.935 | 0.0088644 | 0.38994 | 690.354 | 750.505
Populus deltoides | LI1A |  | RACiR 300 to 800 | 568.761 | 42.5539 | 24.9872 | 999.928 | 0.0088684 | 0.39042 | 692.077 | 752.335
Populus deltoides | LI1A |  | RACiR 300 to 800 | 569.8205 | 42.75 | 24.9854 | 999.93 | 0.0088689 | 0.39057 | 693.62 | 754.103
Populus deltoides | LI1A |  | RACiR 300 to 800 | 571.4167 | 42.791 | 24.9843 | 1000.28 | 0.0088654 | 0.3905 | 695.375 | 755.932
Populus deltoides | LI1A |  | RACiR 300 to 800 | 573.1292 | 42.7465 | 24.9828 | 999.926 | 0.0088701 | 0.39082 | 696.919 | 757.45
Populus deltoides | LI1A |  | RACiR 300 to 800 | 574.9767 | 42.6994 | 24.9806 | 999.924 | 0.0088692 | 0.39093 | 698.65 | 759.156
Populus deltoides | LI1A |  | RACiR 300 to 800 | 576.7952 | 42.5345 | 24.9799 | 999.923 | 0.0088573 | 0.39043 | 700.222 | 760.593
Populus deltoides | LI1A |  | RACiR 300 to 800 | 578.6132 | 42.4309 | 24.9775 | 1000.25 | 0.0088469 | 0.39013 | 701.892 | 762.16
Populus deltoides | LI1A |  | RACiR 300 to 800 | 579.7262 | 42.4937 | 24.98 | 999.907 | 0.0088333 | 0.38942 | 703.398 | 763.749
Populus deltoides | LI1A |  | RACiR 300 to 800 | 580.5787 | 42.6971 | 24.9872 | 999.921 | 0.0088192 | 0.38843 | 705.095 | 765.67
Populus deltoides | LI1A |  | RACiR 300 to 800 | 581.852 | 42.8049 | 24.9891 | 999.905 | 0.0088217 | 0.3886 | 706.628 | 767.323
Populus deltoides | LI1A |  | RACiR 300 to 800 | 583.4346 | 42.8973 | 24.9889 | 999.916 | 0.008824 | 0.3889 | 708.395 | 769.222
Populus deltoides | LI1A |  | RACiR 300 to 800 | 585.1406 | 42.8796 | 24.9869 | 999.922 | 0.0088264 | 0.38925 | 709.984 | 770.82
Populus deltoides | LI1A |  | RACiR 300 to 800 | 587.2862 | 42.7742 | 24.9849 | 999.912 | 0.0088327 | 0.38979 | 711.742 | 772.49
Populus deltoides | LI1A |  | RACiR 300 to 800 | 588.89 | 42.7343 | 24.9853 | 999.914 | 0.0088321 | 0.38979 | 713.278 | 774.024
Populus deltoides | LI1A |  | RACiR 300 to 800 | 590.6437 | 42.7356 | 24.9849 | 999.922 | 0.008838 | 0.39017 | 714.96 | 775.73
Populus deltoides | LI1A |  | RACiR 300 to 800 | 592.3302 | 42.6704 | 24.9869 | 999.921 | 0.0088391 | 0.39016 | 716.523 | 777.248
Populus deltoides | LI1A |  | RACiR 300 to 800 | 593.5688 | 42.8197 | 24.9901 | 1000.27 | 0.0088394 | 0.39 | 718.23 | 779.142
Populus deltoides | LI1A |  | RACiR 300 to 800 | 594.7161 | 42.9565 | 24.9929 | 999.919 | 0.0088397 | 0.38983 | 719.812 | 780.882
Populus deltoides | LI1A |  | RACiR 300 to 800 | 596.3873 | 42.9428 | 24.9981 | 999.925 | 0.008841 | 0.38956 | 721.576 | 782.652
Populus deltoides | LI1A |  | RACiR 300 to 800 | 597.672 | 43.0444 | 24.9991 | 1000.27 | 0.0088432 | 0.38963 | 723.137 | 784.34
Populus deltoides | LI1A |  | RACiR 300 to 800 | 599.407 | 43.0202 | 24.9993 | 999.907 | 0.008842 | 0.38954 | 724.876 | 786.082
Populus deltoides | LI1A | 3 | Standard | 277.3052 | 27.2096 | 25.033 | 999.995 | 0.0071786 | 0.31794 | 370.197 | 400.065
Populus deltoides | LI1A | 3 | Standard | 239.7984 | 24.4077 | 25.01 | 999.968 | 0.0071483 | 0.31645 | 323.288 | 350.006
Populus deltoides | LI1A | 3 | Standard | 204.3874 | 21.1304 | 25.0287 | 999.95 | 0.0071402 | 0.31486 | 276.956 | 300.064
Populus deltoides | LI1A | 3 | Standard | 171.7776 | 17.2536 | 25.0136 | 999.919 | 0.0071191 | 0.31476 | 231.149 | 250.049
Populus deltoides | LI1A | 3 | Standard | 139.7045 | 13.2626 | 25.0224 | 999.9 | 0.0071313 | 0.31513 | 185.47 | 200.055
Populus deltoides | LI1A | 3 | Standard | 108.6515 | 9.0329 | 25.0202 | 999.895 | 0.0071531 | 0.31617 | 140.034 | 150.069
Populus deltoides | LI1A | 3 | Standard | 78.3641 | 4.6153 | 25.0114 | 1000.18 | 0.0072209 | 0.31997 | 94.7416 | 100.041
Populus deltoides | LI1A | 3 | Standard | 48.2866 | 0.078108 | 25.0064 | 1000.18 | 0.0072508 | 0.32151 | 49.6337 | 50.0717
Populus deltoides | LI1A | 3 | Standard | 274.5319 | 28.2767 | 25.027 | 999.903 | 0.007349 | 0.32398 | 369.111 | 400.1
Populus deltoides | LI1A | 3 | Standard | 349.2953 | 34.0223 | 25.0246 | 1000.27 | 0.0073971 | 0.32662 | 462.655 | 500.098
Populus deltoides | LI1A | 3 | Standard | 425.5815 | 39.063 | 25.0357 | 999.972 | 0.0073325 | 0.32484 | 556.925 | 600.07
Populus deltoides | LI1A | 3 | Standard | 506.0858 | 42.5999 | 25.0434 | 1000 | 0.0072331 | 0.31895 | 652.789 | 700.113
Populus deltoides | LI1A | 3 | Standard | 590.693 | 44.7167 | 25.0382 | 999.988 | 0.0070582 | 0.31024 | 750.082 | 800.093
Populus deltoides | LI1A | 3 | Standard | 674.0639 | 46.4024 | 25.0596 | 999.988 | 0.0067562 | 0.29638 | 847.978 | 900.108
Populus deltoides | LI1A | 3 | Standard | 760.1149 | 47.4272 | 25.0402 | 999.967 | 0.006448 | 0.28398 | 946.496 | 1000.03
Populus deltoides | LI1A | 3 | Standard | 847.169 | 48.2898 | 25.0123 | 999.98 | 0.0062373 | 0.27368 | 1045.18 | 1099.99
Populus deltoides | LI1A | 2 | Standard | 249.596 | 26.7424 | 25.0268 | 1000.23 | 0.0053126 | 0.23341 | 371.232 | 399.946
Populus deltoides | LI1A | 2 | Standard | 249.2926 | 26.8108 | 25.0281 | 1000.22 | 0.0053126 | 0.23337 | 371.24 | 400.023
Populus deltoides | LI1A | 2 | Standard | 252.7516 | 27.5925 | 24.9979 | 1000.02 | 0.0056671 | 0.24972 | 370.317 | 400.007
Populus deltoides | LI1A | 2 | Standard | 222.3867 | 24.43 | 25.0019 | 999.964 | 0.0058083 | 0.2566 | 323.774 | 350.084
Populus deltoides | LI1A | 2 | Standard | 192.2301 | 21.0433 | 25.0022 | 999.916 | 0.0059826 | 0.26382 | 277.317 | 300.02
Populus deltoides | LI1A | 2 | Standard | 163.6637 | 17.1897 | 24.9994 | 1000.14 | 0.0061333 | 0.27139 | 231.468 | 250.078
Populus deltoides | LI1A | 2 | Standard | 134.2878 | 13.3496 | 25.003 | 1000.13 | 0.0063489 | 0.28039 | 185.519 | 200.046
Populus deltoides | LI1A | 2 | Standard | 105.7789 | 9.1097 | 25.0041 | 1000.08 | 0.0065572 | 0.28941 | 140.041 | 150.069
Populus deltoides | LI1A | 2 | Standard | 76.9275 | 4.7 | 24.9752 | 1000.01 | 0.0067464 | 0.29792 | 94.6465 | 99.9847
Populus deltoides | LI1A | 2 | Standard | 48.404 | 0.0253 | 24.9921 | 1000.01 | 0.0069702 | 0.30751 | 49.597 | 49.968
Populus deltoides | LI1A | 2 | Standard | 271.7587 | 28.4586 | 24.8841 | 1000.06 | 0.0071737 | 0.31653 | 368.928 | 400.032
Populus deltoides | LI1A | 2 | Standard | 348.4607 | 34.2237 | 25.011 | 1000.05 | 0.0073576 | 0.32668 | 462.354 | 499.981
Populus deltoides | LI1A | 2 | Standard | 426.9656 | 39.4167 | 25.0118 | 1000.12 | 0.0074775 | 0.33278 | 556.461 | 600.04
Populus deltoides | LI1A | 2 | Standard | 512.1639 | 42.6098 | 25.0077 | 1000.12 | 0.0075705 | 0.33505 | 652.495 | 700.046
Populus deltoides | LI1A | 2 | Standard | 602.9458 | 44.6774 | 25.009 | 1000.08 | 0.0076377 | 0.33969 | 749.675 | 800.079
Populus deltoides | LI1A | 2 | Standard | 693.8559 | 46.2732 | 25.0111 | 1000.08 | 0.0077166 | 0.34054 | 847.2 | 900.016
Populus deltoides | LI1A | 2 | Standard | 785.2388 | 47.5152 | 25.0367 | 1000.05 | 0.0076281 | 0.33785 | 945.416 | 1000.14
Populus deltoides | LI1A | 2 | Standard | 878.5589 | 48.5546 | 24.9994 | 1000.04 | 0.0076631 | 0.339 | 1043.51 | 1100.06
Populus deltoides | LI1A | 2 | Standard | 972.5841 | 49.1025 | 25.021 | 1000.06 | 0.0075971 | 0.33648 | 1142.39 | 1200.17
Populus deltoides | LI1A | 2 | Standard | 1066.6574 | 49.8541 | 25.0141 | 1000.02 | 0.0076232 | 0.33685 | 1240.77 | 1300.08
Populus deltoides | LI1A | 2 | Standard | 1161.5909 | 50.2353 | 25.0236 | 999.711 | 0.0075675 | 0.33513 | 1339.73 | 1400.1
Populus deltoides | LI1A | 2 | Standard | 1256.5482 | 50.6486 | 25.0149 | 1000.05 | 0.0075568 | 0.33402 | 1438.67 | 1500.19
Populus deltoides | LI1A | 2 | Standard | 1544.2267 | 51.2758 | 25.0181 | 999.993 | 0.0074927 | 0.33194 | 1735.72 | 1800
Populus deltoides | LI1A | 2 | Standard | 1731.5335 | 52.3355 | 25.0034 | 1000.03 | 0.0073462 | 0.32578 | 1933.5 | 2000.04
Populus deltoides | LI1B | 1 | Standard | 272.5545 | 22.1549 | 25.007 | 1000.14 | 0.0052245 | 0.23071 | 375.926 | 400.044
Populus deltoides | LI1B | 1 | Standard | 238.946 | 19.7579 | 25.0121 | 1000.14 | 0.0053664 | 0.2375 | 328.547 | 350.069
Populus deltoides | LI1B | 1 | Standard | 206.8285 | 16.9485 | 24.999 | 1000.13 | 0.0055268 | 0.2449 | 281.546 | 300.051
Populus deltoides | LI1B | 1 | Standard | 175.5897 | 13.7748 | 24.9898 | 999.818 | 0.0056705 | 0.25187 | 234.901 | 250.009
Populus deltoides | LI1B | 1 | Standard | 143.3036 | 10.6809 | 25.0049 | 1000.13 | 0.0058398 | 0.25898 | 188.285 | 200.066
Populus deltoides | LI1B | 1 | Standard | 112.0512 | 7.1886 | 25.0082 | 999.809 | 0.0059949 | 0.26609 | 141.928 | 149.968
Populus deltoides | LI1B | 1 | Standard | 80.4859 | 3.6532 | 25.0051 | 1000.12 | 0.0061799 | 0.27434 | 95.7879 | 100.033
Populus deltoides | LI1B | 1 | Standard | 48.6618 | -0.018448 | 25.0031 | 1000.15 | 0.0063248 | 0.28097 | 49.7032 | 49.9991
Populus deltoides | LI1B | 1 | Standard | 294.8222 | 21.4843 | 25.0279 | 1000.1 | 0.0066327 | 0.29204 | 376.006 | 399.983
Populus deltoides | LI1B | 1 | Standard | 378.6139 | 25.0443 | 25.0082 | 999.83 | 0.0067548 | 0.29928 | 471.894 | 500.125
Populus deltoides | LI1B | 1 | Standard | 467.1913 | 27.4271 | 25 | 999.865 | 0.0069284 | 0.30539 | 568.754 | 600.122
Populus deltoides | LI1B | 1 | Standard | 559.83 | 28.3717 | 25.0388 | 999.864 | 0.0068822 | 0.30386 | 667.096 | 700.061
Populus deltoides | LI1B | 1 | Standard | 655.8371 | 28.4054 | 25.0063 | 999.884 | 0.0068017 | 0.30068 | 766.395 | 800.014
Populus deltoides | LI1B | 1 | Standard | 750.6466 | 28.5306 | 25.0308 | 999.879 | 0.0066707 | 0.29455 | 865.812 | 900.118
Populus deltoides | LI1B | 1 | Standard | 848.0417 | 28.0887 | 25.0248 | 999.879 | 0.0065477 | 0.28937 | 965.63 | 1000.04
Populus deltoides | LI1B | 1 | Standard | 943.4214 | 28.0855 | 25.0259 | 999.895 | 0.0064171 | 0.28362 | 1065.17 | 1100.09
Populus deltoides | LI1B | 1 | Standard | 1040.2417 | 27.8414 | 25.0155 | 999.862 | 0.0063118 | 0.27937 | 1164.81 | 1200
Populus deltoides | LI1B | 1 | Standard | 1135.7497 | 27.958 | 25.0389 | 1000.18 | 0.0062515 | 0.27584 | 1264.3 | 1300.16
Populus deltoides | LI1B | 1 | Standard | 1232.2811 | 27.6922 | 25.0333 | 999.832 | 0.0061504 | 0.27105 | 1363.9 | 1399.98
Populus deltoides | LI1B | 1 | Standard | 1326.2981 | 28.1889 | 25.0258 | 1000.15 | 0.0060696 | 0.26766 | 1463.24 | 1500.31
Populus deltoides | LI1B | 1 | Standard | 1614.4364 | 28.3405 | 25.0187 | 999.835 | 0.0059325 | 0.26128 | 1761.23 | 1800.02
Populus deltoides | LI1B | 1 | Standard | 1805.7931 | 28.538 | 25.0217 | 1000.11 | 0.0058106 | 0.25577 | 1960.15 | 2000.08
Populus deltoides | LI1B | 2 | Standard | 272.362 | 22.8602 | 24.9793 | 1000.14 | 0.0054305 | 0.23956 | 375.127 | 400.026
Populus deltoides | LI1B | 2 | Standard | 239.7075 | 20.1183 | 25.0074 | 999.846 | 0.005574 | 0.24549 | 328.105 | 350.052
Populus deltoides | LI1B | 2 | Standard | 205.455 | 17.3082 | 25.0173 | 1000.13 | 0.0056139 | 0.24652 | 281.206 | 300.093
Populus deltoides | LI1B | 2 | Standard | 173.5725 | 13.9075 | 25.0152 | 1000.1 | 0.0056059 | 0.24571 | 234.832 | 250.056
Populus deltoides | LI1B | 2 | Standard | 140.5222 | 10.754 | 25.0164 | 1000.12 | 0.0055816 | 0.24446 | 188.267 | 200.072
Populus deltoides | LI1B | 2 | Standard | 109.8192 | 7.2218 | 25.0175 | 999.793 | 0.0056085 | 0.24616 | 142.026 | 150.044
Populus deltoides | LI1B | 2 | Standard | 79.2709 | 3.6349 | 25.0104 | 1000.11 | 0.0056716 | 0.24935 | 95.8396 | 100.018
Populus deltoides | LI1B | 2 | Standard | 48.7379 | -0.0047833 | 25.0172 | 1000.1 | 0.0057338 | 0.25219 | 49.8396 | 50.1206
Populus deltoides | LI1B | 2 | Standard | 282.7732 | 22.0474 | 25.0263 | 999.779 | 0.0058953 | 0.25788 | 375.797 | 400.06
Populus deltoides | LI1B | 2 | Standard | 365.4277 | 25.4924 | 25.0171 | 999.814 | 0.0060023 | 0.26328 | 471.796 | 500.122
Populus deltoides | LI1B | 2 | Standard | 450.2257 | 28.2829 | 25.0267 | 1000.11 | 0.0060594 | 0.26537 | 568.436 | 600.166
Populus deltoides | LI1B | 2 | Standard | 539.7569 | 29.797 | 25.0277 | 1000.11 | 0.0060102 | 0.26398 | 666.363 | 700.164
Populus deltoides | LI1B | 2 | Standard | 632.6761 | 30.5292 | 25.0225 | 999.822 | 0.005989 | 0.26273 | 764.804 | 799.914
Populus deltoides | LI1B | 2 | Standard | 724.6477 | 31.2985 | 25.0212 | 1000.12 | 0.005916 | 0.25883 | 863.721 | 900.13
Populus deltoides | LI1B | 2 | Standard | 819.7115 | 31.4421 | 25.0118 | 999.836 | 0.0058518 | 0.25585 | 962.992 | 1000.07
Populus deltoides | LI1B | 2 | Standard | 914.1774 | 31.7999 | 25.0143 | 999.836 | 0.0057884 | 0.25354 | 1062.16 | 1100.11
Populus deltoides | LI1B | 2 | Standard | 1009.1853 | 31.845 | 25.0171 | 999.825 | 0.0057204 | 0.2498 | 1161.52 | 1200.01
Populus deltoides | LI1B | 2 | Standard | 1103.5643 | 32.1104 | 25.0131 | 1000.16 | 0.0056368 | 0.24632 | 1260.98 | 1300.2
Populus deltoides | LI1B | 2 | Standard | 1200.3722 | 32.1255 | 25.0111 | 1000.2 | 0.0056254 | 0.24587 | 1360.33 | 1400.11
Populus deltoides | LI1B | 2 | Standard | 1293.3459 | 32.4173 | 25.0136 | 1000.21 | 0.0055176 | 0.24068 | 1459.59 | 1500.06
Populus deltoides | LI1B | 2 | Standard | 1580.2329 | 33.1186 | 25.0072 | 999.872 | 0.0054565 | 0.23809 | 1757.58 | 1800.29
Populus deltoides | LI1B | 2 | Standard | 1771.839 | 33.1464 | 25.0181 | 999.909 | 0.0053578 | 0.23345 | 1956.61 | 2000.24
Populus deltoides | LI1B | 3 | Standard | 285.7126 | 19.2383 | 25.0088 | 999.898 | 0.0050909 | 0.22479 | 378.822 | 399.988
Populus deltoides | LI1B | 3 | Standard | 242.0867 | 18.2294 | 25.0065 | 1000.24 | 0.0051147 | 0.22347 | 330.212 | 350.131
Populus deltoides | LI1B | 3 | Standard | 205.9551 | 15.8715 | 25.022 | 1000.23 | 0.0051048 | 0.22313 | 282.675 | 299.99
Populus deltoides | LI1B | 3 | Standard | 173.3866 | 12.9114 | 25.009 | 999.886 | 0.005118 | 0.22329 | 235.901 | 250.019
Populus deltoides | LI1B | 3 | Standard | 140.0205 | 10.1149 | 25 | 999.941 | 0.0051186 | 0.22365 | 189.012 | 200.094
Populus deltoides | LI1B | 3 | Standard | 109.4068 | 6.8027 | 25.0131 | 999.896 | 0.0051609 | 0.22508 | 142.518 | 150.056
Populus deltoides | LI1B | 3 | Standard | 78.8129 | 3.4827 | 24.9999 | 999.917 | 0.0051695 | 0.22783 | 96.0835 | 100.063
Populus deltoides | LI1B | 3 | Standard | 48.3898 | 0.053146 | 25.0014 | 999.965 | 0.0052007 | 0.22944 | 49.7335 | 50.0453
Populus deltoides | LI1B | 3 | Standard | 275.5298 | 21.9243 | 25.0127 | 999.903 | 0.005406 | 0.2357 | 376.018 | 399.973
Populus deltoides | LI1B | 3 | Standard | 354.561 | 25.6874 | 25.0137 | 1000.29 | 0.0054506 | 0.23824 | 471.838 | 500.096
Populus deltoides | LI1B | 3 | Standard | 436.4423 | 28.6332 | 25.0201 | 999.99 | 0.0054499 | 0.23784 | 568.34 | 600.07
Populus deltoides | LI1B | 3 | Standard | 523.079 | 30.3272 | 25.0373 | 999.994 | 0.0053345 | 0.23405 | 666.21 | 700.09
Populus deltoides | LI1B | 3 | Standard | 612.4202 | 31.1888 | 25.0168 | 1000 | 0.0051881 | 0.22794 | 764.925 | 800.083
Populus deltoides | LI1B | 3 | Standard | 701.578 | 31.803 | 25.0244 | 999.993 | 0.0050402 | 0.22024 | 863.896 | 900.054
Populus deltoides | LI1B | 3 | Standard | 791.7694 | 32.0634 | 25.0241 | 999.999 | 0.0048447 | 0.21165 | 963.35 | 1000.08
Populus deltoides | LI1B | 3 | Standard | 884.5216 | 32.3566 | 25.0116 | 999.995 | 0.0047356 | 0.20756 | 1062.62 | 1100.01
Populus deltoides | LI1B | 3 | Standard | 977.3975 | 32.5658 | 25.0196 | 999.987 | 0.0046248 | 0.203 | 1162.19 | 1200.13
Populus deltoides | LI1B | 3 | Standard | 1072.2209 | 32.6143 | 25.017 | 999.94 | 0.0045639 | 0.20023 | 1261.7 | 1300.07
Populus deltoides | LI1B | 3 | Standard | 1166.6802 | 32.5318 | 25.0202 | 999.968 | 0.0044711 | 0.19612 | 1361.37 | 1399.99
Populus deltoides | LI1B | 3 | Standard | 1256.6372 | 33.098 | 25.0207 | 999.939 | 0.004367 | 0.19127 | 1460.7 | 1500.18
Populus deltoides | LI1B | 3 | Standard | 1545.7 | 33.0075 | 25.0254 | 999.985 | 0.0042864 | 0.1876 | 1759.41 | 1799.96
Populus deltoides | LI1B | 3 | Standard | 1734.1174 | 33.4336 | 25.0252 | 999.98 | 0.0041956 | 0.18363 | 1958.37 | 2000.02
Populus deltoides | LI1B |  | RACiR 800 to 300 | 639.3806 | 12.3759 | 25.0084 | 999.934 | 0.0070326 | 0.30789 | 694.813 | 722.844
Populus deltoides | LI1B |  | RACiR 800 to 300 | 638.2459 | 12.2677 | 25.0099 | 999.954 | 0.0070447 | 0.30835 | 693.24 | 721.15
Populus deltoides | LI1B |  | RACiR 800 to 300 | 636.6889 | 12.2644 | 25.0061 | 999.946 | 0.0070635 | 0.30936 | 691.496 | 719.392
Populus deltoides | LI1B |  | RACiR 800 to 300 | 634.9278 | 12.3663 | 25.0033 | 999.916 | 0.0070814 | 0.3102 | 689.915 | 717.901
Populus deltoides | LI1B |  | RACiR 800 to 300 | 633.4464 | 12.3157 | 25.0062 | 999.905 | 0.0071004 | 0.31076 | 688.176 | 716.098
Populus deltoides | LI1B |  | RACiR 800 to 300 | 631.9076 | 12.3177 | 25.0055 | 999.946 | 0.0071057 | 0.31086 | 686.602 | 714.507
Populus deltoides | LI1B |  | RACiR 800 to 300 | 629.7477 | 12.4227 | 25.0112 | 999.944 | 0.0071019 | 0.31017 | 684.848 | 712.827
Populus deltoides | LI1B |  | RACiR 800 to 300 | 628.2945 | 12.3854 | 25.0139 | 999.91 | 0.0070943 | 0.30954 | 683.338 | 711.253
Populus deltoides | LI1B |  | RACiR 800 to 300 | 626.7252 | 12.3569 | 25.0084 | 999.943 | 0.0070867 | 0.30931 | 681.663 | 709.52
Populus deltoides | LI1B |  | RACiR 800 to 300 | 624.8312 | 12.4599 | 25.0088 | 999.933 | 0.0070824 | 0.309 | 680.11 | 708.042
Populus deltoides | LI1B |  | RACiR 800 to 300 | 623.6204 | 12.316 | 25.0036 | 999.94 | 0.0070711 | 0.30864 | 678.439 | 706.196
Populus deltoides | LI1B |  | RACiR 800 to 300 | 622.2193 | 12.2644 | 25.0006 | 1000.25 | 0.0070644 | 0.30842 | 676.862 | 704.539
Populus deltoides | LI1B |  | RACiR 800 to 300 | 620.1599 | 12.3463 | 25.0004 | 999.935 | 0.0070537 | 0.30789 | 675.096 | 702.824
Populus deltoides | LI1B |  | RACiR 800 to 300 | 618.7403 | 12.3042 | 24.9972 | 999.937 | 0.007046 | 0.30769 | 673.525 | 701.183
Populus deltoides | LI1B |  | RACiR 800 to 300 | 616.6604 | 12.4005 | 24.9992 | 1000.25 | 0.007042 | 0.30743 | 671.752 | 699.48
Populus deltoides | LI1B |  | RACiR 800 to 300 | 614.8637 | 12.491 | 24.9982 | 999.944 | 0.0070417 | 0.3075 | 670.199 | 697.994
Populus deltoides | LI1B |  | RACiR 800 to 300 | 613.2789 | 12.4625 | 24.9983 | 999.941 | 0.0070434 | 0.30762 | 668.466 | 696.207
Populus deltoides | LI1B |  | RACiR 800 to 300 | 611.836 | 12.4375 | 25.0024 | 999.94 | 0.0070451 | 0.30752 | 666.929 | 694.623
Populus deltoides | LI1B |  | RACiR 800 to 300 | 609.6564 | 12.5766 | 25.0058 | 1000.25 | 0.0070449 | 0.30731 | 665.194 | 693.003
Populus deltoides | LI1B |  | RACiR 800 to 300 | 608.5542 | 12.4417 | 25.0095 | 999.942 | 0.0070469 | 0.30724 | 663.639 | 691.29
Populus deltoides | LI1B |  | RACiR 800 to 300 | 606.9573 | 12.4111 | 25.0113 | 999.943 | 0.0070453 | 0.30707 | 661.931 | 689.527
Populus deltoides | LI1B |  | RACiR 800 to 300 | 605.1606 | 12.4873 | 25.0093 | 999.953 | 0.0070456 | 0.30717 | 660.326 | 687.974
Populus deltoides | LI1B |  | RACiR 800 to 300 | 604.2938 | 12.2513 | 25.0049 | 999.916 | 0.0070447 | 0.30735 | 658.628 | 686.012
Populus deltoides | LI1B |  | RACiR 800 to 300 | 602.6303 | 12.292 | 25.0008 | 999.911 | 0.0070425 | 0.30742 | 657.044 | 684.446
Populus deltoides | LI1B |  | RACiR 800 to 300 | 600.3349 | 12.4628 | 24.9987 | 999.943 | 0.0070419 | 0.30747 | 655.246 | 682.792
Populus deltoides | LI1B |  | RACiR 800 to 300 | 599.1175 | 12.3803 | 24.9972 | 999.943 | 0.0070419 | 0.30753 | 653.718 | 681.158
Populus deltoides | LI1B |  | RACiR 800 to 300 | 597.2714 | 12.4275 | 24.9962 | 1000.26 | 0.007041 | 0.30753 | 651.983 | 679.447
Populus deltoides | LI1B |  | RACiR 800 to 300 | 595.4579 | 12.5164 | 24.9985 | 999.946 | 0.0070419 | 0.30745 | 650.436 | 677.965
Populus deltoides | LI1B |  | RACiR 800 to 300 | 594.3265 | 12.3621 | 24.9971 | 999.949 | 0.0070437 | 0.30757 | 648.754 | 676.103
Populus deltoides | LI1B |  | RACiR 800 to 300 | 593.0197 | 12.2952 | 24.9973 | 999.94 | 0.0070435 | 0.30752 | 647.204 | 674.463
Populus deltoides | LI1B |  | RACiR 800 to 300 | 591.1084 | 12.3512 | 24.9999 | 1000.26 | 0.0070414 | 0.30729 | 645.47 | 672.758
Populus deltoides | LI1B |  | RACiR 800 to 300 | 589.8875 | 12.2322 | 25.0003 | 999.94 | 0.0070268 | 0.3066 | 643.923 | 671.06
Populus deltoides | LI1B |  | RACiR 800 to 300 | 588.2725 | 12.1845 | 24.9975 | 999.956 | 0.0070082 | 0.30592 | 642.195 | 669.247
Populus deltoides | LI1B |  | RACiR 800 to 300 | 586.1741 | 12.3436 | 24.9941 | 999.962 | 0.0069925 | 0.30546 | 640.625 | 667.803
Populus deltoides | LI1B |  | RACiR 800 to 300 | 584.4672 | 12.3335 | 24.9902 | 999.918 | 0.0069773 | 0.30508 | 638.883 | 666.017
Populus deltoides | LI1B |  | RACiR 800 to 300 | 583.1388 | 12.2944 | 24.987 | 999.921 | 0.0069757 | 0.3053 | 637.352 | 664.427
Populus deltoides | LI1B |  | RACiR 800 to 300 | 580.74 | 12.5131 | 24.9919 | 999.952 | 0.006979 | 0.30535 | 635.62 | 662.891
Populus deltoides | LI1B |  | RACiR 800 to 300 | 579.3805 | 12.475 | 24.9965 | 999.957 | 0.0069846 | 0.30548 | 634.09 | 661.303
Populus deltoides | LI1B |  | RACiR 800 to 300 | 577.86 | 12.4304 | 25.0038 | 999.954 | 0.0069888 | 0.30542 | 632.407 | 659.553
Populus deltoides | LI1B |  | RACiR 800 to 300 | 576.031 | 12.5308 | 25.0089 | 999.926 | 0.006991 | 0.30534 | 630.886 | 658.108
Populus deltoides | LI1B |  | RACiR 800 to 300 | 574.7351 | 12.4332 | 25.0092 | 999.961 | 0.0069937 | 0.3055 | 629.213 | 656.316
Populus deltoides | LI1B |  | RACiR 800 to 300 | 573.4012 | 12.3847 | 25.0098 | 999.961 | 0.0069942 | 0.30555 | 627.681 | 654.716
Populus deltoides | LI1B |  | RACiR 800 to 300 | 571.5043 | 12.4549 | 25.0101 | 999.969 | 0.0069958 | 0.30563 | 625.96 | 653.04
Populus deltoides | LI1B |  | RACiR 800 to 300 | 570.2603 | 12.3793 | 25.0087 | 999.955 | 0.006998 | 0.30584 | 624.405 | 651.387
Populus deltoides | LI1B |  | RACiR 800 to 300 | 568.5593 | 12.3905 | 25.0083 | 1000.27 | 0.0069984 | 0.3059 | 622.691 | 649.66
Populus deltoides | LI1B |  | RACiR 800 to 300 | 566.8938 | 12.4412 | 25.0091 | 999.913 | 0.0069999 | 0.30592 | 621.154 | 648.15
Populus deltoides | LI1B |  | RACiR 800 to 300 | 565.6799 | 12.3167 | 25.0054 | 999.95 | 0.0069984 | 0.30603 | 619.477 | 646.325
Populus deltoides | LI1B |  | RACiR 800 to 300 | 564.2822 | 12.2946 | 25.0032 | 999.941 | 0.0069986 | 0.30615 | 617.953 | 644.757
Populus deltoides | LI1B |  | RACiR 800 to 300 | 562.4929 | 12.3265 | 25.0057 | 999.95 | 0.0070005 | 0.30612 | 616.238 | 643.049
Populus deltoides | LI1B |  | RACiR 800 to 300 | 561.3566 | 12.2112 | 25.0064 | 1000.26 | 0.0069991 | 0.30602 | 614.71 | 641.383
Populus deltoides | LI1B |  | RACiR 800 to 300 | 559.796 | 12.175 | 25.0057 | 999.912 | 0.0069985 | 0.30602 | 612.992 | 639.603
Populus deltoides | LI1B |  | RACiR 800 to 300 | 557.9578 | 12.266 | 25.006 | 999.948 | 0.0069983 | 0.30598 | 611.418 | 638.096
Populus deltoides | LI1B |  | RACiR 800 to 300 | 556.5704 | 12.178 | 25.0054 | 999.953 | 0.0069987 | 0.306 | 609.705 | 636.27
Populus deltoides | LI1B |  | RACiR 800 to 300 | 555.0259 | 12.1839 | 25.0042 | 1000.26 | 0.0069989 | 0.30606 | 608.134 | 634.682
Populus deltoides | LI1B |  | RACiR 800 to 300 | 553.0213 | 12.287 | 25.0051 | 999.91 | 0.0069994 | 0.30604 | 606.429 | 633.056
Populus deltoides | LI1B |  | RACiR 800 to 300 | 551.7575 | 12.2208 | 25.0066 | 999.946 | 0.0069996 | 0.30597 | 604.929 | 631.468
Populus deltoides | LI1B |  | RACiR 800 to 300 | 550.264 | 12.1712 | 25.004 | 999.945 | 0.0069987 | 0.30606 | 603.22 | 629.684
Populus deltoides | LI1B |  | RACiR 800 to 300 | 548.5462 | 12.227 | 25.0051 | 1000.26 | 0.0069977 | 0.30595 | 601.664 | 628.16
Populus deltoides | LI1B |  | RACiR 800 to 300 | 547.1952 | 12.1174 | 25.0074 | 999.956 | 0.0069982 | 0.30584 | 599.94 | 626.3
Populus deltoides | LI1B |  | RACiR 800 to 300 | 545.9379 | 12.0334 | 25.0076 | 999.944 | 0.0069977 | 0.30579 | 598.383 | 624.636
Populus deltoides | LI1B |  | RACiR 800 to 300 | 543.8776 | 12.1378 | 25.0079 | 999.952 | 0.0069978 | 0.30574 | 596.629 | 622.961
Populus deltoides | LI1B |  | RACiR 800 to 300 | 542.6292 | 12.0688 | 25.0052 | 999.957 | 0.006996 | 0.30582 | 595.107 | 621.347
Populus deltoides | LI1B |  | RACiR 800 to 300 | 541.0903 | 12.0369 | 25.0026 | 1000.27 | 0.0069955 | 0.30594 | 593.405 | 619.589
Populus deltoides | LI1B |  | RACiR 800 to 300 | 539.2403 | 12.1385 | 25.0045 | 999.949 | 0.0069948 | 0.30583 | 591.866 | 618.127
Populus deltoides | LI1B |  | RACiR 800 to 300 | 537.9589 | 12.0171 | 25.0103 | 999.948 | 0.0069945 | 0.30553 | 590.205 | 616.32
Populus deltoides | LI1B |  | RACiR 800 to 300 | 536.2682 | 12.0561 | 25.0124 | 999.951 | 0.0069936 | 0.30534 | 588.635 | 614.767
Populus deltoides | LI1B |  | RACiR 800 to 300 | 534.1992 | 12.1727 | 25.0105 | 999.955 | 0.006992 | 0.30532 | 586.904 | 613.127
Populus deltoides | LI1B |  | RACiR 800 to 300 | 532.639 | 12.1906 | 25.0069 | 999.919 | 0.0069929 | 0.30552 | 585.334 | 611.553
Populus deltoides | LI1B |  | RACiR 800 to 300 | 530.977 | 12.1777 | 25.0096 | 999.953 | 0.0069933 | 0.30542 | 583.609 | 609.792
Populus deltoides | LI1B |  | RACiR 800 to 300 | 529.3166 | 12.2185 | 25.0156 | 999.965 | 0.0069943 | 0.30519 | 582.088 | 608.289
Populus deltoides | LI1B |  | RACiR 800 to 300 | 528.0456 | 12.0766 | 25.0213 | 999.958 | 0.0069926 | 0.30481 | 580.379 | 606.409
Populus deltoides | LI1B |  | RACiR 800 to 300 | 526.8207 | 11.9773 | 25.0252 | 999.966 | 0.0069876 | 0.30436 | 578.866 | 604.772
Populus deltoides | LI1B |  | RACiR 800 to 300 | 524.8473 | 12.0696 | 25.0221 | 999.916 | 0.0069876 | 0.3045 | 577.128 | 603.1
Populus deltoides | LI1B |  | RACiR 800 to 300 | 523.8219 | 11.9437 | 25.0173 | 999.925 | 0.0069878 | 0.30474 | 575.621 | 601.445
Populus deltoides | LI1B |  | RACiR 800 to 300 | 521.9003 | 12.0168 | 25.0162 | 999.958 | 0.0069872 | 0.30476 | 573.893 | 599.768
Populus deltoides | LI1B |  | RACiR 800 to 300 | 520.0052 | 12.1504 | 25.0156 | 999.956 | 0.0069908 | 0.30498 | 572.365 | 598.353
Populus deltoides | LI1B |  | RACiR 800 to 300 | 518.6539 | 12.0589 | 25.0151 | 1000.27 | 0.0069907 | 0.305 | 570.676 | 596.547
Populus deltoides | LI1B |  | RACiR 800 to 300 | 517.3972 | 11.985 | 25.0174 | 999.917 | 0.0069907 | 0.30488 | 569.164 | 594.937
Populus deltoides | LI1B |  | RACiR 800 to 300 | 515.3153 | 12.1051 | 25.0153 | 999.954 | 0.0069877 | 0.30478 | 567.445 | 593.312
Populus deltoides | LI1B |  | RACiR 800 to 300 | 514.3169 | 11.9618 | 25.0145 | 999.948 | 0.0069864 | 0.30471 | 565.958 | 591.659
Populus deltoides | LI1B |  | RACiR 800 to 300 | 513.0628 | 11.8357 | 25.0169 | 1000.26 | 0.006986 | 0.30453 | 564.285 | 589.834
Populus deltoides | LI1B |  | RACiR 800 to 300 | 511.1701 | 11.9454 | 25.0167 | 999.921 | 0.0069835 | 0.30438 | 562.734 | 588.371
Populus deltoides | LI1B |  | RACiR 800 to 300 | 510.1303 | 11.7585 | 25.0191 | 999.919 | 0.006986 | 0.30437 | 561.056 | 586.48
Populus deltoides | LI1B |  | RACiR 800 to 300 | 508.9004 | 11.6684 | 25.0208 | 999.959 | 0.0069859 | 0.30427 | 559.515 | 584.827
Populus deltoides | LI1B |  | RACiR 800 to 300 | 506.7644 | 11.8084 | 25.017 | 999.963 | 0.0069839 | 0.30433 | 557.78 | 583.207
Populus deltoides | LI1B |  | RACiR 800 to 300 | 505.6841 | 11.7091 | 25.0156 | 1000.28 | 0.0069999 | 0.3051 | 556.244 | 581.557
Populus deltoides | LI1B |  | RACiR 800 to 300 | 503.884 | 11.7705 | 25.0158 | 999.925 | 0.0070166 | 0.30579 | 554.52 | 579.88
Populus deltoides | LI1B |  | RACiR 800 to 300 | 502.3761 | 11.8096 | 25.014 | 999.967 | 0.0070354 | 0.30664 | 552.996 | 578.381
Populus deltoides | LI1B |  | RACiR 800 to 300 | 501.2066 | 11.6795 | 25.012 | 999.973 | 0.0070497 | 0.30717 | 551.307 | 576.544
Populus deltoides | LI1B |  | RACiR 800 to 300 | 499.8172 | 11.6441 | 25.0101 | 999.981 | 0.007052 | 0.30718 | 549.771 | 574.953
Populus deltoides | LI1B |  | RACiR 800 to 300 | 497.9514 | 11.6944 | 25.0081 | 999.978 | 0.0070456 | 0.30674 | 548.089 | 573.291
Populus deltoides | LI1B |  | RACiR 800 to 300 | 496.4892 | 11.6453 | 25.0081 | 1000.28 | 0.0070362 | 0.30612 | 546.518 | 571.644
Populus deltoides | LI1B |  | RACiR 800 to 300 | 494.6051 | 11.668 | 25.0104 | 999.938 | 0.0070239 | 0.30531 | 544.778 | 569.896
Populus deltoides | LI1B |  | RACiR 800 to 300 | 492.9444 | 11.6974 | 25.0107 | 999.933 | 0.007016 | 0.30483 | 543.242 | 568.359
Populus deltoides | LI1B |  | RACiR 800 to 300 | 491.3674 | 11.641 | 25.0107 | 999.981 | 0.0070045 | 0.30424 | 541.52 | 566.549
Populus deltoides | LI1B |  | RACiR 800 to 300 | 490.0426 | 11.5841 | 25.0071 | 999.971 | 0.0069947 | 0.3039 | 540.014 | 564.959
Populus deltoides | LI1B |  | RACiR 800 to 300 | 488.0727 | 11.6561 | 25.004 | 999.971 | 0.0069809 | 0.30341 | 538.297 | 563.281
Populus deltoides | LI1B |  | RACiR 800 to 300 | 486.5872 | 11.6261 | 25.002 | 999.977 | 0.0069648 | 0.30281 | 536.75 | 561.675
Populus deltoides | LI1B |  | RACiR 800 to 300 | 484.7765 | 11.6522 | 24.9995 | 999.976 | 0.00695 | 0.30232 | 535.041 | 559.963
Populus deltoides | LI1B |  | RACiR 800 to 300 | 482.9105 | 11.7439 | 25.0008 | 1000.3 | 0.006939 | 0.30186 | 533.498 | 558.482
Populus deltoides | LI1B |  | RACiR 800 to 300 | 481.6037 | 11.6519 | 25.0021 | 999.936 | 0.0069397 | 0.30194 | 531.84 | 556.707
Populus deltoides | LI1B |  | RACiR 800 to 300 | 480.4554 | 11.5658 | 25.0028 | 999.936 | 0.0069431 | 0.30213 | 530.35 | 555.11
Populus deltoides | LI1B |  | RACiR 800 to 300 | 478.6588 | 11.6138 | 25.0067 | 999.974 | 0.0069486 | 0.30228 | 528.658 | 553.442
Populus deltoides | LI1B |  | RACiR 800 to 300 | 477.5113 | 11.5121 | 25.01 | 999.987 | 0.0069515 | 0.30229 | 527.145 | 551.807
Populus deltoides | LI1B |  | RACiR 800 to 300 | 475.8658 | 11.5061 | 25.01 | 999.984 | 0.0069515 | 0.3023 | 525.44 | 550.073
Populus deltoides | LI1B |  | RACiR 800 to 300 | 474.1043 | 11.5906 | 25.0093 | 999.972 | 0.0069519 | 0.30237 | 523.909 | 548.602
Populus deltoides | LI1B |  | RACiR 800 to 300 | 472.7869 | 11.4891 | 25.0085 | 999.986 | 0.0069504 | 0.30233 | 522.226 | 546.793
Populus deltoides | LI1B |  | RACiR 800 to 300 | 471.5952 | 11.3999 | 25.0098 | 999.996 | 0.0069489 | 0.30219 | 520.728 | 545.184
Populus deltoides | LI1B |  | RACiR 800 to 300 | 469.7738 | 11.4368 | 25.0093 | 999.989 | 0.0069416 | 0.3019 | 519.024 | 543.487
Populus deltoides | LI1B |  | RACiR 800 to 300 | 468.567 | 11.339 | 25.0091 | 999.996 | 0.0069351 | 0.30164 | 517.493 | 541.834
Populus deltoides | LI1B |  | RACiR 800 to 300 | 467.1895 | 11.2508 | 25.0075 | 999.994 | 0.0069276 | 0.30141 | 515.814 | 540.037
Populus deltoides | LI1B |  | RACiR 800 to 300 | 465.615 | 11.2636 | 25.0051 | 999.981 | 0.0069227 | 0.30134 | 514.249 | 538.459
Populus deltoides | LI1B |  | RACiR 800 to 300 | 464.3442 | 11.1406 | 25.0074 | 1000.29 | 0.006922 | 0.30125 | 512.549 | 536.611
Populus deltoides | LI1B |  | RACiR 800 to 300 | 462.9494 | 11.1052 | 25.0105 | 999.952 | 0.0069235 | 0.3012 | 511.013 | 535.016
Populus deltoides | LI1B |  | RACiR 800 to 300 | 460.9771 | 11.1906 | 25.0135 | 999.961 | 0.0069248 | 0.30118 | 509.287 | 533.351
Populus deltoides | LI1B |  | RACiR 800 to 300 | 459.4342 | 11.1957 | 25.016 | 999.994 | 0.0069256 | 0.30111 | 507.738 | 531.785
Populus deltoides | LI1B |  | RACiR 800 to 300 | 457.9501 | 11.142 | 25.0172 | 999.997 | 0.0069272 | 0.30115 | 506.036 | 530.005
Populus deltoides | LI1B |  | RACiR 800 to 300 | 456.2584 | 11.2066 | 25.0161 | 999.995 | 0.006928 | 0.30124 | 504.508 | 528.52
Populus deltoides | LI1B |  | RACiR 800 to 300 | 454.9895 | 11.0948 | 25.0144 | 1000 | 0.0069267 | 0.30126 | 502.829 | 526.705
Populus deltoides | LI1B |  | RACiR 800 to 300 | 453.7614 | 11.0218 | 25.0121 | 999.988 | 0.0069262 | 0.30135 | 501.311 | 525.091
Populus deltoides | LI1B |  | RACiR 800 to 300 | 451.9391 | 11.0738 | 25.0107 | 999.991 | 0.0069229 | 0.30125 | 499.632 | 523.438
Populus deltoides | LI1B |  | RACiR 800 to 300 | 450.6192 | 11.0034 | 25.0139 | 999.995 | 0.0069213 | 0.30102 | 498.08 | 521.793
Populus deltoides | LI1B |  | RACiR 800 to 300 | 448.9003 | 11.0114 | 25.0153 | 999.994 | 0.0069197 | 0.30081 | 496.378 | 520.072
Populus deltoides | LI1B |  | RACiR 800 to 300 | 447.2245 | 11.0685 | 25.0117 | 999.995 | 0.0069166 | 0.30081 | 494.851 | 518.58
Populus deltoides | LI1B |  | RACiR 800 to 300 | 445.8975 | 10.9892 | 25.0097 | 999.992 | 0.0069178 | 0.30096 | 493.204 | 516.829
Populus deltoides | LI1B |  | RACiR 800 to 300 | 444.6354 | 10.9284 | 25.0069 | 999.997 | 0.0069147 | 0.30091 | 491.712 | 515.252
Populus deltoides | LI1B |  | RACiR 800 to 300 | 442.8376 | 10.9714 | 25.0075 | 999.997 | 0.0069137 | 0.30083 | 490.027 | 513.586
Populus deltoides | LI1B |  | RACiR 800 to 300 | 441.6711 | 10.8664 | 25.0137 | 999.996 | 0.0069155 | 0.3006 | 488.52 | 511.952
Populus deltoides | LI1B |  | RACiR 800 to 300 | 440.2041 | 10.7949 | 25.0166 | 999.999 | 0.0069134 | 0.30033 | 486.817 | 510.153
Populus deltoides | LI1B |  | RACiR 800 to 300 | 438.5154 | 10.8596 | 25.015 | 1000 | 0.0069157 | 0.3005 | 485.284 | 508.661
Populus deltoides | LI1B |  | RACiR 800 to 300 | 437.403 | 10.7104 | 25.0106 | 1000 | 0.006915 | 0.30069 | 483.613 | 506.817
Populus deltoides | LI1B |  | RACiR 800 to 300 | 436.0008 | 10.6817 | 25.0066 | 999.998 | 0.006908 | 0.30054 | 482.094 | 505.245
Populus deltoides | LI1B |  | RACiR 800 to 300 | 434.1547 | 10.7305 | 25.0049 | 999.999 | 0.0068996 | 0.30022 | 480.406 | 503.575
Populus deltoides | LI1B |  | RACiR 800 to 300 | 432.9861 | 10.6326 | 25.0054 | 999.999 | 0.006889 | 0.29974 | 478.938 | 501.982
Populus deltoides | LI1B |  | RACiR 800 to 300 | 431.1312 | 10.6709 | 25.0062 | 1000 | 0.0068766 | 0.29913 | 477.246 | 500.299
Populus deltoides | LI1B |  | RACiR 800 to 300 | 429.6398 | 10.6845 | 25.0053 | 999.995 | 0.0068726 | 0.29906 | 475.771 | 498.809
Populus deltoides | LI1B |  | RACiR 800 to 300 | 428.2633 | 10.6109 | 25.0056 | 999.995 | 0.0068742 | 0.29916 | 474.1 | 497.045
Populus deltoides | LI1B |  | RACiR 800 to 300 | 426.9406 | 10.5751 | 25.0058 | 999.992 | 0.0068767 | 0.29933 | 472.603 | 495.491
Populus deltoides | LI1B |  | RACiR 800 to 300 | 425.1974 | 10.6013 | 25.0076 | 1000 | 0.0068786 | 0.29936 | 470.905 | 493.796
Populus deltoides | LI1B |  | RACiR 800 to 300 | 424.1242 | 10.4761 | 25.0123 | 999.993 | 0.0068803 | 0.29925 | 469.404 | 492.149
Populus deltoides | LI1B |  | RACiR 800 to 300 | 422.9256 | 10.3565 | 25.0088 | 999.997 | 0.0068807 | 0.29945 | 467.744 | 490.343
Populus deltoides | LI1B |  | RACiR 800 to 300 | 421.4138 | 10.3655 | 25.0078 | 1000.01 | 0.0068815 | 0.29952 | 466.218 | 488.805
Populus deltoides | LI1B |  | RACiR 800 to 300 | 420.4622 | 10.1676 | 25.0066 | 1000.01 | 0.006883 | 0.29965 | 464.558 | 486.922
Populus deltoides | LI1B |  | RACiR 800 to 300 | 418.7817 | 10.2193 | 25.0046 | 1000 | 0.0068843 | 0.29978 | 462.996 | 485.391
Populus deltoides | LI1B |  | RACiR 800 to 300 | 416.6902 | 10.3428 | 25.0082 | 1000.01 | 0.0068814 | 0.29946 | 461.317 | 483.81
Populus deltoides | LI1B |  | RACiR 800 to 300 | 415.5272 | 10.236 | 25.0096 | 1000.01 | 0.0068805 | 0.29936 | 459.779 | 482.141
Populus deltoides | LI1B |  | RACiR 800 to 300 | 413.8445 | 10.2351 | 25.0111 | 1000.01 | 0.0068772 | 0.29911 | 458.087 | 480.424
Populus deltoides | LI1B |  | RACiR 800 to 300 | 412.2718 | 10.265 | 25.0109 | 1000.01 | 0.0068745 | 0.29897 | 456.595 | 478.937
Populus deltoides | LI1B |  | RACiR 800 to 300 | 411.4051 | 10.0201 | 25.0133 | 1000.02 | 0.0068741 | 0.29883 | 454.901 | 476.971
Populus deltoides | LI1B |  | RACiR 800 to 300 | 410.3496 | 9.8832 | 25.0155 | 1000.03 | 0.0068712 | 0.29854 | 453.395 | 475.306
Populus deltoides | LI1B |  | RACiR 800 to 300 | 408.4026 | 9.9678 | 25.0133 | 999.993 | 0.00687 | 0.29858 | 451.681 | 473.65
Populus deltoides | LI1B |  | RACiR 800 to 300 | 407.6532 | 9.7509 | 25.0131 | 999.987 | 0.0068701 | 0.29857 | 450.182 | 471.911
Populus deltoides | LI1B |  | RACiR 800 to 300 | 406.1234 | 9.7042 | 25.0122 | 999.986 | 0.0068671 | 0.29845 | 448.47 | 470.128
Populus deltoides | LI1B |  | RACiR 800 to 300 | 404.1626 | 9.8284 | 25.0081 | 1000.02 | 0.0068535 | 0.298 | 446.931 | 468.683
Populus deltoides | LI1B |  | RACiR 800 to 300 | 403.0832 | 9.6461 | 25.0084 | 999.987 | 0.0068388 | 0.29734 | 445.281 | 466.82
Populus deltoides | LI1B |  | RACiR 800 to 300 | 401.6708 | 9.6057 | 25.0106 | 999.983 | 0.0068239 | 0.29664 | 443.777 | 465.248
Populus deltoides | LI1B |  | RACiR 800 to 300 | 399.5422 | 9.7369 | 25.0113 | 999.988 | 0.0068095 | 0.29602 | 442.115 | 463.687
Populus deltoides | LI1B |  | RACiR 800 to 300 | 398.2258 | 9.7017 | 25.0161 | 1000.02 | 0.0068104 | 0.29596 | 440.658 | 462.174
Populus deltoides | LI1B |  | RACiR 800 to 300 | 396.7381 | 9.677 | 25.0162 | 999.99 | 0.0068122 | 0.29616 | 439.025 | 460.495
Populus deltoides | LI1B |  | RACiR 800 to 300 | 395.0744 | 9.7582 | 25.0129 | 1000.04 | 0.0068167 | 0.2966 | 437.543 | 459.072
Populus deltoides | LI1B |  | RACiR 800 to 300 | 394.0443 | 9.5953 | 25.0118 | 1000.04 | 0.006818 | 0.29679 | 435.907 | 457.25
Populus deltoides | LI1B |  | RACiR 800 to 300 | 393.0026 | 9.4907 | 25.0096 | 999.995 | 0.0068195 | 0.29702 | 434.454 | 455.673
Populus deltoides | LI1B |  | RACiR 800 to 300 | 391.5032 | 9.458 | 25.0113 | 1000.01 | 0.0068204 | 0.29701 | 432.812 | 453.973
Populus deltoides | LI1B |  | RACiR 800 to 300 | 390.4526 | 9.3384 | 25.01 | 1000.01 | 0.00682 | 0.29706 | 431.322 | 452.342
Populus deltoides | LI1B |  | RACiR 800 to 300 | 389.0011 | 9.2897 | 25.0075 | 1000 | 0.0068206 | 0.29721 | 429.651 | 450.599
Populus deltoides | LI1B |  | RACiR 800 to 300 | 387.5358 | 9.2905 | 25.0062 | 1000.01 | 0.0068205 | 0.29726 | 428.147 | 449.073
Populus deltoides | LI1B |  | RACiR 800 to 300 | 386.2238 | 9.1976 | 25.001 | 1000.01 | 0.0068199 | 0.29748 | 426.458 | 447.266
Populus deltoides | LI1B |  | RACiR 800 to 300 | 385.0683 | 9.1084 | 25.0018 | 1000 | 0.0068196 | 0.2974 | 424.982 | 445.677
Populus deltoides | LI1B |  | RACiR 800 to 300 | 383.158 | 9.1778 | 25.0052 | 999.999 | 0.0068193 | 0.29719 | 423.292 | 444.032
Populus deltoides | LI1B |  | RACiR 800 to 300 | 382.2336 | 9.0119 | 25.0073 | 1000 | 0.0068175 | 0.29699 | 421.806 | 442.357
Populus deltoides | LI1B |  | RACiR 800 to 300 | 380.7365 | 8.9536 | 25.0148 | 1000 | 0.0068159 | 0.29654 | 420.134 | 440.604
Populus deltoides | LI1B |  | RACiR 800 to 300 | 379.3681 | 8.9197 | 25.0175 | 1000 | 0.0068152 | 0.29637 | 418.64 | 439.053
Populus deltoides | LI1B |  | RACiR 800 to 300 | 378.3719 | 8.7313 | 25.0158 | 1000.02 | 0.0068139 | 0.29638 | 416.975 | 437.175
Populus deltoides | LI1B |  | RACiR 800 to 300 | 377.1926 | 8.6391 | 25.0183 | 1000.01 | 0.0068123 | 0.29618 | 415.477 | 435.563
Populus deltoides | LI1B |  | RACiR 800 to 300 | 375.5136 | 8.65 | 25.0188 | 1000.01 | 0.0068119 | 0.29614 | 413.801 | 433.871
Populus deltoides | LI1B |  | RACiR 800 to 300 | 374.3321 | 8.5585 | 25.0207 | 1000.02 | 0.0068116 | 0.29601 | 412.295 | 432.252
Populus deltoides | LI1B |  | RACiR 800 to 300 | 372.7285 | 8.5404 | 25.022 | 1000.02 | 0.0068117 | 0.29596 | 410.599 | 430.514
Populus deltoides | LI1B |  | RACiR 800 to 300 | 371.033 | 8.6082 | 25.0211 | 1000.01 | 0.0068115 | 0.29599 | 409.092 | 429.052
Populus deltoides | LI1B |  | RACiR 800 to 300 | 369.8908 | 8.4626 | 25.0214 | 1000.01 | 0.0068116 | 0.29596 | 407.43 | 427.218
Populus deltoides | LI1B |  | RACiR 800 to 300 | 368.7399 | 8.3717 | 25.0189 | 1000.02 | 0.0068094 | 0.29596 | 405.938 | 425.612
Populus deltoides | LI1B |  | RACiR 800 to 300 | 366.8824 | 8.4328 | 25.0206 | 1000.01 | 0.0068063 | 0.29569 | 404.277 | 423.987
Populus deltoides | LI1B |  | RACiR 800 to 300 | 365.8651 | 8.2951 | 25.021 | 1000.02 | 0.006807 | 0.29569 | 402.766 | 422.316
Populus deltoides | LI1B |  | RACiR 800 to 300 | 364.6056 | 8.1971 | 25.019 | 1000.02 | 0.006805 | 0.29564 | 401.145 | 420.574
Populus deltoides | LI1B |  | RACiR 800 to 300 | 363.0055 | 8.2378 | 25.0181 | 1000.01 | 0.0068054 | 0.29568 | 399.642 | 419.089
Populus deltoides | LI1B |  | RACiR 800 to 300 | 361.9206 | 8.0831 | 25.0149 | 1000.02 | 0.0068041 | 0.29576 | 397.992 | 417.262
Populus deltoides | LI1B |  | RACiR 800 to 300 | 360.7117 | 8.0119 | 25.0108 | 1000.01 | 0.0068029 | 0.29586 | 396.497 | 415.673
Populus deltoides | LI1B |  | RACiR 800 to 300 | 359.0153 | 8.0268 | 25.0107 | 1000.01 | 0.0068051 | 0.29597 | 394.801 | 413.967
Populus deltoides | LI1B |  | RACiR 800 to 300 | 357.8314 | 7.9414 | 25.014 | 1000.01 | 0.0068023 | 0.29565 | 393.334 | 412.392
Populus deltoides | LI1B |  | RACiR 800 to 300 | 356.1996 | 7.9416 | 25.0151 | 1000.02 | 0.0067998 | 0.29544 | 391.686 | 410.717
Populus deltoides | LI1B |  | RACiR 800 to 300 | 354.5213 | 8.0039 | 25.019 | 1000.01 | 0.0067981 | 0.29513 | 390.219 | 409.29
Populus deltoides | LI1B |  | RACiR 800 to 300 | 353.2909 | 7.8932 | 25.0201 | 1000.02 | 0.0067921 | 0.29475 | 388.618 | 407.555
Populus deltoides | LI1B |  | RACiR 800 to 300 | 352.1029 | 7.8098 | 25.02 | 1000.01 | 0.0067914 | 0.29467 | 387.125 | 405.956
Populus deltoides | LI1B |  | RACiR 800 to 300 | 350.7079 | 7.751 | 25.0208 | 1000.02 | 0.0067908 | 0.29457 | 385.507 | 404.255
Populus deltoides | LI1B |  | RACiR 800 to 300 | 349.7794 | 7.5992 | 25.0165 | 1000.02 | 0.0067892 | 0.29465 | 384.024 | 402.598
Populus deltoides | LI1B |  | RACiR 800 to 300 | 348.5143 | 7.5001 | 25.0132 | 1000.01 | 0.0067884 | 0.29475 | 382.376 | 400.826
Populus deltoides | LI1B |  | RACiR 800 to 300 | 347.1261 | 7.4844 | 25.0099 | 1000.02 | 0.0067883 | 0.29491 | 380.883 | 399.295
Populus deltoides | LI1B |  | RACiR 800 to 300 | 345.7974 | 7.3887 | 25.0079 | 1000.02 | 0.0067875 | 0.29494 | 379.191 | 397.481
Populus deltoides | LI1B |  | RACiR 800 to 300 | 344.543 | 7.3267 | 25.0097 | 1000.02 | 0.0067819 | 0.29457 | 377.73 | 395.935
Populus deltoides | LI1B |  | RACiR 800 to 300 | 342.8186 | 7.3579 | 25.0094 | 999.702 | 0.0067846 | 0.2947 | 376.061 | 394.272
Populus deltoides | LI1B |  | RACiR 800 to 300 | 341.6387 | 7.2635 | 25.0112 | 1000.02 | 0.0067748 | 0.29411 | 374.584 | 392.675
Populus deltoides | LI1B |  | RACiR 800 to 300 | 340.5527 | 7.1066 | 25.0135 | 1000.02 | 0.0067635 | 0.2935 | 372.988 | 390.896
Populus deltoides | LI1B |  | RACiR 800 to 300 | 339.1192 | 7.0831 | 25.0143 | 1000.01 | 0.0067548 | 0.29312 | 371.471 | 389.33
Populus deltoides | LI1B |  | RACiR 800 to 300 | 338.2303 | 6.8532 | 25.0181 | 1000.01 | 0.0067429 | 0.29241 | 369.831 | 387.431
Populus deltoides | LI1B |  | RACiR 800 to 300 | 337.2674 | 6.7149 | 25.0184 | 1000.01 | 0.0067372 | 0.29222 | 368.381 | 385.819
Populus deltoides | LI1B |  | RACiR 800 to 300 | 335.6043 | 6.7251 | 25.0149 | 1000.02 | 0.0067374 | 0.29247 | 366.688 | 384.112
Populus deltoides | LI1B |  | RACiR 800 to 300 | 334.4192 | 6.6485 | 25.0155 | 1000.03 | 0.0067391 | 0.29254 | 365.205 | 382.531
Populus deltoides | LI1B |  | RACiR 800 to 300 | 332.9003 | 6.6195 | 25.015 | 1000.02 | 0.0067398 | 0.2926 | 363.544 | 380.819
Populus deltoides | LI1B |  | RACiR 800 to 300 | 331.338 | 6.6559 | 25.0144 | 1000.02 | 0.0067417 | 0.29273 | 362.06 | 379.348
Populus deltoides | LI1B |  | RACiR 800 to 300 | 330.2887 | 6.4944 | 25.0151 | 1000.03 | 0.0067441 | 0.29281 | 360.423 | 377.526
Populus deltoides | LI1B |  | RACiR 800 to 300 | 329.2784 | 6.3801 | 25.0117 | 1000.02 | 0.0067423 | 0.29287 | 358.985 | 375.952
Populus deltoides | LI1B |  | RACiR 800 to 300 | 327.8953 | 6.313 | 25.0086 | 1000.02 | 0.0067419 | 0.29303 | 357.322 | 374.196
Populus deltoides | LI1B |  | RACiR 800 to 300 | 326.9521 | 6.1515 | 25.0077 | 1000.03 | 0.0067337 | 0.29269 | 355.823 | 372.51
Populus deltoides | LI1B |  | RACiR 800 to 300 | 325.5877 | 6.064 | 25.0078 | 1000.02 | 0.0067136 | 0.29176 | 354.193 | 370.763
Populus deltoides | LI1B |  | RACiR 800 to 300 | 323.8832 | 6.1147 | 25.0106 | 1000.03 | 0.0066992 | 0.29103 | 352.681 | 369.274
Populus deltoides | LI1B |  | RACiR 800 to 300 | 322.6115 | 6.0038 | 25.0142 | 1000.02 | 0.0066825 | 0.29015 | 351.061 | 367.514
Populus deltoides | LI1B |  | RACiR 800 to 300 | 321.233 | 5.9812 | 25.0157 | 1000.03 | 0.0066768 | 0.28993 | 349.587 | 365.994
Populus deltoides | LI1B |  | RACiR 800 to 300 | 319.6469 | 5.9849 | 25.0147 | 1000.03 | 0.0066782 | 0.29016 | 347.956 | 364.344
Populus deltoides | LI1B |  | RACiR 800 to 300 | 318.707 | 5.8493 | 25.0132 | 1000.02 | 0.0066838 | 0.29058 | 346.487 | 362.719
Populus deltoides | LI1B |  | RACiR 800 to 300 | 317.2906 | 5.8112 | 25.0115 | 1000.02 | 0.0066839 | 0.29076 | 344.888 | 361.06
Populus deltoides | LI1B |  | RACiR 800 to 300 | 315.9938 | 5.7794 | 25.0113 | 1000.02 | 0.006685 | 0.29086 | 343.442 | 359.561
Populus deltoides | LI1B |  | RACiR 800 to 300 | 315.079 | 5.5807 | 25.0192 | 1000.02 | 0.0066854 | 0.29056 | 341.843 | 357.739
Populus deltoides | LI1B |  | RACiR 800 to 300 | 313.9939 | 5.4795 | 25.0209 | 999.713 | 0.0066836 | 0.29038 | 340.395 | 356.168
Populus deltoides | LI1B |  | RACiR 800 to 300 | 312.5775 | 5.4351 | 25.0177 | 1000.03 | 0.0066865 | 0.29065 | 338.769 | 354.474
Populus deltoides | LI1B |  | RACiR 800 to 300 | 311.5394 | 5.3155 | 25.0175 | 1000.02 | 0.0066859 | 0.29063 | 337.292 | 352.857
Populus deltoides | LI1B |  | RACiR 800 to 300 | 310.1723 | 5.2547 | 25.0144 | 1000.02 | 0.0066871 | 0.29084 | 335.664 | 351.144
Populus deltoides | LI1B |  | RACiR 800 to 300 | 308.7681 | 5.2462 | 25.0149 | 1000.01 | 0.0066874 | 0.29085 | 334.197 | 349.647
Populus deltoides | LI1B |  | RACiR 800 to 300 | 307.8588 | 5.0412 | 25.0199 | 1000.02 | 0.0066868 | 0.29059 | 332.575 | 347.796
Populus deltoides | LI1B |  | RACiR 800 to 300 | 306.852 | 4.923 | 25.0195 | 1000.02 | 0.0066856 | 0.29055 | 331.136 | 346.215
Populus deltoides | LI1B |  | RACiR 800 to 300 | 305.1242 | 4.9594 | 25.018 | 1000.01 | 0.0066854 | 0.29058 | 329.492 | 344.584
Populus deltoides | LI1B |  | RACiR 800 to 300 | 304.1414 | 4.8294 | 25.019 | 999.999 | 0.0066833 | 0.29039 | 328.047 | 342.987
Populus deltoides | LI1B |  | RACiR 800 to 300 | 303.006 | 4.6934 | 25.0217 | 1000.01 | 0.0066815 | 0.29015 | 326.429 | 341.21
Populus deltoides | LI1B |  | RACiR 800 to 300 | 301.712 | 4.6485 | 25.0242 | 1000.01 | 0.0066801 | 0.28995 | 324.963 | 339.677
Populus deltoides | LI1B |  | RACiR 800 to 300 | 300.9418 | 4.41 | 25.024 | 1000.01 | 0.006678 | 0.28984 | 323.349 | 337.799
Populus deltoides | LI1B |  | RACiR 800 to 300 | 299.9326 | 4.2784 | 25.0219 | 1000 | 0.0066749 | 0.28977 | 321.859 | 336.154
Populus deltoides | LI1B |  | RACiR 800 to 300 | 298.4229 | 4.2502 | 25.0179 | 999.999 | 0.006675 | 0.28996 | 320.201 | 334.442
Populus deltoides | LI1B |  | RACiR 800 to 300 | 297.2397 | 4.1756 | 25.0155 | 1000.01 | 0.0066739 | 0.29 | 318.726 | 332.871
Populus deltoides | LI1B |  | RACiR 800 to 300 | 295.9514 | 4.0908 | 25.0143 | 1000.01 | 0.006673 | 0.29002 | 317.11 | 331.147
Populus deltoides | LI1B |  | RACiR 800 to 300 | 294.2817 | 4.1599 | 25.0136 | 1000 | 0.0066731 | 0.29002 | 315.642 | 329.727
Populus deltoides | LI1B |  | RACiR 800 to 300 | 293.3024 | 3.9861 | 25.0141 | 1000 | 0.0066718 | 0.28989 | 314.042 | 327.929
Populus deltoides | LI1B |  | RACiR 800 to 300 | 292.4745 | 3.8146 | 25.0138 | 999.997 | 0.0066698 | 0.28976 | 312.603 | 326.296
Populus deltoides | LI1B |  | RACiR 800 to 300 | 290.8681 | 3.8102 | 25.016 | 999.993 | 0.0066673 | 0.2895 | 310.96 | 324.624
Populus deltoides | LI1B |  | RACiR 800 to 300 | 290.1147 | 3.6129 | 25.0198 | 999.997 | 0.0066664 | 0.28924 | 309.516 | 322.961
Populus deltoides | LI1B |  | RACiR 800 to 300 | 288.8682 | 3.5121 | 25.0199 | 1000 | 0.0066664 | 0.28925 | 307.887 | 321.208
Populus deltoides | LI1B |  | RACiR 800 to 300 | 287.3951 | 3.5147 | 25.0228 | 999.992 | 0.0066648 | 0.28902 | 306.402 | 319.703
Populus deltoides | LI1B |  | RACiR 800 to 300 | 286.1983 | 3.3961 | 25.0222 | 1000 | 0.0066642 | 0.28899 | 304.764 | 317.922
Populus deltoides | LI1B |  | RACiR 800 to 300 | 284.9329 | 3.3467 | 25.0184 | 1000.01 | 0.0066627 | 0.28906 | 303.29 | 316.377
Populus deltoides | LI1B |  | RACiR 800 to 300 | 283.3087 | 3.3629 | 25.0132 | 999.996 | 0.0066606 | 0.28916 | 301.676 | 314.755
Populus deltoides | LI1B |  | RACiR 800 to 300 | 282.2098 | 3.2798 | 25.008 | 999.989 | 0.0066578 | 0.28926 | 300.252 | 313.227
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 345.5895 | 29.3187 | 25.2166 | 999.887 | 0.0046244 | 0.28306 | 455.713 | 478.23
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 344.3892 | 29.0026 | 25.2169 | 999.873 | 0.0046401 | 0.28448 | 452.842 | 475.021
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 342.9853 | 28.7277 | 25.2175 | 999.873 | 0.0046559 | 0.28586 | 449.94 | 471.812
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 341.6437 | 28.4235 | 25.2116 | 999.867 | 0.0046675 | 0.28739 | 446.949 | 468.495
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 339.6198 | 28.2498 | 25.2114 | 999.876 | 0.0046761 | 0.28833 | 443.952 | 465.293
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 337.5607 | 28.0665 | 25.2137 | 999.864 | 0.0046853 | 0.28909 | 440.954 | 462.078
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 335.5609 | 27.8339 | 25.2121 | 999.857 | 0.0046898 | 0.28974 | 437.885 | 458.75
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 333.3029 | 27.6734 | 25.2093 | 1000.17 | 0.0046936 | 0.29033 | 434.83 | 455.502
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 330.7449 | 27.5304 | 25.2131 | 1000.17 | 0.0046941 | 0.29021 | 431.777 | 452.271
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 328.2377 | 27.32 | 25.2214 | 1000.17 | 0.004694 | 0.28972 | 428.668 | 448.919
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 325.5054 | 27.2423 | 25.222 | 1000.16 | 0.004696 | 0.28985 | 425.576 | 445.719
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 323.1661 | 27.0556 | 25.2217 | 999.847 | 0.0046957 | 0.28988 | 422.541 | 442.46
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 321.1127 | 26.7497 | 25.2194 | 999.841 | 0.0046901 | 0.28964 | 419.465 | 439.042
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 318.5566 | 26.5686 | 25.2184 | 1000.15 | 0.0046847 | 0.28932 | 416.339 | 435.699
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 315.9762 | 26.4015 | 25.2172 | 1000.15 | 0.0046791 | 0.28898 | 413.242 | 432.397
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 313.8187 | 26.1347 | 25.2106 | 1000.13 | 0.004673 | 0.28896 | 410.117 | 428.971
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 311.2976 | 25.9107 | 25.213 | 1000.14 | 0.0046674 | 0.28838 | 406.961 | 425.553
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 308.517 | 25.7837 | 25.2133 | 999.821 | 0.0046637 | 0.28803 | 403.801 | 422.228
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 306.1138 | 25.5828 | 25.2054 | 999.82 | 0.0046561 | 0.28798 | 400.664 | 418.856
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 303.1354 | 25.5135 | 25.2028 | 999.797 | 0.0046476 | 0.28745 | 397.557 | 415.642
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 300.1841 | 25.4404 | 25.2015 | 1000.12 | 0.0046404 | 0.28696 | 394.456 | 412.429
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 297.9711 | 25.1255 | 25.2036 | 1000.12 | 0.0046294 | 0.28595 | 391.419 | 409.043
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 295.4377 | 24.9713 | 25.194 | 1000.11 | 0.0046214 | 0.28591 | 388.302 | 405.732
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 293.0652 | 24.8506 | 25.1749 | 1000.1 | 0.0046166 | 0.28675 | 385.189 | 402.46
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 290.7522 | 24.6159 | 25.1695 | 1000.1 | 0.0046078 | 0.2864 | 382.114 | 399.117
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 288.0021 | 24.4943 | 25.1673 | 1000.1 | 0.0045999 | 0.28592 | 379.037 | 395.884
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 285.3613 | 24.3916 | 25.1566 | 999.776 | 0.0045917 | 0.28592 | 375.981 | 392.685
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 283.2364 | 24.1341 | 25.1427 | 1000.09 | 0.0045791 | 0.28583 | 372.927 | 389.339
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 281.2405 | 23.8489 | 25.1289 | 1000.08 | 0.0045692 | 0.28593 | 369.851 | 385.939
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 278.9015 | 23.6414 | 25.1187 | 1000.08 | 0.0045599 | 0.28584 | 366.759 | 382.602
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 276.6375 | 23.3813 | 25.1114 | 1000.08 | 0.0045508 | 0.28555 | 363.622 | 379.166
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 274.002 | 23.2002 | 25.1089 | 1000.06 | 0.0045414 | 0.2849 | 360.493 | 375.82
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 271.1145 | 23.1236 | 25.101 | 999.767 | 0.004532 | 0.28462 | 357.363 | 372.575
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 268.7726 | 22.9097 | 25.0906 | 1000.06 | 0.0045244 | 0.28466 | 354.205 | 369.167
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 266.1165 | 22.7782 | 25.0833 | 1000.1 | 0.0045156 | 0.28442 | 351.101 | 365.896
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 263.4944 | 22.6666 | 25.0713 | 999.754 | 0.0045054 | 0.28442 | 348.031 | 362.676
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 261.5152 | 22.3701 | 25.0599 | 1000.05 | 0.0044956 | 0.28442 | 344.959 | 359.272
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 259.0358 | 22.2166 | 25.0485 | 1000.05 | 0.004487 | 0.28446 | 341.877 | 356.003
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 256.515 | 22.08 | 25.0398 | 1000.1 | 0.0044809 | 0.28446 | 338.824 | 352.775
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 254.41 | 21.8189 | 25.031 | 1000.06 | 0.0044733 | 0.28443 | 335.763 | 349.418
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 251.9547 | 21.684 | 25.0177 | 1000.08 | 0.0044656 | 0.28473 | 332.697 | 346.184
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 249.7002 | 21.5106 | 25.0049 | 1000.04 | 0.0044593 | 0.28511 | 329.679 | 342.954
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 247.6989 | 21.2424 | 24.9951 | 999.745 | 0.0044527 | 0.28525 | 326.652 | 339.625
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 245.3775 | 21.0286 | 24.9911 | 1000.05 | 0.0044465 | 0.28503 | 323.592 | 336.318
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 243.0276 | 20.827 | 24.9894 | 1000.01 | 0.0044436 | 0.2849 | 320.525 | 333.011
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 241.555 | 20.4601 | 24.9855 | 1000.02 | 0.0044506 | 0.28593 | 317.463 | 329.553
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 240.5116 | 20.1893 | 24.9831 | 1000.04 | 0.004495 | 0.29036 | 314.337 | 326.138
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 239.0415 | 20.1612 | 24.9865 | 1000.04 | 0.0045647 | 0.29686 | 311.186 | 322.946
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 237.6195 | 20.0123 | 24.9976 | 1000.03 | 0.0046261 | 0.30214 | 308.031 | 319.628
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 235.903 | 19.7901 | 25.0099 | 1000.02 | 0.0046616 | 0.30481 | 304.965 | 316.318
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 233.5334 | 19.6267 | 25.0195 | 1000.02 | 0.0046729 | 0.30525 | 301.928 | 313.086
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 231.1927 | 19.3286 | 25.0339 | 1000.03 | 0.0046689 | 0.3039 | 298.866 | 309.69
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 228.4102 | 19.0984 | 25.0502 | 1000.01 | 0.0046546 | 0.30144 | 295.814 | 306.371
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 225.3344 | 18.9733 | 25.062 | 1000.02 | 0.0046416 | 0.29938 | 292.726 | 303.12
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 223.2301 | 18.6043 | 25.0648 | 999.705 | 0.0046314 | 0.29821 | 289.599 | 299.586
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 220.3282 | 18.4631 | 25.0695 | 1000.01 | 0.0046183 | 0.29671 | 286.499 | 296.308
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 216.9832 | 18.411 | 25.0853 | 1000.01 | 0.0046046 | 0.29433 | 283.449 | 293.168
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 214.6098 | 18.1206 | 25.0957 | 1000.01 | 0.0045926 | 0.29256 | 280.433 | 289.825
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 212.098 | 17.9006 | 25.1011 | 1000.01 | 0.0045819 | 0.29127 | 277.406 | 286.544
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 209.7358 | 17.6368 | 25.1093 | 999.993 | 0.0045729 | 0.28991 | 274.39 | 283.227
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 207.8344 | 17.2399 | 25.1154 | 999.995 | 0.0045617 | 0.28855 | 271.368 | 279.771
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 205.2448 | 17.0466 | 25.1207 | 999.987 | 0.0045512 | 0.28734 | 268.32 | 276.494
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 202.36 | 16.9365 | 25.1294 | 999.986 | 0.0045439 | 0.28612 | 265.266 | 273.294
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 200.0832 | 16.6674 | 25.1344 | 999.98 | 0.0045362 | 0.2851 | 262.223 | 269.948
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 197.6408 | 16.4292 | 25.1412 | 999.983 | 0.0045274 | 0.28389 | 259.155 | 266.606
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 195.0146 | 16.2754 | 25.1427 | 999.989 | 0.0045196 | 0.28314 | 256.096 | 263.357
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 192.7082 | 16.0612 | 25.1364 | 999.986 | 0.0045112 | 0.2829 | 253.036 | 260.048
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 190.2943 | 15.8362 | 25.1417 | 999.982 | 0.004504 | 0.28193 | 249.982 | 256.734
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 187.8348 | 15.6389 | 25.1466 | 999.986 | 0.0044986 | 0.28115 | 246.938 | 253.458
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 185.6956 | 15.3585 | 25.1496 | 1000.01 | 0.0044931 | 0.28053 | 243.884 | 250.089
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 183.0797 | 15.2285 | 25.1506 | 999.976 | 0.0044882 | 0.28007 | 240.85 | 246.892
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 180.5069 | 15.0824 | 25.1504 | 1000.01 | 0.0044803 | 0.27941 | 237.84 | 243.701
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 178.5361 | 14.7676 | 25.1512 | 999.969 | 0.0044723 | 0.27874 | 234.833 | 240.344
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 176.4636 | 14.4966 | 25.1468 | 1000.02 | 0.0044648 | 0.27843 | 231.802 | 237.008
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 174.2412 | 14.2543 | 25.1449 | 1000.02 | 0.0044576 | 0.2779 | 228.767 | 233.695
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 172.5052 | 13.8632 | 25.1416 | 1000.01 | 0.004445 | 0.27713 | 225.723 | 230.224
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 170.3638 | 13.6013 | 25.1352 | 1000.01 | 0.0044371 | 0.27689 | 222.634 | 226.838
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 167.5316 | 13.5372 | 25.1272 | 1000.01 | 0.0044308 | 0.27687 | 219.523 | 223.628
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 165.1887 | 13.3054 | 25.13 | 1000 | 0.0044223 | 0.27599 | 216.456 | 220.295
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 162.8424 | 13.0804 | 25.1365 | 999.999 | 0.0044163 | 0.27505 | 213.42 | 217
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 160.6819 | 12.8659 | 25.1261 | 1000 | 0.0044101 | 0.27521 | 210.405 | 213.736
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 158.839 | 12.5553 | 25.1154 | 1000 | 0.0044026 | 0.27525 | 207.382 | 210.368
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 156.3619 | 12.3781 | 25.1125 | 1000.01 | 0.0043922 | 0.27461 | 204.321 | 207.095
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 153.955 | 12.1778 | 25.113 | 999.97 | 0.0043833 | 0.27385 | 201.266 | 203.805
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 152.085 | 11.83 | 25.1122 | 1000.01 | 0.0043745 | 0.27316 | 198.197 | 200.353
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 149.4949 | 11.6913 | 25.1067 | 999.975 | 0.0043673 | 0.27288 | 195.096 | 197.079
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 146.8883 | 11.5908 | 25.0954 | 999.974 | 0.0043599 | 0.273 | 192.052 | 193.901
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 145.2337 | 11.2479 | 25.0812 | 1000.02 | 0.0043538 | 0.2734 | 189.036 | 190.508
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 143.0006 | 11.0292 | 25.0785 | 999.981 | 0.0043493 | 0.27314 | 186 | 187.22
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 140.4903 | 10.8642 | 25.0797 | 1000.02 | 0.0043401 | 0.27229 | 182.968 | 183.989
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 138.7931 | 10.4838 | 25.0746 | 1000.02 | 0.0043274 | 0.27161 | 179.931 | 180.536
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 136.8935 | 10.1547 | 25.0729 | 999.975 | 0.0043202 | 0.27113 | 176.847 | 177.088
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 134.1718 | 10.063 | 25.0673 | 1000.01 | 0.0043133 | 0.27092 | 173.767 | 173.883
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 131.963 | 9.851 | 25.056 | 1000 | 0.0043045 | 0.27094 | 170.726 | 170.596
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 129.67 | 9.6534 | 25.0493 | 1000.02 | 0.0042984 | 0.27078 | 167.68 | 167.319
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 127.26 | 9.4874 | 25.0444 | 999.981 | 0.004292 | 0.27051 | 164.648 | 164.088
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 125.5858 | 9.1225 | 25.0414 | 1000 | 0.0042839 | 0.27005 | 161.645 | 160.686
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 123.54 | 8.865 | 25.0368 | 1000 | 0.0042771 | 0.26977 | 158.639 | 157.389
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 121.139 | 8.7106 | 25.0279 | 1000 | 0.0042737 | 0.27005 | 155.584 | 154.147
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 119.2045 | 8.4281 | 25.0174 | 1000 | 0.0042679 | 0.27024 | 152.538 | 150.785
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 116.9229 | 8.2335 | 25.0109 | 999.963 | 0.0042602 | 0.27004 | 149.514 | 147.533
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 114.4639 | 8.0886 | 25.0069 | 1000.01 | 0.0042567 | 0.26997 | 146.482 | 144.324
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 112.7988 | 7.7176 | 25.0072 | 1000.01 | 0.0042489 | 0.26932 | 143.476 | 140.913
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 110.6949 | 7.4666 | 25.0044 | 999.965 | 0.0042432 | 0.26901 | 140.431 | 137.584
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 108.1828 | 7.3361 | 24.9939 | 1000 | 0.0042396 | 0.26935 | 137.352 | 134.342
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 106.618 | 6.9418 | 24.9909 | 999.994 | 0.0042337 | 0.26906 | 134.314 | 130.875
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 104.6066 | 6.6579 | 24.9891 | 999.973 | 0.0042302 | 0.26885 | 131.226 | 127.47
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 102.09 | 6.5191 | 24.9851 | 999.99 | 0.0042291 | 0.26904 | 128.131 | 124.204
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 100.129 | 6.2598 | 24.9823 | 999.994 | 0.0042382 | 0.27015 | 125.067 | 120.849
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 98.0217 | 6.0723 | 24.979 | 1000.01 | 0.0042636 | 0.27281 | 121.999 | 117.564
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 95.8481 | 5.9136 | 24.9825 | 999.997 | 0.0042984 | 0.27581 | 118.963 | 114.342
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 93.9685 | 5.6301 | 24.996 | 999.995 | 0.004322 | 0.27719 | 115.916 | 110.981
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 92.0603 | 5.3321 | 25.0013 | 999.988 | 0.0043322 | 0.27783 | 112.85 | 107.585
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 89.4812 | 5.2106 | 25.0044 | 999.989 | 0.0043343 | 0.27781 | 109.792 | 104.374
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 87.2064 | 4.991 | 25.0148 | 999.978 | 0.0043312 | 0.27685 | 106.751 | 101.08
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 84.8735 | 4.7965 | 25.0185 | 999.974 | 0.0043246 | 0.27603 | 103.728 | 97.8297
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 82.5014 | 4.615 | 25.0212 | 999.987 | 0.0043138 | 0.27491 | 100.726 | 94.6134
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 80.6763 | 4.2715 | 25.0319 | 999.984 | 0.0043026 | 0.27326 | 97.7128 | 91.2227
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 78.5069 | 4.0284 | 25.036 | 999.985 | 0.0042943 | 0.27226 | 94.6688 | 87.9022
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 75.999 | 3.8755 | 25.0361 | 999.986 | 0.004284 | 0.27135 | 91.6043 | 84.6522
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 74.1999 | 3.5243 | 25.0471 | 999.991 | 0.0042768 | 0.26996 | 88.5436 | 81.2059
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 72.4782 | 3.1605 | 25.0452 | 999.986 | 0.0042677 | 0.26931 | 85.4653 | 77.7289
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 69.8801 | 3.0226 | 25.0448 | 999.995 | 0.004257 | 0.26843 | 82.3474 | 74.4403
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 67.5653 | 2.8085 | 25.0548 | 999.992 | 0.0042475 | 0.26699 | 79.2522 | 71.0973
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 65.3336 | 2.588 | 25.0609 | 999.985 | 0.0042408 | 0.26607 | 76.1884 | 67.7795
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 62.7147 | 2.4736 | 25.0609 | 999.985 | 0.0042345 | 0.26551 | 73.1143 | 64.5587
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 60.8004 | 2.1766 | 25.0588 | 999.994 | 0.0042276 | 0.26504 | 70.0564 | 61.1713
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 58.8172 | 1.8976 | 25.0616 | 999.987 | 0.0042206 | 0.2643 | 67.0017 | 57.8037
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 56.4032 | 1.7388 | 25.0599 | 999.985 | 0.0042145 | 0.26394 | 63.952 | 54.5638
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 54.3417 | 1.4828 | 25.0573 | 999.99 | 0.0042083 | 0.26359 | 60.8869 | 51.2088
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 51.1481 | 1.3084 | 25.0641 | 1000.29 | 0.0042039 | 0.26275 | 56.9926 | 47.0996
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 49.1892 | 1.0369 | 25.0677 | 999.975 | 0.0041988 | 0.2621 | 53.9716 | 43.775
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 47.1191 | 0.78394 | 25.0683 | 999.978 | 0.0041909 | 0.26146 | 50.903 | 40.4202
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 45.2078 | 0.48955 | 25.0683 | 999.976 | 0.004187 | 0.26112 | 47.8285 | 37.0176
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 42.9341 | 0.29077 | 25.0677 | 999.978 | 0.0041815 | 0.2607 | 44.7527 | 33.7114
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 40.3482 | 0.17482 | 25.0751 | 1000.29 | 0.0041781 | 0.25996 | 41.6799 | 30.4908
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 38.6333 | -0.16144 | 25.0754 | 999.97 | 0.0041732 | 0.25956 | 38.6325 | 27.0738
Populus deltoides | LI1B |  | RACiR 500 to 0 no VPD control | 36.715 | -0.45078 | 25.0721 | 999.958 | 0.0041677 | 0.25934 | 35.5575 | 23.677
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 295.234 | 22.2646 | 24.9398 | 999.969 | 0.0035135 | 0.20984 | 407.217 | 413.624
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 291.4843 | 22.157 | 24.9323 | 1000.28 | 0.0035556 | 0.21352 | 401.02 | 407.266
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 289.6292 | 21.5107 | 24.9355 | 999.969 | 0.0035895 | 0.2161 | 394.857 | 400.396
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 286.6977 | 20.9822 | 24.9388 | 999.97 | 0.0036137 | 0.21783 | 388.623 | 393.569
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 281.234 | 20.9387 | 24.9364 | 999.959 | 0.003627 | 0.21896 | 382.357 | 387.194
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 277.2542 | 20.4654 | 24.9416 | 999.961 | 0.0036304 | 0.21895 | 376.14 | 380.432
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 272.0284 | 20.2316 | 24.9484 | 999.961 | 0.0036318 | 0.21864 | 369.896 | 373.882
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 265.4901 | 20.2728 | 24.9547 | 999.944 | 0.00363 | 0.21806 | 363.695 | 367.651
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 261.5916 | 19.7118 | 24.9624 | 1000.26 | 0.0036223 | 0.21703 | 357.582 | 360.902
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 257.2803 | 19.2713 | 24.9647 | 999.936 | 0.0036174 | 0.21645 | 351.401 | 354.206
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 251.657 | 19.1121 | 24.9671 | 999.936 | 0.0036141 | 0.21597 | 345.146 | 347.718
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 248.3436 | 18.4604 | 24.9655 | 999.931 | 0.0036079 | 0.21549 | 338.928 | 340.772
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 243.9941 | 18.0101 | 24.9703 | 999.924 | 0.0036034 | 0.21483 | 332.663 | 333.981
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 237.6105 | 17.9768 | 24.9784 | 999.931 | 0.0035988 | 0.21401 | 326.354 | 327.566
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 232.8532 | 17.6585 | 24.976 | 999.925 | 0.0035949 | 0.21373 | 320.124 | 320.944
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 227.4019 | 17.4782 | 24.9768 | 1000.22 | 0.0035891 | 0.2132 | 313.937 | 314.504
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 221.5809 | 17.3727 | 24.9822 | 999.909 | 0.0035844 | 0.21255 | 307.78 | 308.17
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 218.2258 | 16.7944 | 24.9834 | 999.906 | 0.003586 | 0.21249 | 301.647 | 301.386
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 214.4434 | 16.3196 | 24.9769 | 999.902 | 0.0035849 | 0.21264 | 295.49 | 294.681
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 208.5727 | 16.2488 | 24.9793 | 1000.21 | 0.0035826 | 0.21227 | 289.323 | 288.373
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 204.4171 | 15.8076 | 24.9879 | 999.904 | 0.0035811 | 0.21171 | 283.208 | 281.744
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 200.1105 | 15.4243 | 24.9961 | 999.907 | 0.0035826 | 0.21135 | 277.136 | 275.218
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 195.2269 | 15.179 | 25.001 | 999.909 | 0.003584 | 0.21115 | 271.071 | 268.839
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 191.7408 | 14.6493 | 24.9985 | 1000.22 | 0.003583 | 0.21113 | 265.003 | 262.168
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 187.7334 | 14.221 | 24.9981 | 1000.21 | 0.0035839 | 0.21111 | 258.889 | 255.554
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 182.4595 | 14.0316 | 25.0047 | 999.906 | 0.0035838 | 0.21073 | 252.747 | 249.153
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 178.3215 | 13.6315 | 25.0008 | 1000.21 | 0.0035812 | 0.21067 | 246.639 | 242.572
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 174.4369 | 13.177 | 24.9983 | 1000.21 | 0.0035785 | 0.21049 | 240.565 | 235.971
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 170.1413 | 12.7896 | 25.0001 | 1000.21 | 0.0035749 | 0.21006 | 234.469 | 229.416
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 166.6142 | 12.2463 | 24.9975 | 999.903 | 0.0035726 | 0.20994 | 228.307 | 222.636
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 162.1291 | 11.9068 | 24.9971 | 999.899 | 0.0035746 | 0.21 | 222.098 | 216.017
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 156.5725 | 11.7852 | 25.0005 | 999.891 | 0.0035733 | 0.20967 | 215.955 | 209.682
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 152.822 | 11.2871 | 25.0058 | 1000.2 | 0.0035727 | 0.20929 | 209.849 | 203.005
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 148.8192 | 10.835 | 25.0112 | 1000.21 | 0.0035705 | 0.20884 | 203.716 | 196.347
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 143.8701 | 10.6087 | 25.011 | 1000.21 | 0.0035728 | 0.20897 | 197.556 | 189.892
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 140.3512 | 10.0768 | 25.0147 | 999.879 | 0.0035745 | 0.20889 | 191.429 | 183.159
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 135.5146 | 9.8301 | 25.0103 | 1000.2 | 0.0035753 | 0.20909 | 185.27 | 176.681
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 129.8577 | 9.7815 | 25.0057 | 999.886 | 0.0035812 | 0.20961 | 179.162 | 170.46
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 126.8787 | 9.154 | 25.0098 | 999.879 | 0.0035839 | 0.20954 | 173.131 | 163.727
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 123.547 | 8.5889 | 25.0127 | 999.884 | 0.003585 | 0.20943 | 167.046 | 157.005
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 118.717 | 8.3227 | 25.0195 | 999.878 | 0.0035857 | 0.20911 | 160.914 | 150.538
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 115.8816 | 7.6472 | 25.0145 | 999.877 | 0.0035838 | 0.20916 | 154.761 | 143.632
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 112.1898 | 7.1456 | 25.0114 | 999.875 | 0.0035857 | 0.20936 | 148.554 | 136.852
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 106.6335 | 7.0312 | 25.0111 | 999.878 | 0.0035884 | 0.20949 | 142.33 | 130.444
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 102.6025 | 6.5888 | 25.0158 | 1000.19 | 0.0035893 | 0.20926 | 136.137 | 123.734
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 98.169 | 6.2282 | 25.0261 | 999.877 | 0.0035902 | 0.20878 | 129.962 | 117.128
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 92.736 | 6.0955 | 25.0252 | 999.87 | 0.0035886 | 0.20866 | 123.811 | 110.772
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 89.1813 | 5.5841 | 25.0209 | 999.857 | 0.0035891 | 0.20882 | 117.7 | 104.077
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 85.4361 | 5.1016 | 25.0233 | 1000.17 | 0.0035904 | 0.20873 | 111.572 | 97.3983
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 79.9328 | 4.9714 | 25.0302 | 999.867 | 0.0035891 | 0.20827 | 105.4 | 91.0249
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 76.1327 | 4.5072 | 25.0276 | 999.866 | 0.0035887 | 0.2083 | 99.2821 | 84.3669
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 71.8128 | 4.1541 | 25.0266 | 1000.18 | 0.003591 | 0.2084 | 93.1679 | 77.8295
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 67.0108 | 3.9053 | 25.0298 | 1000.18 | 0.0035903 | 0.20812 | 87.1044 | 71.4499
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 63.9154 | 3.3029 | 25.0284 | 1000.18 | 0.0035912 | 0.20817 | 81.0317 | 64.6965
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 59.9962 | 2.853 | 25.0309 | 999.87 | 0.0035929 | 0.20809 | 74.8708 | 58.016
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 54.7845 | 2.6644 | 25.0318 | 1000.18 | 0.0035892 | 0.20776 | 68.6753 | 51.5637
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 51.164 | 2.1534 | 25.0318 | 1000.18 | 0.0035868 | 0.20753 | 62.5228 | 44.8226
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 46.8971 | 1.7872 | 25.0336 | 999.872 | 0.0035893 | 0.20754 | 56.4016 | 38.2648
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 41.824 | 1.5872 | 25.0317 | 999.87 | 0.0035859 | 0.20734 | 50.2755 | 31.8691
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 38.7673 | 0.96997 | 25.0361 | 999.864 | 0.0035825 | 0.20685 | 44.1748 | 25.0756
Populus deltoides | LI1B |  | RACiR 500 to 0 2.5 min | 32.0886 | 1.1043 | 25.0317 | 999.864 | 0.0035815 | 0.20691 | 38.0327 | 19.0025
Populus deltoides | LI1B |  | RACiR 500 to 0 | 311.3534 | 26.8393 | 24.9663 | 999.764 | 0.0050569 | 0.22256 | 440.491 | 460.459
Populus deltoides | LI1B |  | RACiR 500 to 0 | 308.9979 | 26.8898 | 24.9661 | 1000.05 | 0.0050607 | 0.22276 | 438.2 | 458.193
Populus deltoides | LI1B |  | RACiR 500 to 0 | 308.4173 | 26.6736 | 24.9672 | 1000.06 | 0.0050609 | 0.22273 | 436.638 | 456.399
Populus deltoides | LI1B |  | RACiR 500 to 0 | 307.2838 | 26.4554 | 24.9665 | 999.744 | 0.0050617 | 0.22282 | 434.441 | 453.96
Populus deltoides | LI1B |  | RACiR 500 to 0 | 306.0473 | 26.2849 | 24.9736 | 1000.07 | 0.005077 | 0.22326 | 432.173 | 451.502
Populus deltoides | LI1B |  | RACiR 500 to 0 | 305.7615 | 26.1233 | 24.9745 | 1000.06 | 0.0051043 | 0.22444 | 430.526 | 449.685
Populus deltoides | LI1B |  | RACiR 500 to 0 | 304.7361 | 26.0252 | 24.9729 | 1000.06 | 0.0051415 | 0.22608 | 428.186 | 447.237
Populus deltoides | LI1B |  | RACiR 500 to 0 | 303.8434 | 25.8347 | 24.9696 | 1000.06 | 0.0051676 | 0.22718 | 425.863 | 444.705
Populus deltoides | LI1B |  | RACiR 500 to 0 | 301.584 | 25.9526 | 24.9663 | 1000.07 | 0.0051632 | 0.22694 | 424.2 | 443.141
Populus deltoides | LI1B |  | RACiR 500 to 0 | 299.371 | 25.8539 | 24.9701 | 999.753 | 0.0051537 | 0.22614 | 421.916 | 440.727
Populus deltoides | LI1B |  | RACiR 500 to 0 | 297.2479 | 25.7854 | 24.9672 | 1000.07 | 0.0051449 | 0.22568 | 419.674 | 438.388
Populus deltoides | LI1B |  | RACiR 500 to 0 | 296.5658 | 25.5426 | 24.9677 | 1000.06 | 0.0051413 | 0.2254 | 418.035 | 436.485
Populus deltoides | LI1B |  | RACiR 500 to 0 | 294.9274 | 25.3805 | 24.9696 | 1000.06 | 0.0051393 | 0.2252 | 415.736 | 433.996
Populus deltoides | LI1B |  | RACiR 500 to 0 | 293.3136 | 25.1861 | 24.9724 | 1000.06 | 0.0051347 | 0.22485 | 413.398 | 431.434
Populus deltoides | LI1B |  | RACiR 500 to 0 | 291.9642 | 25.1116 | 24.9755 | 1000.06 | 0.0051354 | 0.22476 | 411.728 | 429.671
Populus deltoides | LI1B |  | RACiR 500 to 0 | 289.7751 | 25.0841 | 24.9727 | 1000.06 | 0.0051316 | 0.22466 | 409.416 | 427.303
Populus deltoides | LI1B |  | RACiR 500 to 0 | 287.9161 | 25.0003 | 24.9732 | 1000.06 | 0.0051326 | 0.22473 | 407.098 | 424.877
Populus deltoides | LI1B |  | RACiR 500 to 0 | 286.0252 | 25.0539 | 24.9737 | 999.753 | 0.0051303 | 0.2246 | 405.469 | 423.281
Populus deltoides | LI1B |  | RACiR 500 to 0 | 284.3299 | 24.9387 | 24.9761 | 1000.05 | 0.0051328 | 0.22468 | 403.179 | 420.849
Populus deltoides | LI1B |  | RACiR 500 to 0 | 282.1899 | 24.9292 | 24.9755 | 1000.06 | 0.0051361 | 0.22481 | 400.883 | 418.518
Populus deltoides | LI1B |  | RACiR 500 to 0 | 281.3349 | 24.7783 | 24.9733 | 1000.06 | 0.0051367 | 0.22489 | 399.288 | 416.755
Populus deltoides | LI1B |  | RACiR 500 to 0 | 279.6164 | 24.6395 | 24.9763 | 999.751 | 0.0051341 | 0.22462 | 397.042 | 414.345
Populus deltoides | LI1B |  | RACiR 500 to 0 | 278.1741 | 24.458 | 24.972 | 1000.06 | 0.0051266 | 0.22449 | 394.81 | 411.9
Populus deltoides | LI1B |  | RACiR 500 to 0 | 277.3593 | 24.264 | 24.9736 | 1000.05 | 0.0051231 | 0.22435 | 393.168 | 410.048
Populus deltoides | LI1B |  | RACiR 500 to 0 | 275.3551 | 24.2109 | 24.9725 | 999.771 | 0.0051221 | 0.22443 | 390.835 | 407.632
Populus deltoides | LI1B |  | RACiR 500 to 0 | 273.8051 | 24.0385 | 24.9816 | 1000.05 | 0.0051259 | 0.22434 | 388.522 | 405.122
Populus deltoides | LI1B |  | RACiR 500 to 0 | 271.9493 | 24.1045 | 24.9859 | 1000.05 | 0.0051307 | 0.22447 | 386.864 | 403.512
Populus deltoides | LI1B |  | RACiR 500 to 0 | 270.743 | 23.8964 | 24.9882 | 1000.05 | 0.0051347 | 0.22462 | 384.617 | 401.033
Populus deltoides | LI1B |  | RACiR 500 to 0 | 268.6655 | 23.915 | 24.9817 | 1000.05 | 0.0051402 | 0.22517 | 382.305 | 398.715
Populus deltoides | LI1B |  | RACiR 500 to 0 | 267.811 | 23.7657 | 24.9793 | 1000.05 | 0.0051406 | 0.22531 | 380.69 | 396.932
Populus deltoides | LI1B |  | RACiR 500 to 0 | 265.9976 | 23.7032 | 24.9789 | 1000.06 | 0.0051457 | 0.22561 | 378.411 | 394.566
Populus deltoides | LI1B |  | RACiR 500 to 0 | 264.6762 | 23.4779 | 24.9872 | 999.745 | 0.005147 | 0.2254 | 376.154 | 392.056
Populus deltoides | LI1B |  | RACiR 500 to 0 | 263.1787 | 23.455 | 24.9872 | 1000.05 | 0.0051481 | 0.22545 | 374.494 | 390.356
Populus deltoides | LI1B |  | RACiR 500 to 0 | 261.9072 | 23.2326 | 24.9918 | 1000.05 | 0.0051509 | 0.22546 | 372.198 | 387.812
Populus deltoides | LI1B |  | RACiR 500 to 0 | 260.7527 | 23.0089 | 24.9893 | 1000.06 | 0.0051541 | 0.22574 | 369.88 | 385.244
Populus deltoides | LI1B |  | RACiR 500 to 0 | 259.3797 | 22.9685 | 24.9902 | 1000.07 | 0.0051585 | 0.22593 | 368.205 | 383.511
Populus deltoides | LI1B |  | RACiR 500 to 0 | 257.8642 | 22.8309 | 24.9895 | 1000.06 | 0.0051635 | 0.22622 | 365.907 | 381.052
Populus deltoides | LI1B |  | RACiR 500 to 0 | 255.8553 | 22.9035 | 24.9894 | 1000.06 | 0.0051918 | 0.2275 | 363.599 | 378.8
Populus deltoides | LI1B |  | RACiR 500 to 0 | 255.229 | 22.7666 | 24.9927 | 999.741 | 0.0052131 | 0.2283 | 361.996 | 377.049
Populus deltoides | LI1B |  | RACiR 500 to 0 | 254.2246 | 22.6145 | 24.9869 | 1000.04 | 0.0052391 | 0.22957 | 359.737 | 374.621
Populus deltoides | LI1B |  | RACiR 500 to 0 | 252.9425 | 22.3955 | 24.9862 | 1000.05 | 0.0052403 | 0.22949 | 357.501 | 372.141
Populus deltoides | LI1B |  | RACiR 500 to 0 | 251.9948 | 22.1967 | 24.9878 | 1000.05 | 0.0052323 | 0.22898 | 355.875 | 370.294
Populus deltoides | LI1B |  | RACiR 500 to 0 | 250.4055 | 21.9319 | 24.9968 | 1000.05 | 0.0052161 | 0.22786 | 353.57 | 367.688
Populus deltoides | LI1B |  | RACiR 500 to 0 | 248.9096 | 21.6755 | 24.9969 | 1000.04 | 0.0052007 | 0.22714 | 351.207 | 365.036
Populus deltoides | LI1B |  | RACiR 500 to 0 | 246.7277 | 21.7733 | 24.9927 | 1000.05 | 0.0051927 | 0.22694 | 349.493 | 363.397
Populus deltoides | LI1B |  | RACiR 500 to 0 | 244.3335 | 21.8132 | 24.9896 | 1000.06 | 0.0051929 | 0.22706 | 347.164 | 361.083
Populus deltoides | LI1B |  | RACiR 500 to 0 | 241.8724 | 21.8879 | 24.9895 | 1000.05 | 0.005197 | 0.2273 | 344.876 | 358.844
Populus deltoides | LI1B |  | RACiR 500 to 0 | 241.0869 | 21.7329 | 24.9863 | 1000.05 | 0.0051987 | 0.2275 | 343.295 | 357.091
Populus deltoides | LI1B |  | RACiR 500 to 0 | 239.6736 | 21.5481 | 24.9842 | 1000.04 | 0.005195 | 0.22745 | 341.042 | 354.626
Populus deltoides | LI1B |  | RACiR 500 to 0 | 238.2879 | 21.3027 | 24.991 | 999.733 | 0.0051867 | 0.22687 | 338.783 | 352.091
Populus deltoides | LI1B |  | RACiR 500 to 0 | 237.2266 | 21.1472 | 24.9951 | 1000.05 | 0.0051808 | 0.22652 | 337.154 | 350.286
Populus deltoides | LI1B |  | RACiR 500 to 0 | 235.2804 | 21.0697 | 24.9971 | 1000.04 | 0.0051798 | 0.22647 | 334.836 | 347.862
Populus deltoides | LI1B |  | RACiR 500 to 0 | 233.7367 | 20.9403 | 24.9968 | 1000.05 | 0.0051852 | 0.22681 | 332.533 | 345.407
Populus deltoides | LI1B |  | RACiR 500 to 0 | 231.6433 | 21.061 | 24.998 | 1000.06 | 0.0051869 | 0.2269 | 330.895 | 343.872
Populus deltoides | LI1B |  | RACiR 500 to 0 | 230.6319 | 20.8046 | 25.0024 | 1000.05 | 0.005191 | 0.22701 | 328.671 | 341.367
Populus deltoides | LI1B |  | RACiR 500 to 0 | 229.131 | 20.6645 | 25.0061 | 1000.05 | 0.005197 | 0.22722 | 326.428 | 338.958
Populus deltoides | LI1B |  | RACiR 500 to 0 | 228.7564 | 20.3878 | 25.0098 | 1000.03 | 0.0052005 | 0.22729 | 324.787 | 337.022
Populus deltoides | LI1B |  | RACiR 500 to 0 | 227.4771 | 20.2089 | 25.0077 | 1000.04 | 0.0052066 | 0.22771 | 322.512 | 334.543
Populus deltoides | LI1B |  | RACiR 500 to 0 | 226.2916 | 20.0008 | 25.0009 | 1000.05 | 0.0052082 | 0.22807 | 320.228 | 332.025
Populus deltoides | LI1B |  | RACiR 500 to 0 | 224.9937 | 19.9537 | 24.9968 | 1000.04 | 0.0052104 | 0.22835 | 318.578 | 330.309
Populus deltoides | LI1B |  | RACiR 500 to 0 | 223.0543 | 19.8768 | 25.0003 | 1000.05 | 0.0052113 | 0.22827 | 316.287 | 327.914
Populus deltoides | LI1B |  | RACiR 500 to 0 | 221.5668 | 19.7028 | 25.0057 | 999.742 | 0.0052146 | 0.22827 | 313.998 | 325.426
Populus deltoides | LI1B |  | RACiR 500 to 0 | 219.3863 | 19.8227 | 25.0076 | 1000.04 | 0.0052134 | 0.22813 | 312.354 | 323.884
Populus deltoides | LI1B |  | RACiR 500 to 0 | 218.0578 | 19.634 | 25.0104 | 1000.04 | 0.0052175 | 0.22824 | 310.118 | 321.435
Populus deltoides | LI1B |  | RACiR 500 to 0 | 215.5602 | 19.6916 | 25.015 | 1000.04 | 0.0052184 | 0.22811 | 307.873 | 319.222
Populus deltoides | LI1B |  | RACiR 500 to 0 | 215.3613 | 19.4005 | 25.0152 | 1000.04 | 0.0052217 | 0.22828 | 306.315 | 317.354
Populus deltoides | LI1B |  | RACiR 500 to 0 | 214.0402 | 19.222 | 25.0093 | 1000.04 | 0.005224 | 0.22862 | 304.037 | 314.871
Populus deltoides | LI1B |  | RACiR 500 to 0 | 213.158 | 18.9467 | 25.0041 | 1000.03 | 0.0052264 | 0.22897 | 301.784 | 312.316
Populus deltoides | LI1B |  | RACiR 500 to 0 | 211.6833 | 18.9003 | 25.009 | 1000.02 | 0.0052259 | 0.22878 | 300.141 | 310.608
Populus deltoides | LI1B |  | RACiR 500 to 0 | 210.1137 | 18.7768 | 25.0104 | 1000.03 | 0.005233 | 0.22907 | 297.885 | 308.204
Populus deltoides | LI1B |  | RACiR 500 to 0 | 208.7443 | 18.6804 | 25.0123 | 1000.04 | 0.0052643 | 0.23037 | 295.595 | 305.801
Populus deltoides | LI1B |  | RACiR 500 to 0 | 207.4606 | 18.7086 | 25.0081 | 1000.04 | 0.0052884 | 0.23153 | 293.993 | 304.216
Populus deltoides | LI1B |  | RACiR 500 to 0 | 206.1299 | 18.5744 | 25.0088 | 1000.04 | 0.0053122 | 0.23238 | 291.75 | 301.82
Populus deltoides | LI1B |  | RACiR 500 to 0 | 204.4088 | 18.4438 | 25.0087 | 1000.03 | 0.0053093 | 0.23203 | 289.549 | 299.461
Populus deltoides | LI1B |  | RACiR 500 to 0 | 203.8556 | 18.1783 | 25.0082 | 1000.03 | 0.0053018 | 0.23159 | 287.979 | 297.604
Populus deltoides | LI1B |  | RACiR 500 to 0 | 202.2259 | 18.0051 | 25.0102 | 1000.03 | 0.005295 | 0.2311 | 285.727 | 295.15
Populus deltoides | LI1B |  | RACiR 500 to 0 | 200.8268 | 17.8001 | 25.0112 | 1000.02 | 0.005293 | 0.23091 | 283.464 | 292.654
Populus deltoides | LI1B |  | RACiR 500 to 0 | 199.719 | 17.6963 | 25.0074 | 999.715 | 0.0052932 | 0.23104 | 281.83 | 290.898
Populus deltoides | LI1B |  | RACiR 500 to 0 | 197.9107 | 17.6031 | 25.0034 | 1000.02 | 0.0052912 | 0.23107 | 279.558 | 288.505
Populus deltoides | LI1B |  | RACiR 500 to 0 | 196.8384 | 17.3445 | 24.9994 | 1000.02 | 0.0052903 | 0.23116 | 277.297 | 285.958
Populus deltoides | LI1B |  | RACiR 500 to 0 | 195.0931 | 17.365 | 25.0011 | 1000.02 | 0.0052876 | 0.23095 | 275.67 | 284.332
Populus deltoides | LI1B |  | RACiR 500 to 0 | 193.8807 | 17.1373 | 25.0018 | 1000.02 | 0.0052901 | 0.23104 | 273.406 | 281.815
Populus deltoides | LI1B |  | RACiR 500 to 0 | 191.8791 | 17.0927 | 24.999 | 1000.01 | 0.0052921 | 0.23123 | 271.099 | 279.437
Populus deltoides | LI1B |  | RACiR 500 to 0 | 191.2038 | 16.889 | 24.9984 | 1000.02 | 0.0052948 | 0.23137 | 269.47 | 277.585
Populus deltoides | LI1B |  | RACiR 500 to 0 | 189.6783 | 16.7283 | 24.9942 | 1000.01 | 0.0052933 | 0.23145 | 267.177 | 275.104
Populus deltoides | LI1B |  | RACiR 500 to 0 | 188.3107 | 16.5234 | 24.9954 | 1000.02 | 0.0052942 | 0.23144 | 264.887 | 272.582
Populus deltoides | LI1B |  | RACiR 500 to 0 | 187.4332 | 16.364 | 24.9917 | 1000.01 | 0.0052933 | 0.23157 | 263.251 | 270.767
Populus deltoides | LI1B |  | RACiR 500 to 0 | 185.6743 | 16.2431 | 24.9971 | 1000 | 0.0052964 | 0.23152 | 260.941 | 268.309
Populus deltoides | LI1B |  | RACiR 500 to 0 | 184.1529 | 16.0588 | 25.0039 | 1000.01 | 0.0052981 | 0.23137 | 258.631 | 265.788
Populus deltoides | LI1B |  | RACiR 500 to 0 | 182.3076 | 16.1072 | 25.0101 | 1000 | 0.0053007 | 0.23127 | 256.989 | 264.176
Populus deltoides | LI1B |  | RACiR 500 to 0 | 180.8719 | 15.9443 | 25.0064 | 1000.01 | 0.0053035 | 0.23153 | 254.725 | 261.724
Populus deltoides | LI1B |  | RACiR 500 to 0 | 178.9336 | 15.9086 | 24.9985 | 1000.01 | 0.0053032 | 0.23184 | 252.488 | 259.425
Populus deltoides | LI1B |  | RACiR 500 to 0 | 178.2932 | 15.7151 | 24.9943 | 1000.01 | 0.0053062 | 0.23216 | 250.89 | 257.615
Populus deltoides | LI1B |  | RACiR 500 to 0 | 176.8693 | 15.5371 | 24.9957 | 1000.01 | 0.0053063 | 0.23215 | 248.659 | 255.18
Populus deltoides | LI1B |  | RACiR 500 to 0 | 175.7983 | 15.2907 | 24.9968 | 1000 | 0.0053112 | 0.23238 | 246.425 | 252.673
Populus deltoides | LI1B |  | RACiR 500 to 0 | 174.882 | 15.1309 | 24.9995 | 999.996 | 0.00531 | 0.23223 | 244.836 | 250.905
Populus deltoides | LI1B |  | RACiR 500 to 0 | 173.3755 | 14.9831 | 24.9979 | 999.998 | 0.0053162 | 0.23259 | 242.548 | 248.443
Populus deltoides | LI1B |  | RACiR 500 to 0 | 171.9186 | 14.8234 | 24.9939 | 999.994 | 0.005319 | 0.23288 | 240.277 | 245.987
Populus deltoides | LI1B |  | RACiR 500 to 0 | 170.081 | 14.902 | 24.9914 | 999.998 | 0.0053224 | 0.23315 | 238.661 | 244.432
Populus deltoides | LI1B |  | RACiR 500 to 0 | 168.219 | 14.8194 | 24.9956 | 1000 | 0.0053211 | 0.23294 | 236.459 | 242.122
Populus deltoides | LI1B |  | RACiR 500 to 0 | 166.3269 | 14.7378 | 25.0019 | 999.994 | 0.0053213 | 0.23273 | 234.233 | 239.788
Populus deltoides | LI1B |  | RACiR 500 to 0 | 166.0246 | 14.4412 | 25.0051 | 1000 | 0.0053223 | 0.23265 | 232.659 | 237.898
Populus deltoides | LI1B |  | RACiR 500 to 0 | 164.8402 | 14.2044 | 25.0043 | 999.98 | 0.0053227 | 0.23272 | 230.397 | 235.372
Populus deltoides | LI1B |  | RACiR 500 to 0 | 163.6554 | 13.9754 | 25.0012 | 999.991 | 0.0053257 | 0.23301 | 228.111 | 232.83
Populus deltoides | LI1B |  | RACiR 500 to 0 | 162.6649 | 13.8627 | 25.0001 | 999.995 | 0.0053367 | 0.23357 | 226.461 | 231.051
Populus deltoides | LI1B |  | RACiR 500 to 0 | 160.9711 | 13.813 | 25.0004 | 999.989 | 0.0053681 | 0.23494 | 224.165 | 228.685
Populus deltoides | LI1B |  | RACiR 500 to 0 | 159.7368 | 13.645 | 25.0008 | 999.98 | 0.0053979 | 0.23615 | 221.88 | 226.212
Populus deltoides | LI1B |  | RACiR 500 to 0 | 157.91 | 13.7072 | 24.9971 | 999.984 | 0.0054019 | 0.2363 | 220.241 | 224.617
Populus deltoides | LI1B |  | RACiR 500 to 0 | 156.2846 | 13.4932 | 24.9959 | 999.985 | 0.0053728 | 0.23488 | 218.012 | 222.141
Populus deltoides | LI1B |  | RACiR 500 to 0 | 153.8163 | 13.475 | 24.9953 | 999.989 | 0.0053416 | 0.23346 | 215.764 | 219.842
Populus deltoides | LI1B |  | RACiR 500 to 0 | 153.3063 | 13.2049 | 24.9968 | 999.991 | 0.0053286 | 0.23287 | 214.216 | 218.002
Populus deltoides | LI1B |  | RACiR 500 to 0 | 151.7137 | 13.0788 | 24.9972 | 999.998 | 0.0053327 | 0.23318 | 211.962 | 215.596
Populus deltoides | LI1B |  | RACiR 500 to 0 | 150.3669 | 12.8847 | 25.0027 | 999.996 | 0.0053368 | 0.23323 | 209.731 | 213.145
Populus deltoides | LI1B |  | RACiR 500 to 0 | 149.3712 | 12.7599 | 25.0032 | 999.991 | 0.0053419 | 0.23354 | 208.097 | 211.368
Populus deltoides | LI1B |  | RACiR 500 to 0 | 147.4315 | 12.7078 | 24.9999 | 999.976 | 0.0053459 | 0.23388 | 205.804 | 208.997
Populus deltoides | LI1B |  | RACiR 500 to 0 | 146.3301 | 12.4728 | 24.9959 | 1000.01 | 0.0053491 | 0.23425 | 203.57 | 206.502
Populus deltoides | LI1B |  | RACiR 500 to 0 | 144.9789 | 12.4304 | 24.993 | 1000.01 | 0.0053496 | 0.23444 | 201.959 | 204.83
Populus deltoides | LI1B |  | RACiR 500 to 0 | 143.7151 | 12.216 | 24.9955 | 1000 | 0.0053501 | 0.23443 | 199.744 | 202.374
Populus deltoides | LI1B |  | RACiR 500 to 0 | 141.8335 | 12.1573 | 24.9945 | 999.97 | 0.0053562 | 0.2348 | 197.482 | 200.028
Populus deltoides | LI1B |  | RACiR 500 to 0 | 140.7711 | 12.0349 | 25.0007 | 1000.01 | 0.0053582 | 0.23465 | 195.903 | 198.308
Populus deltoides | LI1B |  | RACiR 500 to 0 | 139.1929 | 11.9059 | 25.0027 | 1000.01 | 0.0053634 | 0.23485 | 193.69 | 195.941
Populus deltoides | LI1B |  | RACiR 500 to 0 | 137.8636 | 11.7123 | 25.0017 | 1000.01 | 0.005364 | 0.23492 | 191.481 | 193.512
Populus deltoides | LI1B |  | RACiR 500 to 0 | 137.0918 | 11.5429 | 24.9973 | 999.969 | 0.0053658 | 0.23518 | 189.904 | 191.747
Populus deltoides | LI1B |  | RACiR 500 to 0 | 136.0234 | 11.2859 | 24.9951 | 1000.02 | 0.0053669 | 0.2353 | 187.678 | 189.237
Populus deltoides | LI1B |  | RACiR 500 to 0 | 135.1573 | 10.971 | 24.9969 | 999.978 | 0.0053709 | 0.23544 | 185.412 | 186.629
Populus deltoides | LI1B |  | RACiR 500 to 0 | 133.7472 | 10.9297 | 24.9983 | 1000.01 | 0.0053729 | 0.23552 | 183.776 | 184.933
Populus deltoides | LI1B |  | RACiR 500 to 0 | 131.9286 | 10.8397 | 24.9987 | 999.968 | 0.005376 | 0.23568 | 181.497 | 182.538
Populus deltoides | LI1B |  | RACiR 500 to 0 | 129.6419 | 10.8728 | 24.996 | 999.98 | 0.0053806 | 0.23603 | 179.228 | 180.277
Populus deltoides | LI1B |  | RACiR 500 to 0 | 128.7898 | 10.7137 | 24.9942 | 1000.03 | 0.0053812 | 0.23613 | 177.654 | 178.525
Populus deltoides | LI1B |  | RACiR 500 to 0 | 127.3107 | 10.5602 | 24.9876 | 1000.02 | 0.005381 | 0.23636 | 175.435 | 176.126
Populus deltoides | LI1B |  | RACiR 500 to 0 | 125.6148 | 10.4486 | 24.9844 | 1000.02 | 0.0053804 | 0.23644 | 173.206 | 173.759
Populus deltoides | LI1B |  | RACiR 500 to 0 | 124.7194 | 10.2979 | 24.9865 | 999.986 | 0.0053844 | 0.23654 | 171.627 | 172.011
Populus deltoides | LI1B |  | RACiR 500 to 0 | 123.1861 | 10.1437 | 24.9934 | 1000 | 0.0053881 | 0.2365 | 169.411 | 169.615
Populus deltoides | LI1B |  | RACiR 500 to 0 | 122.4793 | 9.796 | 24.9945 | 1000.01 | 0.0053905 | 0.23659 | 167.184 | 167.013
Populus deltoides | LI1B |  | RACiR 500 to 0 | 121.1726 | 9.7263 | 24.9945 | 1000.01 | 0.0053893 | 0.23656 | 165.555 | 165.295
Populus deltoides | LI1B |  | RACiR 500 to 0 | 119.8545 | 9.5125 | 24.9948 | 999.982 | 0.0053909 | 0.23662 | 163.281 | 162.78
Populus deltoides | LI1B |  | RACiR 500 to 0 | 118.1316 | 9.3951 | 24.9961 | 1000.01 | 0.0053922 | 0.23667 | 161.008 | 160.363
Populus deltoides | LI1B |  | RACiR 500 to 0 | 117.2071 | 9.2424 | 24.9968 | 999.978 | 0.0053931 | 0.23672 | 159.402 | 158.585
Populus deltoides | LI1B |  | RACiR 500 to 0 | 116.0692 | 9.0121 | 24.9984 | 1000.01 | 0.0054113 | 0.2375 | 157.127 | 156.055
Populus deltoides | LI1B |  | RACiR 500 to 0 | 114.7809 | 8.8355 | 25.0004 | 1000.01 | 0.0054429 | 0.23882 | 154.85 | 153.579
Populus deltoides | LI1B |  | RACiR 500 to 0 | 113.8909 | 8.7047 | 24.9969 | 1000 | 0.0054658 | 0.23989 | 153.22 | 151.802
Populus deltoides | LI1B |  | RACiR 500 to 0 | 112.3571 | 8.5644 | 24.9887 | 999.975 | 0.0054818 | 0.24075 | 150.928 | 149.345
Populus deltoides | LI1B |  | RACiR 500 to 0 | 110.734 | 8.3913 | 24.9907 | 999.972 | 0.0054742 | 0.24011 | 148.639 | 146.854
Populus deltoides | LI1B |  | RACiR 500 to 0 | 109.1986 | 8.3699 | 24.9884 | 1000.01 | 0.0054696 | 0.23989 | 147.01 | 145.184
Populus deltoides | LI1B |  | RACiR 500 to 0 | 107.2728 | 8.2829 | 24.9889 | 1000.01 | 0.0054613 | 0.23936 | 144.752 | 142.811
Populus deltoides | LI1B |  | RACiR 500 to 0 | 105.348 | 8.2103 | 24.9894 | 1000.01 | 0.0054594 | 0.23917 | 142.506 | 140.466
Populus deltoides | LI1B |  | RACiR 500 to 0 | 104.6715 | 7.9801 | 24.9915 | 1000 | 0.005457 | 0.23895 | 140.872 | 138.581
Populus deltoides | LI1B |  | RACiR 500 to 0 | 103.6027 | 7.6989 | 24.9935 | 1000 | 0.0054557 | 0.23877 | 138.614 | 136.014
Populus deltoides | LI1B |  | RACiR 500 to 0 | 102.3793 | 7.4647 | 24.9861 | 1000 | 0.005454 | 0.23895 | 136.343 | 133.481
Populus deltoides | LI1B |  | RACiR 500 to 0 | 101.0586 | 7.4005 | 24.9842 | 1000 | 0.0054534 | 0.23896 | 134.718 | 131.772
Populus deltoides | LI1B |  | RACiR 500 to 0 | 99.0397 | 7.3513 | 24.9834 | 999.965 | 0.0054574 | 0.23916 | 132.418 | 129.397
Populus deltoides | LI1B |  | RACiR 500 to 0 | 97.659 | 7.1478 | 24.9881 | 999.991 | 0.0054573 | 0.23898 | 130.17 | 126.918
Populus deltoides | LI1B |  | RACiR 500 to 0 | 96.1045 | 7.1382 | 24.9929 | 999.996 | 0.0054561 | 0.23874 | 128.572 | 125.292
Populus deltoides | LI1B |  | RACiR 500 to 0 | 94.8911 | 6.9067 | 24.9944 | 999.995 | 0.0054558 | 0.23868 | 126.357 | 122.818
Populus deltoides | LI1B |  | RACiR 500 to 0 | 93.0726 | 6.8176 | 24.9977 | 999.992 | 0.0054575 | 0.23865 | 124.123 | 120.469
Populus deltoides | LI1B |  | RACiR 500 to 0 | 92.4028 | 6.6084 | 24.9967 | 999.997 | 0.0054589 | 0.23877 | 122.537 | 118.654
Populus deltoides | LI1B |  | RACiR 500 to 0 | 91.1591 | 6.3814 | 24.9983 | 999.993 | 0.0054608 | 0.23881 | 120.298 | 116.161
Populus deltoides | LI1B |  | RACiR 500 to 0 | 89.6903 | 6.2034 | 24.9981 | 999.994 | 0.0054639 | 0.239 | 118.02 | 113.678
Populus deltoides | LI1B |  | RACiR 500 to 0 | 88.8204 | 6.0421 | 24.9951 | 999.995 | 0.0054681 | 0.23935 | 116.409 | 111.887
Populus deltoides | LI1B |  | RACiR 500 to 0 | 87.0952 | 5.9309 | 24.9907 | 999.992 | 0.0054691 | 0.23961 | 114.144 | 109.484
Populus deltoides | LI1B |  | RACiR 500 to 0 | 86.0616 | 5.6477 | 24.9923 | 999.987 | 0.0054701 | 0.23963 | 111.89 | 106.919
Populus deltoides | LI1B |  | RACiR 500 to 0 | 84.6407 | 5.6066 | 24.9961 | 999.993 | 0.0054719 | 0.23958 | 110.268 | 105.237
Populus deltoides | LI1B |  | RACiR 500 to 0 | 83.0507 | 5.4546 | 24.9994 | 999.982 | 0.0054728 | 0.2395 | 108.008 | 102.798
Populus deltoides | LI1B |  | RACiR 500 to 0 | 80.8615 | 5.4571 | 24.9999 | 999.976 | 0.0054762 | 0.23965 | 105.765 | 100.532
Populus deltoides | LI1B |  | RACiR 500 to 0 | 79.9968 | 5.2912 | 25.0003 | 1000.3 | 0.0054786 | 0.23975 | 104.171 | 98.7529
Populus deltoides | LI1B |  | RACiR 500 to 0 | 78.6094 | 5.0998 | 25.0011 | 999.94 | 0.005478 | 0.23969 | 101.949 | 96.3125
Populus deltoides | LI1B |  | RACiR 500 to 0 | 77.394 | 4.8738 | 25.0014 | 999.981 | 0.0054819 | 0.23988 | 99.736 | 93.8472
Populus deltoides | LI1B |  | RACiR 500 to 0 | 76.237 | 4.7798 | 24.9997 | 999.988 | 0.0054814 | 0.23993 | 98.1505 | 92.148
Populus deltoides | LI1B |  | RACiR 500 to 0 | 74.7703 | 4.6142 | 24.9995 | 999.986 | 0.0054866 | 0.24022 | 95.9278 | 89.7338
Populus deltoides | LI1B |  | RACiR 500 to 0 | 73.7375 | 4.3372 | 24.9955 | 999.981 | 0.0054881 | 0.24048 | 93.6839 | 87.1846
Populus deltoides | LI1B |  | RACiR 500 to 0 | 72.3714 | 4.286 | 24.9954 | 999.988 | 0.0054901 | 0.24057 | 92.064 | 85.4959
Populus deltoides | LI1B |  | RACiR 500 to 0 | 71.1888 | 4.0379 | 24.9953 | 1000.3 | 0.0054906 | 0.24061 | 89.8077 | 82.9635
Populus deltoides | LI1B |  | RACiR 500 to 0 | 69.438 | 3.9282 | 24.9947 | 999.95 | 0.005491 | 0.24066 | 87.5514 | 80.5712
Populus deltoides | LI1B |  | RACiR 500 to 0 | 68.6405 | 3.7389 | 24.9951 | 999.955 | 0.0054933 | 0.24076 | 85.9333 | 78.7429
Populus deltoides | LI1B |  | RACiR 500 to 0 | 67.1162 | 3.5716 | 25.0005 | 999.998 | 0.0054983 | 0.2408 | 83.6698 | 76.2849
Populus deltoides | LI1B |  | RACiR 500 to 0 | 65.5169 | 3.4407 | 25.0073 | 999.998 | 0.0055256 | 0.2418 | 81.4254 | 73.8861
Populus deltoides | LI1B |  | RACiR 500 to 0 | 64.5064 | 3.3207 | 25.0092 | 999.993 | 0.0055475 | 0.24268 | 79.8394 | 72.1623
Populus deltoides | LI1B |  | RACiR 500 to 0 | 62.9343 | 3.1866 | 25.0059 | 1000 | 0.0055766 | 0.24398 | 77.6011 | 69.766
Populus deltoides | LI1B |  | RACiR 500 to 0 | 61.6866 | 2.9568 | 24.9965 | 1000 | 0.005579 | 0.24426 | 75.3568 | 67.2634
Populus deltoides | LI1B |  | RACiR 500 to 0 | 60.3001 | 2.9061 | 24.9952 | 999.991 | 0.0055747 | 0.24396 | 73.7437 | 65.5802
Populus deltoides | LI1B |  | RACiR 500 to 0 | 58.9346 | 2.6955 | 24.995 | 999.988 | 0.0055662 | 0.24341 | 71.4995 | 63.097
Populus deltoides | LI1B |  | RACiR 500 to 0 | 57.2383 | 2.5718 | 24.9989 | 999.981 | 0.0055669 | 0.24321 | 69.2606 | 60.7097
Populus deltoides | LI1B |  | RACiR 500 to 0 | 56.4094 | 2.39 | 25.0009 | 999.99 | 0.0055625 | 0.24287 | 67.671 | 58.9173
Populus deltoides | LI1B |  | RACiR 500 to 0 | 54.9851 | 2.2021 | 24.9937 | 999.992 | 0.005558 | 0.24289 | 65.429 | 56.4599
Populus deltoides | LI1B |  | RACiR 500 to 0 | 53.6399 | 1.9905 | 24.9894 | 999.996 | 0.0055551 | 0.24287 | 63.1714 | 53.9627
Populus deltoides | LI1B |  | RACiR 500 to 0 | 52.8081 | 1.8081 | 24.9865 | 1000 | 0.0055482 | 0.24266 | 61.567 | 52.1551
Populus deltoides | LI1B |  | RACiR 500 to 0 | 51.4587 | 1.5871 | 24.9861 | 1000 | 0.0055385 | 0.24224 | 59.2763 | 49.6155
Populus deltoides | LI1B |  | RACiR 500 to 0 | 49.9899 | 1.3929 | 24.9903 | 1000 | 0.0055297 | 0.24173 | 56.9756 | 47.0909
Populus deltoides | LI1B |  | RACiR 500 to 0 | 48.3461 | 1.3996 | 24.9939 | 1000 | 0.00553 | 0.24167 | 55.3235 | 45.427
Populus deltoides | LI1B |  | RACiR 300 to 800 | 227.0038 | 19.2524 | 24.9598 | 1000.24 | 0.0066929 | 0.29523 | 298.168 | 327.215
Populus deltoides | LI1B |  | RACiR 300 to 800 | 227.879 | 19.3868 | 24.9593 | 999.946 | 0.0066933 | 0.29531 | 299.505 | 328.709
Populus deltoides | LI1B |  | RACiR 300 to 800 | 229.2016 | 19.4326 | 24.9579 | 999.933 | 0.0066928 | 0.29538 | 300.996 | 330.265
Populus deltoides | LI1B |  | RACiR 300 to 800 | 229.8034 | 19.6526 | 24.9584 | 1000.25 | 0.0066944 | 0.29544 | 302.353 | 331.862
Populus deltoides | LI1B |  | RACiR 300 to 800 | 230.4427 | 19.9026 | 24.9589 | 999.902 | 0.0066931 | 0.29536 | 303.882 | 333.669
Populus deltoides | LI1B |  | RACiR 300 to 800 | 231.542 | 19.9726 | 24.9623 | 999.933 | 0.0066938 | 0.29524 | 305.275 | 335.152
Populus deltoides | LI1B |  | RACiR 300 to 800 | 231.979 | 20.2953 | 24.9642 | 999.934 | 0.0066953 | 0.29528 | 306.82 | 337.046
Populus deltoides | LI1B |  | RACiR 300 to 800 | 232.9461 | 20.4341 | 24.9641 | 1000.25 | 0.0066982 | 0.29546 | 308.242 | 338.629
Populus deltoides | LI1B |  | RACiR 300 to 800 | 234.2865 | 20.5081 | 24.9633 | 999.937 | 0.0066988 | 0.29556 | 309.839 | 340.321
Populus deltoides | LI1B |  | RACiR 300 to 800 | 235.1533 | 20.6798 | 24.9614 | 999.933 | 0.0067022 | 0.29585 | 311.243 | 341.919
Populus deltoides | LI1B |  | RACiR 300 to 800 | 236.3579 | 20.7859 | 24.9612 | 1000.26 | 0.0067015 | 0.2958 | 312.85 | 343.658
Populus deltoides | LI1B |  | RACiR 300 to 800 | 237.9293 | 20.7579 | 24.9606 | 999.935 | 0.0067042 | 0.29599 | 314.313 | 345.113
Populus deltoides | LI1B |  | RACiR 300 to 800 | 238.8787 | 20.953 | 24.96 | 999.936 | 0.006707 | 0.2962 | 315.898 | 346.919
Populus deltoides | LI1B |  | RACiR 300 to 800 | 239.8704 | 21.09 | 24.9596 | 1000.25 | 0.0067072 | 0.29627 | 317.363 | 348.546
Populus deltoides | LI1B |  | RACiR 300 to 800 | 241.326 | 21.1433 | 24.9601 | 999.9 | 0.0067091 | 0.29638 | 319.006 | 350.262
Populus deltoides | LI1B |  | RACiR 300 to 800 | 241.9919 | 21.3625 | 24.9626 | 999.94 | 0.0067115 | 0.2964 | 320.432 | 351.929
Populus deltoides | LI1B |  | RACiR 300 to 800 | 243.089 | 21.4965 | 24.9689 | 999.941 | 0.0067133 | 0.29617 | 322.077 | 353.735
Populus deltoides | LI1B |  | RACiR 300 to 800 | 244.4731 | 21.5095 | 24.9711 | 1000.25 | 0.0067128 | 0.29604 | 323.571 | 355.263
Populus deltoides | LI1B |  | RACiR 300 to 800 | 245.6683 | 21.6426 | 24.9694 | 999.908 | 0.006716 | 0.29633 | 325.174 | 357.025
Populus deltoides | LI1B |  | RACiR 300 to 800 | 246.7238 | 21.7458 | 24.9696 | 999.943 | 0.0067145 | 0.29624 | 326.628 | 358.605
Populus deltoides | LI1B |  | RACiR 300 to 800 | 248.1821 | 21.7821 | 24.9683 | 1000.25 | 0.0067124 | 0.29619 | 328.254 | 360.29
Populus deltoides | LI1B |  | RACiR 300 to 800 | 249.174 | 21.9359 | 24.9657 | 999.911 | 0.0067156 | 0.29651 | 329.709 | 361.921
Populus deltoides | LI1B |  | RACiR 300 to 800 | 250.1086 | 22.1305 | 24.9658 | 999.944 | 0.0067139 | 0.29642 | 331.352 | 363.783
Populus deltoides | LI1B |  | RACiR 300 to 800 | 251.3977 | 22.1925 | 24.9665 | 999.938 | 0.0067156 | 0.2965 | 332.861 | 365.374
Populus deltoides | LI1B |  | RACiR 300 to 800 | 252.3113 | 22.4127 | 24.9697 | 999.946 | 0.0067211 | 0.29668 | 334.502 | 367.262
Populus deltoides | LI1B |  | RACiR 300 to 800 | 253.3233 | 22.5261 | 24.9749 | 999.9 | 0.0067204 | 0.29641 | 336 | 368.897
Populus deltoides | LI1B |  | RACiR 300 to 800 | 254.9591 | 22.5228 | 24.9793 | 999.939 | 0.0067229 | 0.29635 | 337.683 | 370.602
Populus deltoides | LI1B |  | RACiR 300 to 800 | 256.0105 | 22.6663 | 24.978 | 999.935 | 0.0067304 | 0.29675 | 339.142 | 372.23
Populus deltoides | LI1B |  | RACiR 300 to 800 | 257.7151 | 22.7275 | 24.9716 | 1000.25 | 0.0067466 | 0.29776 | 340.825 | 374.004
Populus deltoides | LI1B |  | RACiR 300 to 800 | 259.6488 | 22.6673 | 24.968 | 999.895 | 0.0067664 | 0.29883 | 342.318 | 375.464
Populus deltoides | LI1B |  | RACiR 300 to 800 | 260.7055 | 22.886 | 24.9657 | 999.941 | 0.0067851 | 0.2997 | 343.912 | 377.31
Populus deltoides | LI1B |  | RACiR 300 to 800 | 261.7568 | 23.0392 | 24.9662 | 999.941 | 0.0067978 | 0.30013 | 345.397 | 378.975
Populus deltoides | LI1B |  | RACiR 300 to 800 | 263.3476 | 23.0322 | 24.9671 | 999.894 | 0.006796 | 0.29983 | 347.082 | 380.675
Populus deltoides | LI1B |  | RACiR 300 to 800 | 264.0037 | 23.2314 | 24.9704 | 999.929 | 0.0067898 | 0.29923 | 348.591 | 382.406
Populus deltoides | LI1B |  | RACiR 300 to 800 | 265.1269 | 23.3555 | 24.9727 | 999.922 | 0.0067864 | 0.29883 | 350.27 | 384.23
Populus deltoides | LI1B |  | RACiR 300 to 800 | 266.6593 | 23.3171 | 24.9731 | 999.905 | 0.0067843 | 0.29867 | 351.753 | 385.694
Populus deltoides | LI1B |  | RACiR 300 to 800 | 267.5674 | 23.4785 | 24.9749 | 999.926 | 0.0067772 | 0.29822 | 353.353 | 387.478
Populus deltoides | LI1B |  | RACiR 300 to 800 | 268.4479 | 23.6182 | 24.975 | 999.924 | 0.0067722 | 0.29795 | 354.801 | 389.088
Populus deltoides | LI1B |  | RACiR 300 to 800 | 270.1547 | 23.5688 | 24.975 | 1000.25 | 0.0067661 | 0.29763 | 356.466 | 390.726
Populus deltoides | LI1B |  | RACiR 300 to 800 | 271.1533 | 23.6778 | 24.9762 | 999.926 | 0.0067597 | 0.29726 | 357.961 | 392.35
Populus deltoides | LI1B |  | RACiR 300 to 800 | 272.4566 | 23.7744 | 24.9761 | 999.927 | 0.0067592 | 0.29727 | 359.619 | 394.13
Populus deltoides | LI1B |  | RACiR 300 to 800 | 274.1941 | 23.7097 | 24.9753 | 1000.24 | 0.0067609 | 0.29744 | 361.127 | 395.59
Populus deltoides | LI1B |  | RACiR 300 to 800 | 275.4217 | 23.8129 | 24.9788 | 999.93 | 0.0067627 | 0.29744 | 362.737 | 397.329
Populus deltoides | LI1B |  | RACiR 300 to 800 | 276.6197 | 23.9047 | 24.9788 | 999.936 | 0.0067678 | 0.29776 | 364.184 | 398.892
Populus deltoides | LI1B |  | RACiR 300 to 800 | 278.1263 | 23.9411 | 24.9793 | 1000.24 | 0.0067712 | 0.29794 | 365.799 | 400.565
Populus deltoides | LI1B |  | RACiR 300 to 800 | 278.9455 | 24.1247 | 24.9814 | 999.895 | 0.006773 | 0.29796 | 367.256 | 402.232
Populus deltoides | LI1B |  | RACiR 300 to 800 | 279.9268 | 24.3222 | 24.9815 | 999.93 | 0.0067773 | 0.29818 | 368.869 | 404.069
Populus deltoides | LI1B |  | RACiR 300 to 800 | 281.1146 | 24.3995 | 24.9815 | 1000.25 | 0.006776 | 0.2981 | 370.369 | 405.668
Populus deltoides | LI1B |  | RACiR 300 to 800 | 282.2667 | 24.546 | 24.9812 | 999.902 | 0.0067782 | 0.29823 | 372.006 | 407.477
Populus deltoides | LI1B |  | RACiR 300 to 800 | 283.4778 | 24.6343 | 24.9816 | 999.941 | 0.0067808 | 0.2984 | 373.498 | 409.082
Populus deltoides | LI1B |  | RACiR 300 to 800 | 285.3794 | 24.5463 | 24.9832 | 999.941 | 0.006781 | 0.29835 | 375.159 | 410.677
Populus deltoides | LI1B |  | RACiR 300 to 800 | 286.4396 | 24.6699 | 24.9849 | 1000.25 | 0.0067842 | 0.29848 | 376.628 | 412.292
Populus deltoides | LI1B |  | RACiR 300 to 800 | 287.6603 | 24.7739 | 24.9903 | 999.914 | 0.0067848 | 0.29831 | 378.282 | 414.077
Populus deltoides | LI1B |  | RACiR 300 to 800 | 288.8382 | 24.8481 | 24.9917 | 999.934 | 0.0067849 | 0.29827 | 379.749 | 415.637
Populus deltoides | LI1B |  | RACiR 300 to 800 | 289.7146 | 25.0769 | 24.9932 | 1000.25 | 0.0067881 | 0.29839 | 381.389 | 417.535
Populus deltoides | LI1B |  | RACiR 300 to 800 | 290.7453 | 25.1989 | 24.9966 | 1000.25 | 0.0067896 | 0.29829 | 382.89 | 419.181
Populus deltoides | LI1B |  | RACiR 300 to 800 | 292.365 | 25.2006 | 24.9968 | 999.95 | 0.006788 | 0.2982 | 384.579 | 420.892
Populus deltoides | LI1B |  | RACiR 300 to 800 | 293.5341 | 25.2967 | 24.9968 | 999.957 | 0.0067892 | 0.29825 | 386.088 | 422.522
Populus deltoides | LI1B |  | RACiR 300 to 800 | 294.9693 | 25.365 | 24.995 | 999.951 | 0.0067892 | 0.29836 | 387.752 | 424.28
Populus deltoides | LI1B |  | RACiR 300 to 800 | 296.6063 | 25.3167 | 24.9968 | 1000.27 | 0.0067902 | 0.29835 | 389.268 | 425.768
Populus deltoides | LI1B |  | RACiR 300 to 800 | 297.6567 | 25.4771 | 24.9991 | 999.91 | 0.0067908 | 0.29832 | 390.896 | 427.579
Populus deltoides | LI1B |  | RACiR 300 to 800 | 298.8536 | 25.5582 | 25.0013 | 999.947 | 0.0067918 | 0.2983 | 392.402 | 429.191
Populus deltoides | LI1B |  | RACiR 300 to 800 | 300.5494 | 25.5397 | 25.0047 | 999.951 | 0.0067944 | 0.29829 | 394.082 | 430.878
Populus deltoides | LI1B |  | RACiR 300 to 800 | 301.3767 | 25.7402 | 25.0035 | 1000.27 | 0.0067953 | 0.29842 | 395.568 | 432.587
Populus deltoides | LI1B |  | RACiR 300 to 800 | 302.5355 | 25.8858 | 25.0025 | 999.916 | 0.0067947 | 0.29845 | 397.235 | 434.427
Populus deltoides | LI1B |  | RACiR 300 to 800 | 304.0663 | 25.8803 | 25.0004 | 999.954 | 0.0067965 | 0.29866 | 398.719 | 435.924
Populus deltoides | LI1B |  | RACiR 300 to 800 | 305.356 | 25.9893 | 24.9995 | 999.956 | 0.0067963 | 0.29873 | 400.384 | 437.725
Populus deltoides | LI1B |  | RACiR 300 to 800 | 306.5155 | 26.0824 | 25.0001 | 999.959 | 0.0067967 | 0.29874 | 401.883 | 439.34
Populus deltoides | LI1B |  | RACiR 300 to 800 | 308.3255 | 26.045 | 25.0003 | 1000.28 | 0.0068006 | 0.29893 | 403.55 | 440.994
Populus deltoides | LI1B |  | RACiR 300 to 800 | 309.5981 | 26.1001 | 25.0023 | 999.922 | 0.0068001 | 0.29883 | 405.07 | 442.592
Populus deltoides | LI1B |  | RACiR 300 to 800 | 310.7495 | 26.233 | 25.0038 | 999.927 | 0.0068021 | 0.29887 | 406.689 | 444.368
Populus deltoides | LI1B |  | RACiR 300 to 800 | 312.2511 | 26.2264 | 25.0004 | 999.967 | 0.0068017 | 0.29897 | 408.17 | 445.864
Populus deltoides | LI1B |  | RACiR 300 to 800 | 313.2354 | 26.4073 | 25.0016 | 999.958 | 0.0068009 | 0.2989 | 409.811 | 447.711
Populus deltoides | LI1B |  | RACiR 300 to 800 | 314.2349 | 26.5403 | 25.0013 | 999.959 | 0.0068006 | 0.29891 | 411.28 | 449.338
Populus deltoides | LI1B |  | RACiR 300 to 800 | 315.8232 | 26.5659 | 24.9977 | 1000.26 | 0.0067998 | 0.29907 | 412.937 | 451.042
Populus deltoides | LI1B |  | RACiR 300 to 800 | 316.9768 | 26.6527 | 24.9982 | 999.943 | 0.0067988 | 0.29902 | 414.424 | 452.639
Populus deltoides | LI1B |  | RACiR 300 to 800 | 318.2597 | 26.7627 | 24.9932 | 999.954 | 0.0067972 | 0.29917 | 416.057 | 454.406
Populus deltoides | LI1B |  | RACiR 300 to 800 | 319.9703 | 26.7093 | 24.9903 | 999.961 | 0.0067972 | 0.29931 | 417.582 | 455.894
Populus deltoides | LI1B |  | RACiR 300 to 800 | 321.1283 | 26.8474 | 24.9902 | 1000.28 | 0.0067979 | 0.29931 | 419.234 | 457.711
Populus deltoides | LI1B |  | RACiR 300 to 800 | 322.2807 | 26.9442 | 24.9899 | 1000.26 | 0.0067989 | 0.29936 | 420.724 | 459.32
Populus deltoides | LI1B |  | RACiR 300 to 800 | 323.9715 | 26.9227 | 24.9921 | 999.956 | 0.0068022 | 0.29942 | 422.365 | 460.962
Populus deltoides | LI1B |  | RACiR 300 to 800 | 324.7605 | 27.1 | 24.9946 | 999.954 | 0.0068022 | 0.29929 | 423.814 | 462.613
Populus deltoides | LI1B |  | RACiR 300 to 800 | 325.7336 | 27.2817 | 24.9978 | 999.96 | 0.0068015 | 0.29912 | 425.48 | 464.488
Populus deltoides | LI1B |  | RACiR 300 to 800 | 326.9325 | 27.3553 | 24.9988 | 999.968 | 0.0068006 | 0.29901 | 426.992 | 466.092
Populus deltoides | LI1B |  | RACiR 300 to 800 | 328.3099 | 27.4532 | 24.9974 | 1000.27 | 0.0068021 | 0.29917 | 428.682 | 467.909
Populus deltoides | LI1B |  | RACiR 300 to 800 | 329.6084 | 27.5232 | 24.9942 | 999.916 | 0.0068008 | 0.29928 | 430.206 | 469.523
Populus deltoides | LI1B |  | RACiR 300 to 800 | 331.567 | 27.4563 | 24.9903 | 999.961 | 0.0068035 | 0.29958 | 431.885 | 471.161
Populus deltoides | LI1B |  | RACiR 300 to 800 | 333.0573 | 27.5065 | 24.9896 | 999.963 | 0.0068155 | 0.30021 | 433.382 | 472.735
Populus deltoides | LI1B |  | RACiR 300 to 800 | 334.4924 | 27.609 | 24.988 | 999.958 | 0.006828 | 0.30087 | 434.986 | 474.473
Populus deltoides | LI1B |  | RACiR 300 to 800 | 335.9919 | 27.6287 | 24.987 | 999.961 | 0.0068356 | 0.30119 | 436.489 | 476.016
Populus deltoides | LI1B |  | RACiR 300 to 800 | 337.1847 | 27.7944 | 24.987 | 1000.27 | 0.0068433 | 0.3015 | 438.17 | 477.89
Populus deltoides | LI1B |  | RACiR 300 to 800 | 338.239 | 27.9086 | 24.9879 | 999.922 | 0.0068438 | 0.30139 | 439.671 | 479.53
Populus deltoides | LI1B |  | RACiR 300 to 800 | 339.8136 | 27.9166 | 24.9855 | 999.966 | 0.0068386 | 0.30117 | 441.376 | 481.265
Populus deltoides | LI1B |  | RACiR 300 to 800 | 340.6317 | 28.095 | 24.9864 | 999.965 | 0.0068375 | 0.30102 | 442.864 | 482.958
Populus deltoides | LI1B |  | RACiR 300 to 800 | 342.0189 | 28.1667 | 24.9855 | 999.961 | 0.0068354 | 0.30091 | 444.556 | 484.75
Populus deltoides | LI1B |  | RACiR 300 to 800 | 343.7842 | 28.0993 | 24.9812 | 999.925 | 0.0068348 | 0.30108 | 446.078 | 486.224
Populus deltoides | LI1B |  | RACiR 300 to 800 | 344.9314 | 28.2255 | 24.9854 | 999.962 | 0.0068353 | 0.30088 | 447.745 | 488.039
Populus deltoides | LI1B |  | RACiR 300 to 800 | 346.1205 | 28.3022 | 24.985 | 999.96 | 0.0068327 | 0.30079 | 449.249 | 489.643
Populus deltoides | LI1B |  | RACiR 300 to 800 | 348.0075 | 28.2299 | 24.9833 | 999.959 | 0.0068321 | 0.30081 | 450.926 | 491.267
Populus deltoides | LI1B |  | RACiR 300 to 800 | 349.267 | 28.2909 | 24.9823 | 1000.27 | 0.0068309 | 0.30077 | 452.431 | 492.855
Populus deltoides | LI1B |  | RACiR 300 to 800 | 350.9135 | 28.3154 | 24.9761 | 999.919 | 0.0068309 | 0.30106 | 454.098 | 494.571
Populus deltoides | LI1B |  | RACiR 300 to 800 | 352.6987 | 28.2145 | 24.9761 | 999.92 | 0.0068299 | 0.30099 | 455.608 | 496
Populus deltoides | LI1B |  | RACiR 300 to 800 | 353.582 | 28.4293 | 24.9777 | 999.955 | 0.006831 | 0.30097 | 457.243 | 497.874
Populus deltoides | LI1B |  | RACiR 300 to 800 | 354.585 | 28.5369 | 24.9826 | 999.967 | 0.00683 | 0.30069 | 458.727 | 499.494
Populus deltoides | LI1B |  | RACiR 300 to 800 | 355.9094 | 28.6092 | 24.9879 | 1000.27 | 0.0068324 | 0.30057 | 460.371 | 501.232
Populus deltoides | LI1B |  | RACiR 300 to 800 | 356.9687 | 28.7543 | 24.9879 | 999.921 | 0.006836 | 0.30075 | 461.882 | 502.911
Populus deltoides | LI1B |  | RACiR 300 to 800 | 358.8005 | 28.7164 | 24.9882 | 999.925 | 0.0068383 | 0.30084 | 463.6 | 504.621
Populus deltoides | LI1B |  | RACiR 300 to 800 | 360.3952 | 28.7089 | 24.9866 | 999.957 | 0.0068451 | 0.30126 | 465.068 | 506.099
Populus deltoides | LI1B |  | RACiR 300 to 800 | 361.5373 | 28.865 | 24.9848 | 999.961 | 0.0068444 | 0.30127 | 466.757 | 507.974
Populus deltoides | LI1B |  | RACiR 300 to 800 | 362.8365 | 28.921 | 24.9834 | 1000.27 | 0.0068444 | 0.3013 | 468.263 | 509.557
Populus deltoides | LI1B |  | RACiR 300 to 800 | 364.7131 | 28.8767 | 24.9793 | 999.924 | 0.0068462 | 0.3016 | 469.931 | 511.204
Populus deltoides | LI1B |  | RACiR 300 to 800 | 365.6707 | 29.0252 | 24.9772 | 999.971 | 0.0068427 | 0.3015 | 471.441 | 512.881
Populus deltoides | LI1B |  | RACiR 300 to 800 | 367.1469 | 29.0959 | 24.9752 | 999.967 | 0.0068453 | 0.30172 | 473.113 | 514.66
Populus deltoides | LI1B |  | RACiR 300 to 800 | 368.5171 | 29.1277 | 24.9735 | 999.972 | 0.0068421 | 0.30164 | 474.644 | 516.241
Populus deltoides | LI1B |  | RACiR 300 to 800 | 369.3895 | 29.3575 | 24.9739 | 999.97 | 0.0068421 | 0.30162 | 476.315 | 518.164
Populus deltoides | LI1B |  | RACiR 300 to 800 | 370.6214 | 29.4449 | 24.9729 | 999.963 | 0.0068422 | 0.30169 | 477.841 | 519.803
Populus deltoides | LI1B |  | RACiR 300 to 800 | 372.4317 | 29.3854 | 24.9753 | 999.929 | 0.0068411 | 0.30152 | 479.554 | 521.479
Populus deltoides | LI1B |  | RACiR 300 to 800 | 373.5729 | 29.4738 | 24.9792 | 999.979 | 0.0068422 | 0.30138 | 481.069 | 523.102
Populus deltoides | LI1B |  | RACiR 300 to 800 | 375.0871 | 29.5082 | 24.9818 | 999.967 | 0.0068426 | 0.30128 | 482.771 | 524.867
Populus deltoides | LI1B |  | RACiR 300 to 800 | 376.9405 | 29.4189 | 24.9816 | 999.977 | 0.006846 | 0.30147 | 484.306 | 526.335
Populus deltoides | LI1B |  | RACiR 300 to 800 | 378.3032 | 29.5412 | 24.9811 | 999.979 | 0.0068539 | 0.30185 | 485.982 | 528.163
Populus deltoides | LI1B |  | RACiR 300 to 800 | 379.5426 | 29.6384 | 24.9838 | 999.977 | 0.0068612 | 0.30207 | 487.508 | 529.813
Populus deltoides | LI1B |  | RACiR 300 to 800 | 381.3784 | 29.6082 | 24.9856 | 999.979 | 0.006869 | 0.30232 | 489.205 | 531.508
Populus deltoides | LI1B |  | RACiR 300 to 800 | 382.4366 | 29.7422 | 24.9893 | 999.943 | 0.0068729 | 0.3023 | 490.748 | 533.208
Populus deltoides | LI1B |  | RACiR 300 to 800 | 383.7962 | 29.8079 | 24.9908 | 999.934 | 0.0068708 | 0.30208 | 492.438 | 534.992
Populus deltoides | LI1B |  | RACiR 300 to 800 | 385.5198 | 29.7587 | 24.987 | 999.973 | 0.0068711 | 0.30225 | 493.974 | 536.498
Populus deltoides | LI1B |  | RACiR 300 to 800 | 386.94 | 29.827 | 24.9857 | 999.972 | 0.0068683 | 0.30215 | 495.689 | 538.306
Populus deltoides | LI1B |  | RACiR 300 to 800 | 388.4182 | 29.8443 | 24.9826 | 999.969 | 0.0068682 | 0.30225 | 497.221 | 539.877
Populus deltoides | LI1B |  | RACiR 300 to 800 | 390.4751 | 29.7402 | 24.9793 | 999.972 | 0.0068674 | 0.30236 | 498.936 | 541.51
Populus deltoides | LI1B |  | RACiR 300 to 800 | 391.573 | 29.8402 | 24.9799 | 999.959 | 0.0068651 | 0.3022 | 500.449 | 543.146
Populus deltoides | LI1B |  | RACiR 300 to 800 | 393.2934 | 29.8294 | 24.9766 | 1000.27 | 0.0068647 | 0.30234 | 502.119 | 544.828
Populus deltoides | LI1B |  | RACiR 300 to 800 | 395.096 | 29.7458 | 24.9745 | 999.935 | 0.0068643 | 0.30245 | 503.645 | 546.291
Populus deltoides | LI1B |  | RACiR 300 to 800 | 396.1033 | 29.9352 | 24.9761 | 999.959 | 0.0068641 | 0.30241 | 505.322 | 548.183
Populus deltoides | LI1B |  | RACiR 300 to 800 | 397.5465 | 29.9581 | 24.9767 | 999.966 | 0.0068655 | 0.30248 | 506.853 | 549.764
Populus deltoides | LI1B |  | RACiR 300 to 800 | 399.3331 | 29.9219 | 24.9791 | 999.964 | 0.0068674 | 0.30246 | 508.57 | 551.465
Populus deltoides | LI1B |  | RACiR 300 to 800 | 400.4147 | 30.0465 | 24.9816 | 999.925 | 0.0068688 | 0.30242 | 510.11 | 553.159
Populus deltoides | LI1B |  | RACiR 300 to 800 | 401.8952 | 30.1044 | 24.982 | 999.923 | 0.0068702 | 0.30245 | 511.807 | 554.939
Populus deltoides | LI1B |  | RACiR 300 to 800 | 403.5567 | 30.0453 | 24.9832 | 999.97 | 0.0068698 | 0.30234 | 513.349 | 556.437
Populus deltoides | LI1B |  | RACiR 300 to 800 | 405.0034 | 30.1327 | 24.9832 | 999.965 | 0.0068742 | 0.30259 | 515.037 | 558.243
Populus deltoides | LI1B |  | RACiR 300 to 800 | 406.7806 | 30.1279 | 24.982 | 999.967 | 0.0068925 | 0.3035 | 516.533 | 559.769
Populus deltoides | LI1B |  | RACiR 300 to 800 | 409.3016 | 30.008 | 24.979 | 1000.27 | 0.0069209 | 0.30494 | 518.233 | 561.385
Populus deltoides | LI1B |  | RACiR 300 to 800 | 410.9345 | 30.1007 | 24.9739 | 999.919 | 0.006944 | 0.3062 | 519.791 | 563.074
Populus deltoides | LI1B |  | RACiR 300 to 800 | 412.2984 | 30.2725 | 24.9722 | 999.927 | 0.0069658 | 0.30712 | 521.455 | 564.948
Populus deltoides | LI1B |  | RACiR 300 to 800 | 413.9745 | 30.228 | 24.9724 | 999.973 | 0.0069697 | 0.30711 | 523.035 | 566.506
Populus deltoides | LI1B |  | RACiR 300 to 800 | 414.8952 | 30.3758 | 24.9764 | 999.973 | 0.0069586 | 0.30616 | 524.79 | 568.426
Populus deltoides | LI1B |  | RACiR 300 to 800 | 416.078 | 30.4109 | 24.9787 | 999.969 | 0.0069513 | 0.30553 | 526.333 | 570.028
Populus deltoides | LI1B |  | RACiR 300 to 800 | 418.1221 | 30.2682 | 24.979 | 999.972 | 0.0069464 | 0.30513 | 528.09 | 571.661
Populus deltoides | LI1B |  | RACiR 300 to 800 | 419.2726 | 30.3503 | 24.9783 | 999.928 | 0.0069418 | 0.30481 | 529.65 | 573.321
Populus deltoides | LI1B |  | RACiR 300 to 800 | 420.8722 | 30.3708 | 24.975 | 999.93 | 0.0069412 | 0.30486 | 531.335 | 575.056
Populus deltoides | LI1B |  | RACiR 300 to 800 | 422.4876 | 30.3219 | 24.974 | 999.966 | 0.006939 | 0.3047 | 532.878 | 576.567
Populus deltoides | LI1B |  | RACiR 300 to 800 | 423.6445 | 30.4442 | 24.9774 | 999.973 | 0.0069349 | 0.3043 | 534.612 | 578.451
Populus deltoides | LI1B |  | RACiR 300 to 800 | 424.8766 | 30.5182 | 24.9775 | 999.971 | 0.0069339 | 0.30424 | 536.138 | 580.074
Populus deltoides | LI1B |  | RACiR 300 to 800 | 426.6287 | 30.4861 | 24.9775 | 1000.28 | 0.0069311 | 0.30404 | 537.891 | 581.818
Populus deltoides | LI1B |  | RACiR 300 to 800 | 427.9003 | 30.571 | 24.9773 | 999.928 | 0.0069296 | 0.304 | 539.489 | 583.524
Populus deltoides | LI1B |  | RACiR 300 to 800 | 429.7374 | 30.5583 | 24.972 | 999.935 | 0.0069309 | 0.30433 | 541.206 | 585.252
Populus deltoides | LI1B |  | RACiR 300 to 800 | 431.4965 | 30.4775 | 24.971 | 999.969 | 0.0069287 | 0.30426 | 542.758 | 586.746
Populus deltoides | LI1B |  | RACiR 300 to 800 | 432.8639 | 30.5728 | 24.969 | 999.964 | 0.006927 | 0.30428 | 544.463 | 588.566
Populus deltoides | LI1B |  | RACiR 300 to 800 | 434.4838 | 30.5618 | 24.9675 | 999.966 | 0.00693 | 0.30451 | 546.003 | 590.122
Populus deltoides | LI1B |  | RACiR 300 to 800 | 436.4498 | 30.4461 | 24.9712 | 999.981 | 0.0069276 | 0.30423 | 547.732 | 591.759
Populus deltoides | LI1B |  | RACiR 300 to 800 | 437.9834 | 30.4305 | 24.975 | 999.931 | 0.006931 | 0.30421 | 549.262 | 593.297
Populus deltoides | LI1B |  | RACiR 300 to 800 | 439.6696 | 30.4109 | 24.9797 | 999.931 | 0.0069336 | 0.30413 | 550.955 | 594.998
Populus deltoides | LI1B |  | RACiR 300 to 800 | 441.2261 | 30.3726 | 24.9833 | 999.967 | 0.0069348 | 0.304 | 552.468 | 596.49
Populus deltoides | LI1B |  | RACiR 300 to 800 | 442.358 | 30.546 | 24.9816 | 999.962 | 0.0069345 | 0.30407 | 554.179 | 598.402
Populus deltoides | LI1B |  | RACiR 300 to 800 | 443.7236 | 30.6186 | 24.9797 | 999.969 | 0.0069358 | 0.30425 | 555.753 | 600.073
Populus deltoides | LI1B |  | RACiR 300 to 800 | 445.5818 | 30.581 | 24.9777 | 999.975 | 0.0069352 | 0.30438 | 557.48 | 601.787
Populus deltoides | LI1B |  | RACiR 300 to 800 | 446.8486 | 30.6721 | 24.978 | 1000.28 | 0.0069395 | 0.30462 | 558.995 | 603.421
Populus deltoides | LI1B |  | RACiR 300 to 800 | 448.3757 | 30.7069 | 24.9796 | 999.937 | 0.0069405 | 0.30461 | 560.679 | 605.167
Populus deltoides | LI1B |  | RACiR 300 to 800 | 450.0509 | 30.6424 | 24.9814 | 999.972 | 0.006938 | 0.30442 | 562.243 | 606.686
Populus deltoides | LI1B |  | RACiR 300 to 800 | 451.6503 | 30.6631 | 24.9848 | 999.961 | 0.0069386 | 0.30429 | 564.001 | 608.489
Populus deltoides | LI1B |  | RACiR 300 to 800 | 453.2644 | 30.639 | 24.9842 | 999.972 | 0.0069387 | 0.30432 | 565.559 | 610.046
Populus deltoides | LI1B |  | RACiR 300 to 800 | 455.5035 | 30.4797 | 24.9854 | 999.968 | 0.0069396 | 0.30432 | 567.322 | 611.674
Populus deltoides | LI1B |  | RACiR 300 to 800 | 456.6955 | 30.5911 | 24.9846 | 1000.28 | 0.0069434 | 0.30457 | 568.826 | 613.316
Populus deltoides | LI1B |  | RACiR 300 to 800 | 457.7143 | 30.7913 | 24.9831 | 999.937 | 0.0069434 | 0.30465 | 570.503 | 615.217
Populus deltoides | LI1B |  | RACiR 300 to 800 | 459.3688 | 30.7603 | 24.9841 | 999.94 | 0.0069453 | 0.30471 | 572.073 | 616.778
Populus deltoides | LI1B |  | RACiR 300 to 800 | 460.4618 | 30.9274 | 24.9862 | 999.964 | 0.0069443 | 0.30456 | 573.8 | 618.698
Populus deltoides | LI1B |  | RACiR 300 to 800 | 461.9879 | 30.9353 | 24.9865 | 999.961 | 0.0069441 | 0.30456 | 575.387 | 620.318
Populus deltoides | LI1B |  | RACiR 300 to 800 | 464.5778 | 30.6999 | 24.9851 | 999.975 | 0.0069464 | 0.30477 | 577.183 | 621.904
Populus deltoides | LI1B |  | RACiR 300 to 800 | 466.1632 | 30.6911 | 24.9835 | 999.972 | 0.0069467 | 0.30485 | 578.746 | 623.482
Populus deltoides | LI1B |  | RACiR 300 to 800 | 468.1983 | 30.6212 | 24.9798 | 1000.29 | 0.0069474 | 0.3051 | 580.504 | 625.199
Populus deltoides | LI1B |  | RACiR 300 to 800 | 470.0158 | 30.5469 | 24.9771 | 999.933 | 0.0069492 | 0.30531 | 582.04 | 626.678
Populus deltoides | LI1B |  | RACiR 300 to 800 | 471.104 | 30.7284 | 24.9779 | 999.971 | 0.0069494 | 0.3053 | 583.76 | 628.606
Populus deltoides | LI1B |  | RACiR 300 to 800 | 472.2593 | 30.8206 | 24.9807 | 999.965 | 0.0069487 | 0.30514 | 585.306 | 630.27
Populus deltoides | LI1B |  | RACiR 300 to 800 | 474.1706 | 30.76 | 24.983 | 999.973 | 0.006951 | 0.30516 | 587.056 | 631.984
Populus deltoides | LI1B |  | RACiR 300 to 800 | 475.3519 | 30.8635 | 24.9864 | 999.98 | 0.0069537 | 0.30517 | 588.611 | 633.669
Populus deltoides | LI1B |  | RACiR 300 to 800 | 477.1454 | 30.8293 | 24.9898 | 1000.3 | 0.0069544 | 0.30507 | 590.368 | 635.419
Populus deltoides | LI1B |  | RACiR 300 to 800 | 479.2546 | 30.6636 | 24.9902 | 999.935 | 0.006955 | 0.30512 | 591.962 | 636.868
Populus deltoides | LI1B |  | RACiR 300 to 800 | 480.5578 | 30.7831 | 24.9919 | 999.97 | 0.0069554 | 0.30508 | 593.707 | 638.759
Populus deltoides | LI1B |  | RACiR 300 to 800 | 482.1247 | 30.7633 | 24.995 | 999.979 | 0.0069565 | 0.30499 | 595.278 | 640.334
Populus deltoides | LI1B |  | RACiR 300 to 800 | 484.2582 | 30.6328 | 24.9947 | 999.963 | 0.0069571 | 0.30503 | 597.015 | 641.961
Populus deltoides | LI1B |  | RACiR 300 to 800 | 485.6455 | 30.665 | 24.9967 | 999.967 | 0.0069571 | 0.30493 | 598.578 | 643.584
Populus deltoides | LI1B |  | RACiR 300 to 800 | 487.6097 | 30.5756 | 24.9997 | 999.971 | 0.0069596 | 0.30492 | 600.301 | 645.245
Populus deltoides | LI1B |  | RACiR 300 to 800 | 489.5164 | 30.4582 | 24.9994 | 1000.29 | 0.0069612 | 0.30503 | 601.823 | 646.667
Populus deltoides | LI1B |  | RACiR 300 to 800 | 490.7115 | 30.5909 | 25.0016 | 1000.28 | 0.0069611 | 0.30493 | 603.523 | 648.525
Populus deltoides | LI1B |  | RACiR 300 to 800 | 492.1741 | 30.6435 | 24.9977 | 999.942 | 0.0069624 | 0.30523 | 605.085 | 650.166
Populus deltoides | LI1B |  | RACiR 300 to 800 | 493.9698 | 30.6321 | 24.994 | 999.977 | 0.0069617 | 0.30539 | 606.823 | 651.914
Populus deltoides | LI1B |  | RACiR 300 to 800 | 495.3605 | 30.684 | 24.9944 | 999.972 | 0.0069626 | 0.30545 | 608.397 | 653.567
Populus deltoides | LI1B |  | RACiR 300 to 800 | 497.107 | 30.6836 | 24.9943 | 999.976 | 0.0069639 | 0.30552 | 610.156 | 655.355
Populus deltoides | LI1B |  | RACiR 300 to 800 | 499.1471 | 30.5438 | 24.9935 | 999.974 | 0.0069637 | 0.30557 | 611.761 | 656.837
Populus deltoides | LI1B |  | RACiR 300 to 800 | 500.6655 | 30.5904 | 24.9944 | 1000.29 | 0.0069641 | 0.30555 | 613.477 | 658.63
Populus deltoides | LI1B |  | RACiR 300 to 800 | 502.0987 | 30.6197 | 24.9961 | 1000.29 | 0.0069647 | 0.30552 | 615.053 | 660.258
Populus deltoides | LI1B |  | RACiR 300 to 800 | 504.2324 | 30.4945 | 24.9968 | 999.932 | 0.006967 | 0.30563 | 616.784 | 661.887
Populus deltoides | LI1B |  | RACiR 300 to 800 | 505.5523 | 30.5651 | 24.9989 | 999.973 | 0.0069714 | 0.30577 | 618.324 | 663.526
Populus deltoides | LI1B |  | RACiR 300 to 800 | 507.329 | 30.6259 | 24.9989 | 999.973 | 0.0069914 | 0.3067 | 620.036 | 665.339
Populus deltoides | LI1B |  | RACiR 300 to 800 | 509.3675 | 30.5683 | 24.9985 | 999.971 | 0.007014 | 0.30774 | 621.589 | 666.868
Populus deltoides | LI1B |  | RACiR 300 to 800 | 510.6987 | 30.7233 | 25.0036 | 999.969 | 0.0070332 | 0.30824 | 623.311 | 668.783
Populus deltoides | LI1B |  | RACiR 300 to 800 | 512.118 | 30.7884 | 25.0075 | 999.968 | 0.0070474 | 0.30855 | 624.887 | 670.458
Populus deltoides | LI1B |  | RACiR 300 to 800 | 513.9926 | 30.6868 | 25.0116 | 1000.28 | 0.0070444 | 0.30808 | 626.64 | 672.128
Populus deltoides | LI1B |  | RACiR 300 to 800 | 515.1259 | 30.7587 | 25.0099 | 999.934 | 0.0070338 | 0.30751 | 628.226 | 673.801
Populus deltoides | LI1B |  | RACiR 300 to 800 | 516.8366 | 30.7424 | 25.0026 | 999.961 | 0.0070227 | 0.30727 | 629.992 | 675.575
Populus deltoides | LI1B |  | RACiR 0 to 500 | 45.016 | 22.3496 | 24.952 | 999.984 | 0.0057567 | 0.25362 | 135.185 | 147.326
Populus deltoides | LI1B |  | RACiR 0 to 500 | 45.8631 | 22.5315 | 24.9504 | 999.981 | 0.0057581 | 0.25379 | 136.713 | 149.056
Populus deltoides | LI1B |  | RACiR 0 to 500 | 47.1235 | 22.7654 | 24.9476 | 999.982 | 0.0057581 | 0.25393 | 138.884 | 151.486
Populus deltoides | LI1B |  | RACiR 0 to 500 | 47.9904 | 23.1199 | 24.9465 | 999.986 | 0.0057604 | 0.25416 | 141.101 | 154.088
Populus deltoides | LI1B |  | RACiR 0 to 500 | 49.6106 | 23.1173 | 24.9473 | 999.953 | 0.0057631 | 0.25428 | 142.703 | 155.706
Populus deltoides | LI1B |  | RACiR 0 to 500 | 51.3207 | 23.2373 | 24.9527 | 999.954 | 0.0057668 | 0.25429 | 144.928 | 158.08
Populus deltoides | LI1B |  | RACiR 0 to 500 | 53.215 | 23.2838 | 24.9604 | 999.983 | 0.0057698 | 0.25415 | 147.102 | 160.325
Populus deltoides | LI1B |  | RACiR 0 to 500 | 53.9624 | 23.4961 | 24.961 | 999.976 | 0.005773 | 0.25431 | 148.653 | 162.109
Populus deltoides | LI1B |  | RACiR 0 to 500 | 55.4583 | 23.6829 | 24.9594 | 999.984 | 0.0057746 | 0.25451 | 150.851 | 164.522
Populus deltoides | LI1B |  | RACiR 0 to 500 | 56.7387 | 23.9437 | 24.9554 | 999.981 | 0.0057781 | 0.25491 | 153.048 | 167.008
Populus deltoides | LI1B |  | RACiR 0 to 500 | 57.8827 | 24.0569 | 24.9565 | 999.983 | 0.0057808 | 0.25502 | 154.623 | 168.717
Populus deltoides | LI1B |  | RACiR 0 to 500 | 59.8217 | 24.1044 | 24.9615 | 999.966 | 0.005781 | 0.25486 | 156.86 | 171.027
Populus deltoides | LI1B |  | RACiR 0 to 500 | 61.8594 | 24.2101 | 24.9639 | 999.94 | 0.0057986 | 0.2556 | 159.084 | 173.387
Populus deltoides | LI1B |  | RACiR 0 to 500 | 63.868 | 24.1585 | 24.9687 | 999.978 | 0.005819 | 0.25629 | 160.678 | 174.951
Populus deltoides | LI1B |  | RACiR 0 to 500 | 65.7223 | 24.3538 | 24.969 | 999.983 | 0.0058521 | 0.25769 | 162.824 | 177.326
Populus deltoides | LI1B |  | RACiR 0 to 500 | 67.6584 | 24.4556 | 24.9683 | 999.977 | 0.0058669 | 0.2582 | 165.016 | 179.647
Populus deltoides | LI1B |  | RACiR 0 to 500 | 68.2175 | 24.6933 | 24.9646 | 999.974 | 0.0058649 | 0.25816 | 166.537 | 181.427
Populus deltoides | LI1B |  | RACiR 0 to 500 | 69.1864 | 24.9572 | 24.9597 | 999.971 | 0.0058539 | 0.25767 | 168.746 | 183.924
Populus deltoides | LI1B |  | RACiR 0 to 500 | 70.3348 | 25.2159 | 24.9553 | 999.941 | 0.0058512 | 0.25764 | 170.947 | 186.412
Populus deltoides | LI1B |  | RACiR 0 to 500 | 71.2456 | 25.3494 | 24.9571 | 999.929 | 0.005846 | 0.25725 | 172.555 | 188.174
Populus deltoides | LI1B |  | RACiR 0 to 500 | 72.9997 | 25.4452 | 24.9607 | 999.962 | 0.0058457 | 0.25706 | 174.803 | 190.544
Populus deltoides | LI1B |  | RACiR 0 to 500 | 74.2985 | 25.64 | 24.9654 | 999.96 | 0.0058451 | 0.25684 | 176.987 | 192.95
Populus deltoides | LI1B |  | RACiR 0 to 500 | 75.8366 | 25.6158 | 24.9634 | 999.925 | 0.0058369 | 0.25654 | 178.581 | 194.537
Populus deltoides | LI1B |  | RACiR 0 to 500 | 77.4291 | 25.7251 | 24.9643 | 999.92 | 0.0058283 | 0.25618 | 180.786 | 196.877
Populus deltoides | LI1B |  | RACiR 0 to 500 | 79.2803 | 25.7835 | 24.9647 | 999.964 | 0.0058211 | 0.25591 | 183.016 | 199.188
Populus deltoides | LI1B |  | RACiR 0 to 500 | 79.8722 | 26.0369 | 24.9659 | 999.972 | 0.0058222 | 0.256 | 184.588 | 201.035
Populus deltoides | LI1B |  | RACiR 0 to 500 | 81.1103 | 26.327 | 24.9643 | 1000.28 | 0.0058301 | 0.25653 | 186.781 | 203.55
Populus deltoides | LI1B |  | RACiR 0 to 500 | 82.3704 | 26.5863 | 24.9635 | 999.924 | 0.005832 | 0.25671 | 189.018 | 206.074
Populus deltoides | LI1B |  | RACiR 0 to 500 | 83.5552 | 26.7029 | 24.9626 | 999.961 | 0.0058339 | 0.25688 | 190.62 | 207.815
Populus deltoides | LI1B |  | RACiR 0 to 500 | 85.5252 | 26.8026 | 24.9608 | 999.969 | 0.005838 | 0.25724 | 192.878 | 210.198
Populus deltoides | LI1B |  | RACiR 0 to 500 | 86.8435 | 27.0655 | 24.9645 | 1000.27 | 0.0058473 | 0.25765 | 195.09 | 212.705
Populus deltoides | LI1B |  | RACiR 0 to 500 | 88.1325 | 27.1632 | 24.964 | 999.969 | 0.0058513 | 0.2579 | 196.69 | 214.421
Populus deltoides | LI1B |  | RACiR 0 to 500 | 89.7289 | 27.3125 | 24.9718 | 999.955 | 0.005856 | 0.25782 | 198.945 | 216.855
Populus deltoides | LI1B |  | RACiR 0 to 500 | 91.6136 | 27.3969 | 24.9739 | 999.96 | 0.005858 | 0.25785 | 201.191 | 219.212
Populus deltoides | LI1B |  | RACiR 0 to 500 | 92.6483 | 27.54 | 24.9765 | 999.963 | 0.0058597 | 0.25787 | 202.805 | 220.99
Populus deltoides | LI1B |  | RACiR 0 to 500 | 94.0855 | 27.7621 | 24.9728 | 999.919 | 0.0058618 | 0.25821 | 205 | 223.436
Populus deltoides | LI1B |  | RACiR 0 to 500 | 95.822 | 27.9082 | 24.9714 | 999.963 | 0.0058642 | 0.25842 | 207.255 | 225.865
Populus deltoides | LI1B |  | RACiR 0 to 500 | 96.9312 | 28.0404 | 24.9716 | 999.956 | 0.0058668 | 0.25858 | 208.842 | 227.606
Populus deltoides | LI1B |  | RACiR 0 to 500 | 98.3276 | 28.2634 | 24.9741 | 999.94 | 0.0058704 | 0.2587 | 211.089 | 230.103
Populus deltoides | LI1B |  | RACiR 0 to 500 | 99.3139 | 28.5976 | 24.9755 | 999.921 | 0.005874 | 0.25887 | 213.33 | 232.71
Populus deltoides | LI1B |  | RACiR 0 to 500 | 100.9774 | 28.6072 | 24.9752 | 999.964 | 0.0058766 | 0.25905 | 214.992 | 234.401
Populus deltoides | LI1B |  | RACiR 0 to 500 | 102.5477 | 28.8022 | 24.9761 | 999.976 | 0.0058812 | 0.25928 | 217.261 | 236.894
Populus deltoides | LI1B |  | RACiR 0 to 500 | 104.617 | 28.8796 | 24.974 | 999.925 | 0.0058849 | 0.25957 | 219.55 | 239.288
Populus deltoides | LI1B |  | RACiR 0 to 500 | 105.3283 | 29.1173 | 24.9735 | 999.934 | 0.0058857 | 0.25966 | 221.163 | 241.16
Populus deltoides | LI1B |  | RACiR 0 to 500 | 106.9469 | 29.2891 | 24.9743 | 999.966 | 0.0058868 | 0.25974 | 223.456 | 243.657
Populus deltoides | LI1B |  | RACiR 0 to 500 | 108.3997 | 29.4921 | 24.9779 | 999.96 | 0.0058883 | 0.25972 | 225.743 | 246.175
Populus deltoides | LI1B |  | RACiR 0 to 500 | 109.3857 | 29.6612 | 24.9782 | 1000.27 | 0.0058891 | 0.25978 | 227.379 | 248.002
Populus deltoides | LI1B |  | RACiR 0 to 500 | 111.1116 | 29.83 | 24.9759 | 999.961 | 0.005893 | 0.26012 | 229.648 | 250.466
Populus deltoides | LI1B |  | RACiR 0 to 500 | 112.6542 | 30.0261 | 24.977 | 999.961 | 0.0058955 | 0.26024 | 231.934 | 252.98
Populus deltoides | LI1B |  | RACiR 0 to 500 | 114.6288 | 29.981 | 24.9829 | 1000.28 | 0.005909 | 0.26066 | 233.594 | 254.614
Populus deltoides | LI1B |  | RACiR 0 to 500 | 116.8497 | 30.1402 | 24.987 | 999.924 | 0.0059442 | 0.26209 | 235.85 | 257.069
Populus deltoides | LI1B |  | RACiR 0 to 500 | 119.3631 | 30.2626 | 24.9891 | 999.964 | 0.0059884 | 0.26388 | 238.104 | 259.481
Populus deltoides | LI1B |  | RACiR 0 to 500 | 120.2006 | 30.5477 | 24.99 | 999.958 | 0.0060106 | 0.26468 | 239.699 | 261.388
Populus deltoides | LI1B |  | RACiR 0 to 500 | 121.3095 | 30.8276 | 24.9902 | 1000.26 | 0.0060133 | 0.26451 | 241.986 | 263.987
Populus deltoides | LI1B |  | RACiR 0 to 500 | 122.3866 | 31.0621 | 24.9915 | 999.934 | 0.0060049 | 0.26385 | 244.285 | 266.546
Populus deltoides | LI1B |  | RACiR 0 to 500 | 123.3788 | 31.202 | 24.9857 | 999.968 | 0.0059981 | 0.26365 | 245.928 | 268.349
Populus deltoides | LI1B |  | RACiR 0 to 500 | 125.5319 | 31.2376 | 24.978 | 999.968 | 0.0059951 | 0.26371 | 248.234 | 270.716
Populus deltoides | LI1B |  | RACiR 0 to 500 | 127.1377 | 31.3998 | 24.9733 | 1000.27 | 0.0059928 | 0.26374 | 250.483 | 273.158
Populus deltoides | LI1B |  | RACiR 0 to 500 | 128.7951 | 31.3659 | 24.975 | 999.924 | 0.005992 | 0.26357 | 252.127 | 274.784
Populus deltoides | LI1B |  | RACiR 0 to 500 | 129.8685 | 31.5987 | 24.9795 | 999.963 | 0.0059867 | 0.26308 | 254.352 | 277.271
Populus deltoides | LI1B |  | RACiR 0 to 500 | 131.6164 | 31.7078 | 24.9806 | 999.967 | 0.0059861 | 0.26298 | 256.608 | 279.664
Populus deltoides | LI1B |  | RACiR 0 to 500 | 132.1531 | 31.9659 | 24.9809 | 1000.26 | 0.0059855 | 0.26293 | 258.17 | 281.504
Populus deltoides | LI1B |  | RACiR 0 to 500 | 133.6715 | 32.1545 | 24.9821 | 999.95 | 0.0059872 | 0.26298 | 260.428 | 283.977
Populus deltoides | LI1B |  | RACiR 0 to 500 | 134.9547 | 32.3811 | 24.9842 | 999.941 | 0.005987 | 0.26292 | 262.641 | 286.447
Populus deltoides | LI1B |  | RACiR 0 to 500 | 136.2958 | 32.4573 | 24.9847 | 999.951 | 0.0059876 | 0.26297 | 264.28 | 288.183
Populus deltoides | LI1B |  | RACiR 0 to 500 | 138.0326 | 32.5895 | 24.9843 | 999.933 | 0.0059898 | 0.26312 | 266.494 | 290.557
Populus deltoides | LI1B |  | RACiR 0 to 500 | 139.3784 | 32.8157 | 24.9874 | 999.941 | 0.0059907 | 0.26302 | 268.794 | 293.114
Populus deltoides | LI1B |  | RACiR 0 to 500 | 140.6868 | 32.8937 | 24.9882 | 999.909 | 0.0059932 | 0.2631 | 270.39 | 294.806
Populus deltoides | LI1B |  | RACiR 0 to 500 | 142.4415 | 33.0225 | 24.9903 | 999.951 | 0.0059947 | 0.2631 | 272.686 | 297.262
Populus deltoides | LI1B |  | RACiR 0 to 500 | 144.4385 | 33.1273 | 24.984 | 999.949 | 0.0059979 | 0.26356 | 274.902 | 299.61
Populus deltoides | LI1B |  | RACiR 0 to 500 | 145.2437 | 33.3387 | 24.9824 | 999.948 | 0.0059973 | 0.26361 | 276.509 | 301.451
Populus deltoides | LI1B |  | RACiR 0 to 500 | 146.9126 | 33.524 | 24.9819 | 999.906 | 0.0060029 | 0.26396 | 278.756 | 303.91
Populus deltoides | LI1B |  | RACiR 0 to 500 | 148.3847 | 33.7247 | 24.9863 | 999.936 | 0.0060061 | 0.26397 | 281.028 | 306.412
Populus deltoides | LI1B |  | RACiR 0 to 500 | 149.4848 | 33.8458 | 24.9916 | 999.947 | 0.0060092 | 0.26395 | 282.631 | 308.16
Populus deltoides | LI1B |  | RACiR 0 to 500 | 151.0023 | 34.0394 | 24.9935 | 999.947 | 0.0060111 | 0.26402 | 284.89 | 310.641
Populus deltoides | LI1B |  | RACiR 0 to 500 | 152.1146 | 34.3673 | 24.9903 | 999.944 | 0.0060126 | 0.26427 | 287.159 | 313.268
Populus deltoides | LI1B |  | RACiR 0 to 500 | 153.7559 | 34.395 | 24.9889 | 1000.26 | 0.0060158 | 0.26451 | 288.824 | 314.979
Populus deltoides | LI1B |  | RACiR 0 to 500 | 155.2073 | 34.6061 | 24.9888 | 999.904 | 0.0060156 | 0.26449 | 291.124 | 317.521
Populus deltoides | LI1B |  | RACiR 0 to 500 | 157.5152 | 34.5704 | 24.9952 | 999.945 | 0.0060161 | 0.26428 | 293.459 | 319.848
Populus deltoides | LI1B |  | RACiR 0 to 500 | 158.6119 | 34.7197 | 24.9927 | 999.943 | 0.0060163 | 0.26442 | 295.08 | 321.64
Populus deltoides | LI1B |  | RACiR 0 to 500 | 160.3601 | 34.8702 | 24.9919 | 999.903 | 0.006019 | 0.26464 | 297.332 | 324.071
Populus deltoides | LI1B |  | RACiR 0 to 500 | 161.8673 | 35.112 | 24.9871 | 999.934 | 0.0060238 | 0.26511 | 299.553 | 326.566
Populus deltoides | LI1B |  | RACiR 0 to 500 | 162.6103 | 35.3495 | 24.9892 | 1000.25 | 0.0060272 | 0.26522 | 301.163 | 328.435
Populus deltoides | LI1B |  | RACiR 0 to 500 | 164.1109 | 35.5663 | 24.9931 | 999.951 | 0.0060312 | 0.26531 | 303.48 | 331
Populus deltoides | LI1B |  | RACiR 0 to 500 | 165.6902 | 35.779 | 24.995 | 999.953 | 0.0060345 | 0.26546 | 305.828 | 333.589
Populus deltoides | LI1B |  | RACiR 0 to 500 | 167.8409 | 35.7158 | 24.996 | 1000.27 | 0.0060459 | 0.26597 | 307.529 | 335.251
Populus deltoides | LI1B |  | RACiR 0 to 500 | 170.5096 | 35.8188 | 24.9956 | 999.897 | 0.0060802 | 0.26757 | 309.833 | 337.7
Populus deltoides | LI1B |  | RACiR 0 to 500 | 172.8922 | 35.933 | 24.9998 | 999.939 | 0.006113 | 0.26875 | 312.112 | 340.127
Populus deltoides | LI1B |  | RACiR 0 to 500 | 173.9373 | 36.1372 | 25.0001 | 1000.25 | 0.0061258 | 0.26919 | 313.73 | 341.977
Populus deltoides | LI1B |  | RACiR 0 to 500 | 175.1727 | 36.3453 | 24.9983 | 999.939 | 0.0061198 | 0.26873 | 316.012 | 344.491
Populus deltoides | LI1B |  | RACiR 0 to 500 | 176.6181 | 36.4975 | 24.9988 | 999.944 | 0.0061142 | 0.26831 | 318.287 | 346.948
Populus deltoides | LI1B |  | RACiR 0 to 500 | 177.4671 | 36.6487 | 25.0006 | 1000.25 | 0.00611 | 0.26795 | 319.91 | 348.742
Populus deltoides | LI1B |  | RACiR 0 to 500 | 179.0613 | 36.7998 | 24.9997 | 999.944 | 0.0061053 | 0.26773 | 322.227 | 351.236
Populus deltoides | LI1B |  | RACiR 0 to 500 | 180.6041 | 36.9704 | 24.9974 | 999.939 | 0.0061005 | 0.26756 | 324.539 | 353.748
Populus deltoides | LI1B |  | RACiR 0 to 500 | 182.5666 | 36.9069 | 24.9937 | 999.899 | 0.0061018 | 0.26778 | 326.188 | 355.353
Populus deltoides | LI1B |  | RACiR 0 to 500 | 184.2374 | 37.0725 | 24.9916 | 999.927 | 0.0061021 | 0.26788 | 328.467 | 357.828
Populus deltoides | LI1B |  | RACiR 0 to 500 | 186.1596 | 37.1929 | 24.9872 | 1000.24 | 0.0061042 | 0.26817 | 330.733 | 360.241
Populus deltoides | LI1B |  | RACiR 0 to 500 | 187.0129 | 37.4243 | 24.986 | 999.898 | 0.006107 | 0.2684 | 332.358 | 362.119
Populus deltoides | LI1B |  | RACiR 0 to 500 | 188.3684 | 37.648 | 24.9914 | 999.939 | 0.0061093 | 0.2683 | 334.643 | 364.66
Populus deltoides | LI1B |  | RACiR 0 to 500 | 189.6881 | 37.8749 | 24.9973 | 999.891 | 0.0061093 | 0.26813 | 336.945 | 367.218
Populus deltoides | LI1B |  | RACiR 0 to 500 | 190.8595 | 37.9607 | 25.002 | 999.939 | 0.006108 | 0.26788 | 338.605 | 368.986
Populus deltoides | LI1B |  | RACiR 0 to 500 | 192.9627 | 38.041 | 25.0014 | 1000.24 | 0.0061123 | 0.26813 | 340.921 | 371.409
Populus deltoides | LI1B |  | RACiR 0 to 500 | 194.6971 | 38.1736 | 25.004 | 999.907 | 0.0061128 | 0.26807 | 343.233 | 373.883
Populus deltoides | LI1B |  | RACiR 0 to 500 | 196.4481 | 38.1716 | 25.0006 | 999.943 | 0.0061152 | 0.26833 | 344.871 | 375.54
Populus deltoides | LI1B |  | RACiR 0 to 500 | 198.8565 | 38.1204 | 25.0007 | 1000.25 | 0.0061152 | 0.26837 | 347.119 | 377.768
Populus deltoides | LI1B |  | RACiR 0 to 500 | 200.7376 | 38.2451 | 24.9989 | 999.946 | 0.0061193 | 0.26868 | 349.344 | 380.142
Populus deltoides | LI1B |  | RACiR 0 to 500 | 201.4069 | 38.4812 | 25.0002 | 999.943 | 0.0061189 | 0.26863 | 350.945 | 382
Populus deltoides | LI1B |  | RACiR 0 to 500 | 202.1215 | 38.8888 | 25.0002 | 1000.25 | 0.0061173 | 0.2686 | 353.231 | 384.724
Populus deltoides | LI1B |  | RACiR 0 to 500 | 203.5459 | 39.1088 | 24.9992 | 999.946 | 0.0061146 | 0.26851 | 355.565 | 387.307
Populus deltoides | LI1B |  | RACiR 0 to 500 | 204.7172 | 39.2226 | 25.0004 | 999.941 | 0.0061141 | 0.2684 | 357.254 | 389.133
Populus deltoides | LI1B |  | RACiR 0 to 500 | 206.7897 | 39.2948 | 24.9995 | 999.925 | 0.0061167 | 0.26853 | 359.573 | 391.553
Populus deltoides | LI1B |  | RACiR 0 to 500 | 208.9128 | 39.3628 | 24.9973 | 999.945 | 0.0061176 | 0.26865 | 361.931 | 394.012
Populus deltoides | LI1B |  | RACiR 0 to 500 | 211.2646 | 39.2008 | 24.9962 | 999.945 | 0.0061199 | 0.26885 | 363.615 | 395.553
Populus deltoides | LI1B |  | RACiR 0 to 500 | 213.1089 | 39.3359 | 24.9977 | 1000.26 | 0.0061251 | 0.26907 | 365.889 | 397.994
Populus deltoides | LI1B |  | RACiR 0 to 500 | 215.6547 | 39.3593 | 24.9956 | 999.944 | 0.0061403 | 0.26986 | 368.147 | 400.305
Populus deltoides | LI1B |  | RACiR 0 to 500 | 216.5039 | 39.5892 | 24.9956 | 999.942 | 0.0061464 | 0.27013 | 369.729 | 402.14
Populus deltoides | LI1B |  | RACiR 0 to 500 | 218.273 | 39.7832 | 24.9913 | 1000.24 | 0.0061548 | 0.27063 | 371.989 | 404.627
Populus deltoides | LI1B |  | RACiR 0 to 500 | 219.8819 | 39.9419 | 24.9882 | 999.937 | 0.0061503 | 0.27051 | 374.292 | 407.116
Populus deltoides | LI1B |  | RACiR 0 to 500 | 221.1564 | 40.062 | 24.9817 | 999.946 | 0.0061489 | 0.27068 | 375.943 | 408.909
Populus deltoides | LI1B |  | RACiR 0 to 500 | 223.2264 | 40.1167 | 24.9782 | 999.95 | 0.0061443 | 0.27057 | 378.323 | 411.368
Populus deltoides | LI1B |  | RACiR 0 to 500 | 225.087 | 40.3026 | 24.9798 | 999.944 | 0.0061554 | 0.27102 | 380.674 | 413.942
Populus deltoides | LI1B |  | RACiR 0 to 500 | 227.0992 | 40.2691 | 24.9817 | 1000.27 | 0.0061686 | 0.27153 | 382.324 | 415.583
Populus deltoides | LI1B |  | RACiR 0 to 500 | 228.8667 | 40.4798 | 24.9818 | 999.917 | 0.0061836 | 0.27214 | 384.585 | 418.091
Populus deltoides | LI1B |  | RACiR 0 to 500 | 231.074 | 40.5155 | 24.9744 | 999.955 | 0.006182 | 0.27221 | 386.929 | 420.498
Populus deltoides | LI1B |  | RACiR 0 to 500 | 232.3821 | 40.6284 | 24.9684 | 999.954 | 0.0061805 | 0.27236 | 388.599 | 422.301
Populus deltoides | LI1B |  | RACiR 0 to 500 | 233.8932 | 40.7986 | 24.9642 | 1000.26 | 0.0061731 | 0.27212 | 390.91 | 424.81
Populus deltoides | LI1B |  | RACiR 0 to 500 | 235.7158 | 40.9369 | 24.9646 | 999.947 | 0.0061762 | 0.27224 | 393.222 | 427.289
Populus deltoides | LI1B |  | RACiR 0 to 500 | 236.3634 | 41.1701 | 24.9658 | 999.95 | 0.0061742 | 0.27205 | 394.857 | 429.179
Populus deltoides | LI1B |  | RACiR 0 to 500 | 238.448 | 41.2641 | 24.9603 | 1000.25 | 0.0061715 | 0.27219 | 397.255 | 431.696
Populus deltoides | LI1B |  | RACiR 0 to 500 | 240.309 | 41.4277 | 24.9538 | 999.911 | 0.0061704 | 0.27242 | 399.632 | 434.273
Populus deltoides | LI1B |  | RACiR 0 to 500 | 242.3448 | 41.3329 | 24.954 | 999.953 | 0.0061713 | 0.27243 | 401.354 | 435.919
Populus deltoides | LI1B |  | RACiR 0 to 500 | 243.9633 | 41.4787 | 24.9589 | 999.95 | 0.0061717 | 0.27222 | 403.679 | 438.422
Populus deltoides | LI1B |  | RACiR 0 to 500 | 245.8731 | 41.4969 | 24.9673 | 999.91 | 0.0061686 | 0.27171 | 405.996 | 440.78
Populus deltoides | LI1B |  | RACiR 0 to 500 | 247.269 | 41.5817 | 24.9663 | 999.941 | 0.0061708 | 0.27188 | 407.64 | 442.532
Populus deltoides | LI1B |  | RACiR 0 to 500 | 249.1728 | 41.6962 | 24.9654 | 999.95 | 0.0061709 | 0.27194 | 409.978 | 445.013
Populus deltoides | LI1B |  | RACiR 0 to 500 | 251.2939 | 41.798 | 24.9643 | 999.915 | 0.0061772 | 0.27232 | 412.312 | 447.481
Populus deltoides | LI1B |  | RACiR 0 to 500 | 252.1085 | 42.0009 | 24.9668 | 999.96 | 0.0061765 | 0.2722 | 413.965 | 449.359
Populus deltoides | LI1B |  | RACiR 0 to 500 | 254.1394 | 42.0999 | 24.9653 | 999.949 | 0.0061785 | 0.27237 | 416.31 | 451.834
Populus deltoides | LI1B |  | RACiR 0 to 500 | 255.8868 | 42.2675 | 24.9636 | 1000.25 | 0.0061786 | 0.27248 | 418.653 | 454.371
Populus deltoides | LI1B |  | RACiR 0 to 500 | 258.0668 | 42.1877 | 24.9595 | 999.948 | 0.0061833 | 0.2729 | 420.342 | 456.004
Populus deltoides | LI1B |  | RACiR 0 to 500 | 259.872 | 42.304 | 24.9632 | 999.945 | 0.0061845 | 0.27286 | 422.648 | 458.454
Populus deltoides | LI1B |  | RACiR 0 to 500 | 262.0308 | 42.3642 | 24.9606 | 1000.25 | 0.0061853 | 0.27304 | 424.97 | 460.869
Populus deltoides | LI1B |  | RACiR 0 to 500 | 262.8979 | 42.5574 | 24.9627 | 999.909 | 0.0061853 | 0.27297 | 426.617 | 462.731
Populus deltoides | LI1B |  | RACiR 0 to 500 | 264.9296 | 42.7087 | 24.9606 | 999.954 | 0.0061937 | 0.27345 | 428.972 | 465.27
Populus deltoides | LI1B |  | RACiR 0 to 500 | 267.0968 | 42.9129 | 24.9589 | 1000.26 | 0.006216 | 0.27452 | 431.326 | 467.865
Populus deltoides | LI1B |  | RACiR 0 to 500 | 268.7543 | 42.9985 | 24.9585 | 999.917 | 0.0062296 | 0.27508 | 433.015 | 469.669
Populus deltoides | LI1B |  | RACiR 0 to 500 | 271.6503 | 42.9283 | 24.9559 | 999.955 | 0.006243 | 0.27567 | 435.377 | 471.993
Populus deltoides | LI1B |  | RACiR 0 to 500 | 273.2693 | 43.0214 | 24.9612 | 999.955 | 0.0062403 | 0.27516 | 437.679 | 474.42
Populus deltoides | LI1B |  | RACiR 0 to 500 | 275.0185 | 42.9444 | 24.9624 | 999.926 | 0.0062372 | 0.27491 | 439.336 | 476.015
Populus deltoides | LI1B |  | RACiR 0 to 500 | 276.368 | 43.1169 | 24.9664 | 999.913 | 0.0062324 | 0.27441 | 441.645 | 478.528
Populus deltoides | LI1C |  | RACiR 800 to 300 | 590.0671 | 39.4927 | 25.0001 | 999.897 | 0.0073738 | 0.32065 | 728.39 | 766.907
Populus deltoides | LI1C |  | RACiR 800 to 300 | 588.049 | 39.1469 | 25.0052 | 999.885 | 0.0073726 | 0.32038 | 725.35 | 763.486
Populus deltoides | LI1C |  | RACiR 800 to 300 | 585.2886 | 39.1096 | 25.0048 | 999.886 | 0.0073814 | 0.32083 | 722.229 | 760.269
Populus deltoides | LI1C |  | RACiR 800 to 300 | 582.1022 | 39.1532 | 25.0046 | 999.887 | 0.0073788 | 0.32077 | 719.127 | 757.201
Populus deltoides | LI1C |  | RACiR 800 to 300 | 579.4953 | 38.9583 | 25.0148 | 1000.19 | 0.0073842 | 0.32049 | 715.977 | 753.789
Populus deltoides | LI1C |  | RACiR 800 to 300 | 576.9739 | 38.7879 | 25.0143 | 1000.19 | 0.0073858 | 0.32058 | 712.824 | 750.42
Populus deltoides | LI1C |  | RACiR 800 to 300 | 574.2663 | 38.7099 | 25.0073 | 999.894 | 0.0073841 | 0.32091 | 709.663 | 747.158
Populus deltoides | LI1C |  | RACiR 800 to 300 | 572.0756 | 38.4301 | 25.0045 | 1000.19 | 0.0073883 | 0.32122 | 706.418 | 743.567
Populus deltoides | LI1C |  | RACiR 800 to 300 | 568.7661 | 38.3824 | 25.0123 | 999.89 | 0.0073858 | 0.32071 | 703.091 | 740.163
Populus deltoides | LI1C |  | RACiR 800 to 300 | 564.9404 | 38.5946 | 25.0175 | 999.883 | 0.0073938 | 0.32092 | 699.771 | 737.01
Populus deltoides | LI1C |  | RACiR 800 to 300 | 562.1252 | 38.5557 | 25.0072 | 999.883 | 0.0073957 | 0.3215 | 696.528 | 733.672
Populus deltoides | LI1C |  | RACiR 800 to 300 | 558.5717 | 38.6268 | 25.0105 | 999.884 | 0.0073893 | 0.32102 | 693.305 | 730.501
Populus deltoides | LI1C |  | RACiR 800 to 300 | 555.0436 | 38.7574 | 25.0161 | 999.899 | 0.0073967 | 0.32117 | 690.055 | 727.334
Populus deltoides | LI1C |  | RACiR 800 to 300 | 552.7789 | 38.496 | 25.0097 | 999.888 | 0.0073977 | 0.32149 | 686.784 | 723.74
Populus deltoides | LI1C |  | RACiR 800 to 300 | 549.5892 | 38.431 | 25.0076 | 1000.2 | 0.0073852 | 0.32098 | 683.506 | 720.387
Populus deltoides | LI1C |  | RACiR 800 to 300 | 546.1991 | 38.5466 | 25.0091 | 1000.21 | 0.0073955 | 0.3214 | 680.243 | 717.165
Populus deltoides | LI1C |  | RACiR 800 to 300 | 543.1696 | 38.4897 | 25.0098 | 999.898 | 0.0073976 | 0.32141 | 676.963 | 713.772
Populus deltoides | LI1C |  | RACiR 800 to 300 | 539.9094 | 38.4401 | 25.0114 | 999.901 | 0.0073856 | 0.32083 | 673.694 | 710.47
Populus deltoides | LI1C |  | RACiR 800 to 300 | 536.5257 | 38.6297 | 25.0094 | 999.901 | 0.0074034 | 0.32181 | 670.449 | 707.327
Populus deltoides | LI1C |  | RACiR 800 to 300 | 533.7472 | 38.5199 | 25.0075 | 999.904 | 0.0074031 | 0.32183 | 667.25 | 703.975
Populus deltoides | LI1C |  | RACiR 800 to 300 | 530.8369 | 38.3716 | 25.0112 | 1000.21 | 0.007394 | 0.32123 | 664.039 | 700.62
Populus deltoides | LI1C |  | RACiR 800 to 300 | 528.2517 | 38.3529 | 25.0142 | 1000.21 | 0.0074297 | 0.32263 | 660.82 | 697.305
Populus deltoides | LI1C |  | RACiR 800 to 300 | 526.3134 | 38.013 | 25.0134 | 999.904 | 0.0074468 | 0.3231 | 657.608 | 693.728
Populus deltoides | LI1C |  | RACiR 800 to 300 | 522.9741 | 37.9313 | 25.016 | 999.91 | 0.0074343 | 0.32211 | 654.319 | 690.339
Populus deltoides | LI1C |  | RACiR 800 to 300 | 519.4683 | 38.028 | 25.0134 | 999.9 | 0.0074374 | 0.3222 | 651.003 | 687.04
Populus deltoides | LI1C |  | RACiR 800 to 300 | 516.3649 | 37.9707 | 25.0054 | 1000.21 | 0.0074259 | 0.32194 | 647.736 | 683.688
Populus deltoides | LI1C |  | RACiR 800 to 300 | 512.9229 | 38.117 | 24.9984 | 999.908 | 0.0074282 | 0.32236 | 644.503 | 680.547
Populus deltoides | LI1C |  | RACiR 800 to 300 | 509.568 | 38.195 | 25.0005 | 999.909 | 0.0074317 | 0.32236 | 641.319 | 677.384
Populus deltoides | LI1C |  | RACiR 800 to 300 | 507.0507 | 37.8726 | 25.0087 | 999.914 | 0.0074182 | 0.32131 | 638.141 | 673.889
Populus deltoides | LI1C |  | RACiR 800 to 300 | 504.2608 | 37.7993 | 25.0072 | 999.919 | 0.0074268 | 0.32176 | 634.885 | 670.472
Populus deltoides | LI1C |  | RACiR 800 to 300 | 500.9026 | 37.8483 | 25.0057 | 1000.22 | 0.0074258 | 0.3217 | 631.624 | 667.204
Populus deltoides | LI1C |  | RACiR 800 to 300 | 497.9035 | 37.7422 | 25.0044 | 999.898 | 0.007414 | 0.32122 | 628.397 | 663.874
Populus deltoides | LI1C |  | RACiR 800 to 300 | 494.7395 | 37.8187 | 25.0022 | 999.903 | 0.0074247 | 0.32178 | 625.19 | 660.663
Populus deltoides | LI1C |  | RACiR 800 to 300 | 491.4339 | 37.8244 | 25.0077 | 999.907 | 0.0074219 | 0.3213 | 622.019 | 657.453
Populus deltoides | LI1C |  | RACiR 800 to 300 | 489.0955 | 37.5542 | 25.0071 | 999.912 | 0.0074144 | 0.32106 | 618.858 | 654.012
Populus deltoides | LI1C |  | RACiR 800 to 300 | 486.2621 | 37.5469 | 25.0025 | 999.916 | 0.0074256 | 0.32182 | 615.644 | 650.714
Populus deltoides | LI1C |  | RACiR 800 to 300 | 483.055 | 37.51 | 25.0081 | 999.932 | 0.007421 | 0.32131 | 612.447 | 647.455
Populus deltoides | LI1C |  | RACiR 800 to 300 | 480.8922 | 37.2242 | 25.0087 | 999.925 | 0.0074256 | 0.32153 | 609.251 | 643.924
Populus deltoides | LI1C |  | RACiR 800 to 300 | 478.1776 | 37.1046 | 25.0053 | 999.916 | 0.007428 | 0.32175 | 606.014 | 640.503
Populus deltoides | LI1C |  | RACiR 800 to 300 | 474.7314 | 37.1227 | 25.0043 | 1000.22 | 0.0074149 | 0.32121 | 602.738 | 637.251
Populus deltoides | LI1C |  | RACiR 800 to 300 | 471.8671 | 37.0744 | 25.0042 | 999.919 | 0.0074247 | 0.32167 | 599.488 | 633.87
Populus deltoides | LI1C |  | RACiR 800 to 300 | 468.6304 | 37.0892 | 25.0051 | 999.924 | 0.0074269 | 0.32163 | 596.241 | 630.583
Populus deltoides | LI1C |  | RACiR 800 to 300 | 465.182 | 37.1698 | 25.0012 | 1000.24 | 0.007417 | 0.32144 | 593.025 | 627.453
Populus deltoides | LI1C |  | RACiR 800 to 300 | 462.4564 | 37.1017 | 24.9991 | 999.937 | 0.0074253 | 0.32191 | 589.849 | 624.129
Populus deltoides | LI1C |  | RACiR 800 to 300 | 459.7681 | 36.9468 | 25.0013 | 1000.24 | 0.0074246 | 0.32168 | 586.702 | 620.781
Populus deltoides | LI1C |  | RACiR 800 to 300 | 456.5164 | 36.946 | 25.0022 | 999.929 | 0.0074193 | 0.32137 | 583.484 | 617.532
Populus deltoides | LI1C |  | RACiR 800 to 300 | 453.8754 | 36.8001 | 25.0026 | 1000.24 | 0.0074217 | 0.32142 | 580.306 | 614.148
Populus deltoides | LI1C |  | RACiR 800 to 300 | 451.0419 | 36.6883 | 25.0016 | 999.944 | 0.0074141 | 0.32115 | 577.152 | 610.876
Populus deltoides | LI1C |  | RACiR 800 to 300 | 447.8162 | 36.8056 | 25.0038 | 999.946 | 0.0074291 | 0.32178 | 573.992 | 607.764
Populus deltoides | LI1C |  | RACiR 800 to 300 | 445.5511 | 36.5523 | 25.0004 | 999.949 | 0.0074274 | 0.32186 | 570.845 | 604.317
Populus deltoides | LI1C |  | RACiR 800 to 300 | 442.5943 | 36.4168 | 25.0037 | 1000.26 | 0.0074163 | 0.32114 | 567.658 | 600.978
Populus deltoides | LI1C |  | RACiR 800 to 300 | 439.2983 | 36.5081 | 25.0064 | 999.951 | 0.0074257 | 0.32142 | 564.476 | 597.817
Populus deltoides | LI1C |  | RACiR 800 to 300 | 437.3303 | 36.134 | 25.0026 | 999.958 | 0.0074197 | 0.32133 | 561.31 | 594.251
Populus deltoides | LI1C |  | RACiR 800 to 300 | 434.4279 | 36.0276 | 25.0091 | 999.956 | 0.0074255 | 0.32125 | 558.045 | 590.822
Populus deltoides | LI1C |  | RACiR 800 to 300 | 431.1574 | 36.0498 | 25.0099 | 999.956 | 0.0074271 | 0.32121 | 554.787 | 587.529
Populus deltoides | LI1C |  | RACiR 800 to 300 | 428.6015 | 35.8388 | 25.0019 | 1000.26 | 0.0074127 | 0.32098 | 551.575 | 584.11
Populus deltoides | LI1C |  | RACiR 800 to 300 | 425.4068 | 35.9294 | 25.0001 | 999.968 | 0.0074222 | 0.32153 | 548.391 | 580.934
Populus deltoides | LI1C |  | RACiR 800 to 300 | 421.9317 | 35.9866 | 25.0074 | 999.967 | 0.0074212 | 0.32102 | 545.213 | 577.764
Populus deltoides | LI1C |  | RACiR 800 to 300 | 419.5668 | 35.7094 | 25.0068 | 999.967 | 0.0074124 | 0.32065 | 542.048 | 574.31
Populus deltoides | LI1C |  | RACiR 800 to 300 | 416.9792 | 35.6091 | 25.0027 | 999.972 | 0.0074199 | 0.32118 | 538.89 | 570.982
Populus deltoides | LI1C |  | RACiR 800 to 300 | 413.4668 | 35.6841 | 25.0076 | 999.975 | 0.0074163 | 0.32074 | 535.697 | 567.834
Populus deltoides | LI1C |  | RACiR 800 to 300 | 410.7939 | 35.5547 | 25.009 | 999.974 | 0.0074184 | 0.32079 | 532.537 | 564.502
Populus deltoides | LI1C |  | RACiR 800 to 300 | 407.8856 | 35.4793 | 25.0143 | 999.978 | 0.0074225 | 0.32067 | 529.376 | 561.208
Populus deltoides | LI1C |  | RACiR 800 to 300 | 405.0731 | 35.3725 | 25.0158 | 1000.28 | 0.0074189 | 0.32043 | 526.246 | 557.944
Populus deltoides | LI1C |  | RACiR 800 to 300 | 402.8625 | 35.1479 | 25.0091 | 999.978 | 0.0074222 | 0.3209 | 523.1 | 554.523
Populus deltoides | LI1C |  | RACiR 800 to 300 | 400.0779 | 35.0273 | 25.0083 | 999.987 | 0.007419 | 0.32077 | 519.914 | 551.185
Populus deltoides | LI1C |  | RACiR 800 to 300 | 397.1397 | 35.0087 | 25.004 | 999.985 | 0.0074228 | 0.32117 | 516.704 | 547.896
Populus deltoides | LI1C |  | RACiR 800 to 300 | 394.9153 | 34.7189 | 25.0009 | 999.983 | 0.0074216 | 0.32122 | 513.495 | 544.349
Populus deltoides | LI1C |  | RACiR 800 to 300 | 392.1568 | 34.5405 | 25.0006 | 999.989 | 0.0074097 | 0.3207 | 510.287 | 540.953
Populus deltoides | LI1C |  | RACiR 800 to 300 | 388.611 | 34.6821 | 25.0085 | 999.989 | 0.0074228 | 0.32089 | 507.052 | 537.784
Populus deltoides | LI1C |  | RACiR 800 to 300 | 385.8652 | 34.5165 | 25.0151 | 999.988 | 0.0074218 | 0.32045 | 503.881 | 534.404
Populus deltoides | LI1C |  | RACiR 800 to 300 | 383.0323 | 34.4566 | 25.0084 | 999.995 | 0.0074174 | 0.32063 | 500.72 | 531.163
Populus deltoides | LI1C |  | RACiR 800 to 300 | 379.9122 | 34.5239 | 25.0056 | 999.984 | 0.0074246 | 0.3211 | 497.575 | 528.021
Populus deltoides | LI1C |  | RACiR 800 to 300 | 377.5968 | 34.1892 | 25.0137 | 999.989 | 0.0074148 | 0.3202 | 494.47 | 524.556
Populus deltoides | LI1C |  | RACiR 800 to 300 | 375.0098 | 34.031 | 25.0141 | 999.996 | 0.0074173 | 0.32033 | 491.277 | 521.154
Populus deltoides | LI1C |  | RACiR 800 to 300 | 371.8877 | 34.093 | 25.0085 | 999.995 | 0.0074232 | 0.32083 | 488.103 | 517.99
Populus deltoides | LI1C |  | RACiR 800 to 300 | 369.3627 | 33.9273 | 25.0045 | 1000 | 0.0074176 | 0.32075 | 485.017 | 514.722
Populus deltoides | LI1C |  | RACiR 800 to 300 | 366.7785 | 33.8089 | 25.0031 | 1000.01 | 0.0074211 | 0.32095 | 481.932 | 511.456
Populus deltoides | LI1C |  | RACiR 800 to 300 | 363.9524 | 33.6953 | 25.0088 | 1000.01 | 0.00742 | 0.32054 | 478.828 | 508.196
Populus deltoides | LI1C |  | RACiR 800 to 300 | 361.7033 | 33.4293 | 25.0073 | 1000 | 0.0074154 | 0.32036 | 475.746 | 504.809
Populus deltoides | LI1C |  | RACiR 800 to 300 | 359.0695 | 33.3029 | 25.0036 | 1000.01 | 0.0074128 | 0.32039 | 472.637 | 501.525
Populus deltoides | LI1C |  | RACiR 800 to 300 | 356.284 | 33.2075 | 25.0046 | 1000.02 | 0.0074147 | 0.32042 | 469.477 | 498.231
Populus deltoides | LI1C |  | RACiR 800 to 300 | 354.4224 | 32.8445 | 25.0018 | 1000.01 | 0.007418 | 0.32073 | 466.32 | 494.659
Populus deltoides | LI1C |  | RACiR 800 to 300 | 351.644 | 32.6818 | 25.0047 | 1000.01 | 0.0074109 | 0.32023 | 463.128 | 491.277
Populus deltoides | LI1C |  | RACiR 800 to 300 | 348.3468 | 32.7451 | 25.0066 | 999.976 | 0.0074171 | 0.32041 | 459.9 | 488.056
Populus deltoides | LI1C |  | RACiR 800 to 300 | 346.0734 | 32.5413 | 24.9964 | 1000.01 | 0.0074187 | 0.32095 | 456.743 | 484.645
Populus deltoides | LI1C |  | RACiR 800 to 300 | 343.3568 | 32.375 | 24.9977 | 999.983 | 0.0074083 | 0.32041 | 453.613 | 481.335
Populus deltoides | LI1C |  | RACiR 800 to 300 | 340.2512 | 32.3768 | 25.0068 | 999.971 | 0.0074164 | 0.32032 | 450.482 | 478.134
Populus deltoides | LI1C |  | RACiR 800 to 300 | 338.0009 | 32.1316 | 25.0063 | 999.986 | 0.0074152 | 0.32026 | 447.422 | 474.788
Populus deltoides | LI1C |  | RACiR 800 to 300 | 335.4573 | 31.9791 | 25.0041 | 1000.02 | 0.0074138 | 0.32034 | 444.308 | 471.5
Populus deltoides | LI1C |  | RACiR 800 to 300 | 332.5296 | 31.979 | 25.005 | 1000.03 | 0.0074221 | 0.32071 | 441.197 | 468.328
Populus deltoides | LI1C |  | RACiR 800 to 300 | 330.5654 | 31.6065 | 25.0054 | 999.994 | 0.0074148 | 0.32043 | 438.096 | 464.831
Populus deltoides | LI1C |  | RACiR 800 to 300 | 327.7602 | 31.553 | 25.0046 | 1000.02 | 0.0074208 | 0.32071 | 434.971 | 461.594
Populus deltoides | LI1C |  | RACiR 800 to 300 | 324.5135 | 31.6103 | 25.0052 | 999.988 | 0.0074213 | 0.32058 | 431.873 | 458.503
Populus deltoides | LI1C |  | RACiR 800 to 300 | 322.434 | 31.283 | 25.0036 | 999.991 | 0.0074111 | 0.32022 | 428.818 | 455.113
Populus deltoides | LI1C |  | RACiR 800 to 300 | 320.3403 | 31.0276 | 25.0054 | 999.986 | 0.0074218 | 0.32063 | 425.746 | 451.712
Populus deltoides | LI1C |  | RACiR 800 to 300 | 317.9263 | 30.8051 | 25.0068 | 999.977 | 0.0074204 | 0.32044 | 422.634 | 448.333
Populus deltoides | LI1C |  | RACiR 800 to 300 | 315.9718 | 30.4454 | 25.0033 | 999.976 | 0.0074143 | 0.32038 | 419.512 | 444.83
Populus deltoides | LI1C |  | RACiR 800 to 300 | 313.4379 | 30.3087 | 25.0038 | 999.989 | 0.0074239 | 0.32079 | 416.364 | 441.482
Populus deltoides | LI1C |  | RACiR 800 to 300 | 310.2034 | 30.3267 | 25.0048 | 999.98 | 0.0074179 | 0.32045 | 413.215 | 438.326
Populus deltoides | LI1C |  | RACiR 800 to 300 | 307.8515 | 30.11 | 25.007 | 999.981 | 0.0074198 | 0.32047 | 410.119 | 434.963
Populus deltoides | LI1C |  | RACiR 800 to 300 | 305.2808 | 29.9645 | 25.0076 | 1000.01 | 0.0074236 | 0.32054 | 407.011 | 431.655
Populus deltoides | LI1C |  | RACiR 800 to 300 | 302.4061 | 29.8883 | 25.008 | 999.97 | 0.0074169 | 0.32014 | 403.95 | 428.497
Populus deltoides | LI1C |  | RACiR 800 to 300 | 300.5396 | 29.5582 | 25.0067 | 1000.01 | 0.0074225 | 0.32043 | 400.91 | 425.071
Populus deltoides | LI1C |  | RACiR 800 to 300 | 297.9799 | 29.4059 | 25.0025 | 999.977 | 0.0074208 | 0.32053 | 397.776 | 421.742
Populus deltoides | LI1C |  | RACiR 800 to 300 | 295.0494 | 29.3712 | 25.0007 | 1000.02 | 0.0074193 | 0.32058 | 394.649 | 418.543
Populus deltoides | LI1C |  | RACiR 800 to 300 | 292.8277 | 29.1305 | 25.0007 | 1000.01 | 0.0074215 | 0.32064 | 391.601 | 415.205
Populus deltoides | LI1C |  | RACiR 800 to 300 | 290.4647 | 28.9089 | 25 | 999.985 | 0.0074174 | 0.32046 | 388.533 | 411.884
Populus deltoides | LI1C |  | RACiR 800 to 300 | 287.7858 | 28.8205 | 25.0024 | 1000.02 | 0.0074226 | 0.3206 | 385.474 | 408.683
Populus deltoides | LI1C |  | RACiR 800 to 300 | 285.3281 | 28.5945 | 25.0063 | 999.976 | 0.007416 | 0.32009 | 382.396 | 405.348
Populus deltoides | LI1C |  | RACiR 800 to 300 | 282.907 | 28.4051 | 25.008 | 999.979 | 0.0074192 | 0.32014 | 379.306 | 402.019
Populus deltoides | LI1C |  | RACiR 800 to 300 | 280.5488 | 28.2168 | 25.0037 | 999.975 | 0.0074207 | 0.32041 | 376.219 | 398.693
Populus deltoides | LI1C |  | RACiR 800 to 300 | 278.7222 | 27.8283 | 24.9965 | 1000.01 | 0.0074135 | 0.32046 | 373.099 | 395.163
Populus deltoides | LI1C |  | RACiR 800 to 300 | 276.125 | 27.7143 | 24.9977 | 1000.01 | 0.0074261 | 0.32093 | 369.957 | 391.843
Populus deltoides | LI1C |  | RACiR 800 to 300 | 273.1018 | 27.6552 | 25.0047 | 1000 | 0.0074208 | 0.32022 | 366.881 | 388.666
Populus deltoides | LI1C |  | RACiR 800 to 300 | 271.0247 | 27.2894 | 25.0115 | 999.994 | 0.0074104 | 0.31945 | 363.814 | 385.214
Populus deltoides | LI1C |  | RACiR 800 to 300 | 268.5395 | 27.1628 | 25.0087 | 999.993 | 0.0074209 | 0.32008 | 360.696 | 381.903
Populus deltoides | LI1C |  | RACiR 800 to 300 | 265.6564 | 27.1285 | 25.0034 | 1000 | 0.0074199 | 0.32022 | 357.595 | 378.726
Populus deltoides | LI1C |  | RACiR 800 to 300 | 263.6652 | 26.794 | 25.0013 | 999.997 | 0.0074128 | 0.32004 | 354.546 | 375.319
Populus deltoides | LI1C |  | RACiR 800 to 300 | 261.4151 | 26.5882 | 25.0019 | 1000 | 0.0074221 | 0.32042 | 351.492 | 372
Populus deltoides | LI1C |  | RACiR 800 to 300 | 259.0004 | 26.3851 | 24.9995 | 999.997 | 0.0074164 | 0.32026 | 348.421 | 368.691
Populus deltoides | LI1C |  | RACiR 800 to 300 | 257.1421 | 25.9972 | 24.9971 | 999.989 | 0.0074089 | 0.32005 | 345.344 | 365.178
Populus deltoides | LI1C |  | RACiR 800 to 300 | 254.3324 | 25.9015 | 24.9953 | 999.992 | 0.0074006 | 0.31981 | 342.226 | 361.912
Populus deltoides | LI1C |  | RACiR 800 to 300 | 251.0898 | 25.926 | 25.0011 | 999.987 | 0.0073937 | 0.31929 | 339.123 | 358.798
Populus deltoides | LI1C |  | RACiR 800 to 300 | 249.045 | 25.6491 | 25.0058 | 999.992 | 0.0074026 | 0.31956 | 336.086 | 355.44
Populus deltoides | LI1C |  | RACiR 800 to 300 | 246.8459 | 25.4414 | 25.0015 | 999.975 | 0.0074047 | 0.31995 | 333.072 | 352.181
Populus deltoides | LI1C |  | RACiR 800 to 300 | 244.5533 | 25.2463 | 25.0013 | 999.97 | 0.0074071 | 0.3201 | 330.072 | 348.941
Populus deltoides | LI1C |  | RACiR 800 to 300 | 242.7102 | 24.8471 | 25.0082 | 999.977 | 0.0074059 | 0.31967 | 327.036 | 345.462
Populus deltoides | LI1C |  | RACiR 800 to 300 | 240.5207 | 24.5545 | 25.0121 | 999.971 | 0.0074029 | 0.31938 | 323.945 | 342.047
Populus deltoides | LI1C |  | RACiR 800 to 300 | 237.5614 | 24.5463 | 25.0161 | 999.971 | 0.0074117 | 0.3196 | 320.839 | 338.88
Populus deltoides | LI1C |  | RACiR 800 to 300 | 235.3585 | 24.2751 | 25.0173 | 999.968 | 0.007408 | 0.31938 | 317.78 | 335.514
Populus deltoides | LI1C |  | RACiR 800 to 300 | 233.0468 | 24.0801 | 25.0164 | 999.964 | 0.0074119 | 0.31961 | 314.74 | 332.236
Populus deltoides | LI1C |  | RACiR 800 to 300 | 230.4429 | 23.979 | 25.0153 | 999.963 | 0.0074154 | 0.31981 | 311.708 | 329.056
Populus deltoides | LI1C |  | RACiR 800 to 300 | 228.4836 | 23.6729 | 25.0068 | 999.952 | 0.0074077 | 0.31996 | 308.688 | 325.704
Populus deltoides | LI1C |  | RACiR 800 to 300 | 225.7175 | 23.6431 | 25.0082 | 999.951 | 0.007418 | 0.32033 | 305.68 | 322.609
Populus deltoides | LI1C |  | RACiR 800 to 300 | 223.7403 | 23.3173 | 25.0144 | 999.959 | 0.0074185 | 0.31999 | 302.712 | 319.277
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 345.6796 | 28.7998 | 25.3394 | 999.903 | 0.0050791 | 0.33221 | 438.364 | 460.378
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 342.4084 | 28.7666 | 25.3249 | 999.901 | 0.0050546 | 0.33089 | 435.285 | 457.218
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 339.1338 | 28.6722 | 25.3095 | 999.887 | 0.0050223 | 0.32886 | 432.211 | 453.999
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 336.569 | 28.3046 | 25.2923 | 1000.2 | 0.0049814 | 0.32621 | 429.185 | 450.559
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 333.9243 | 28.075 | 25.2576 | 999.903 | 0.0049417 | 0.32495 | 426.101 | 447.192
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 330.6166 | 28.015 | 25.2247 | 999.914 | 0.0048993 | 0.3233 | 422.979 | 443.953
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 327.6719 | 27.7831 | 25.2024 | 999.923 | 0.0048576 | 0.32091 | 419.907 | 440.602
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 324.4846 | 27.6708 | 25.1834 | 999.925 | 0.0048229 | 0.31893 | 416.851 | 437.382
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 322.1617 | 27.4488 | 25.1498 | 1000.23 | 0.004795 | 0.31888 | 413.773 | 434.035
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 320.5545 | 27.1408 | 25.1194 | 999.926 | 0.0047892 | 0.32076 | 410.626 | 430.548
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 318.5737 | 27.0664 | 25.102 | 999.919 | 0.0048079 | 0.32406 | 407.483 | 427.299
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 316.8293 | 26.9843 | 25.0867 | 1000.22 | 0.0048363 | 0.32823 | 404.354 | 424.063
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 315.6734 | 26.6702 | 25.0799 | 999.917 | 0.0048669 | 0.33198 | 401.265 | 420.642
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 313.9663 | 26.4785 | 25.0773 | 999.904 | 0.004898 | 0.33547 | 398.095 | 417.256
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 311.9271 | 26.4301 | 25.0635 | 1000.21 | 0.004924 | 0.33945 | 394.915 | 414.003
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 310.435 | 26.16 | 25.0497 | 999.904 | 0.0049437 | 0.3428 | 391.812 | 410.604
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 307.9546 | 26.0883 | 25.0487 | 999.902 | 0.0049575 | 0.34432 | 388.738 | 407.428
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 305.3857 | 26.0262 | 25.0493 | 999.886 | 0.0049695 | 0.34552 | 385.679 | 404.277
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 303.5162 | 25.663 | 25.0512 | 999.882 | 0.0049751 | 0.34599 | 382.622 | 400.825
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 301.277 | 25.4054 | 25.0501 | 999.884 | 0.0049759 | 0.34622 | 379.548 | 397.458
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 298.3762 | 25.3367 | 25.0505 | 1000.2 | 0.0049747 | 0.34603 | 376.447 | 394.253
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 295.8192 | 25.1464 | 25.0501 | 1000.19 | 0.0049707 | 0.34565 | 373.381 | 390.965
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 293.4684 | 24.9376 | 25.0449 | 1000.19 | 0.0049693 | 0.34583 | 370.347 | 387.684
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 291.1043 | 24.6947 | 25.0391 | 999.883 | 0.0049629 | 0.34565 | 367.276 | 384.328
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 289.058 | 24.313 | 25.0357 | 1000.18 | 0.0049556 | 0.34522 | 364.175 | 380.812
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 286.0132 | 24.2467 | 25.0373 | 999.861 | 0.0049497 | 0.34442 | 361.061 | 377.594
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 282.9136 | 24.2343 | 25.0351 | 999.863 | 0.0049455 | 0.34411 | 357.944 | 374.427
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 280.8516 | 23.8879 | 25.0299 | 999.868 | 0.0049398 | 0.34394 | 354.87 | 370.971
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 278.6543 | 23.6197 | 25.0186 | 999.866 | 0.0049366 | 0.34455 | 351.723 | 367.519
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 275.8505 | 23.5364 | 25.0087 | 1000.18 | 0.0049311 | 0.34486 | 348.565 | 364.242
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 273.2228 | 23.3362 | 25.0107 | 999.862 | 0.0049244 | 0.34401 | 345.487 | 360.927
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 270.1703 | 23.2786 | 25.0165 | 1000.17 | 0.0049227 | 0.34321 | 342.39 | 357.736
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 267.5058 | 23.1927 | 25.0085 | 1000.15 | 0.0049211 | 0.34366 | 339.338 | 354.563
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 265.4351 | 22.8685 | 25.0056 | 1000.16 | 0.0049166 | 0.34353 | 336.311 | 351.18
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 262.9169 | 22.6641 | 25.0066 | 999.846 | 0.004913 | 0.3431 | 333.242 | 347.869
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 260.2978 | 22.5384 | 25.0016 | 1000.16 | 0.0049125 | 0.34341 | 330.152 | 344.616
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 258.2117 | 22.2434 | 24.9923 | 1000.16 | 0.0049082 | 0.34382 | 327.085 | 341.22
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 255.9126 | 22.0542 | 24.9806 | 999.836 | 0.0049069 | 0.34475 | 324.011 | 337.922
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 253.5624 | 21.8739 | 24.9829 | 999.837 | 0.0049167 | 0.34563 | 320.936 | 334.635
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 252.0757 | 21.5452 | 24.984 | 999.856 | 0.0049499 | 0.34916 | 317.821 | 331.167
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 250.1004 | 21.4465 | 24.9883 | 1000.14 | 0.0049957 | 0.35379 | 314.707 | 327.933
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 247.7008 | 21.3831 | 25.0012 | 999.857 | 0.0050303 | 0.35642 | 311.642 | 324.781
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 246.079 | 20.9723 | 25.0078 | 999.846 | 0.0050507 | 0.35799 | 308.575 | 321.275
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 243.5876 | 20.7326 | 25.0232 | 1000.14 | 0.0050574 | 0.35731 | 305.498 | 317.925
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 240.4675 | 20.6742 | 25.0367 | 1000.13 | 0.0050582 | 0.35609 | 302.381 | 314.715
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 237.9363 | 20.455 | 25.0378 | 1000.12 | 0.0050531 | 0.35543 | 299.305 | 311.384
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 235.111 | 20.3271 | 25.0438 | 1000.13 | 0.005049 | 0.35439 | 296.255 | 308.168
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 232.2255 | 20.1996 | 25.0516 | 1000.13 | 0.0050429 | 0.35299 | 293.203 | 304.953
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 230.1195 | 19.8013 | 25.0629 | 1000.12 | 0.0050387 | 0.35144 | 290.193 | 301.51
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 227.6028 | 19.5528 | 25.0748 | 999.83 | 0.0050382 | 0.35016 | 287.146 | 298.178
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 224.8162 | 19.4477 | 25.0729 | 999.823 | 0.0050322 | 0.34971 | 284.088 | 294.978
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 222.4822 | 19.1358 | 25.0792 | 1000.13 | 0.005023 | 0.34828 | 281.057 | 291.601
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 219.8098 | 18.9454 | 25.0899 | 999.8 | 0.0050207 | 0.34694 | 278.02 | 288.336
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 217.3948 | 18.7181 | 25.0917 | 999.814 | 0.00502 | 0.34658 | 274.968 | 285.02
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 215.4386 | 18.2946 | 25.0993 | 1000.12 | 0.0050158 | 0.34554 | 271.921 | 281.516
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 212.627 | 18.2031 | 25.1008 | 999.801 | 0.0050144 | 0.34525 | 268.848 | 278.316
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 209.6474 | 18.1555 | 25.1049 | 1000.11 | 0.0050132 | 0.34471 | 265.774 | 275.16
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 207.6763 | 17.7364 | 25.115 | 1000.12 | 0.0050093 | 0.34342 | 262.753 | 271.684
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 205.4494 | 17.4266 | 25.1176 | 1000.11 | 0.0050054 | 0.3428 | 259.681 | 268.264
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 202.7421 | 17.2948 | 25.1161 | 1000.11 | 0.0050038 | 0.34274 | 256.555 | 264.971
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 200.4549 | 17.0181 | 25.1195 | 999.795 | 0.0050027 | 0.34229 | 253.49 | 261.595
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 197.6708 | 16.9341 | 25.1217 | 1000.1 | 0.0050031 | 0.34201 | 250.462 | 258.449
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 195.0787 | 16.8086 | 25.1183 | 999.789 | 0.0050013 | 0.34208 | 247.45 | 255.278
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 193.2951 | 16.3724 | 25.1185 | 1000.11 | 0.0049961 | 0.34152 | 244.436 | 251.791
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 190.9005 | 16.1291 | 25.1189 | 1000.12 | 0.0049904 | 0.34093 | 241.374 | 248.449
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 187.9505 | 16.0817 | 25.1206 | 1000.1 | 0.0049854 | 0.34022 | 238.342 | 245.335
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 185.6947 | 15.8195 | 25.1217 | 1000.1 | 0.0049811 | 0.33966 | 235.356 | 242.052
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 183.6535 | 15.5117 | 25.1159 | 1000.1 | 0.0049784 | 0.33988 | 232.339 | 238.691
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 181.3968 | 15.232 | 25.1174 | 999.778 | 0.0049764 | 0.33944 | 229.281 | 235.317
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 179.2465 | 14.9126 | 25.1192 | 1000.09 | 0.0049751 | 0.33909 | 226.199 | 231.881
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 176.6949 | 14.7676 | 25.1081 | 1000.08 | 0.0049717 | 0.33966 | 223.099 | 228.601
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 173.808 | 14.7104 | 25.1062 | 1000.09 | 0.0049673 | 0.33932 | 220.043 | 225.452
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 171.92 | 14.3317 | 25.1043 | 1000.09 | 0.004962 | 0.33897 | 217.047 | 222.042
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 169.8808 | 14.0023 | 25.1023 | 1000.09 | 0.0049591 | 0.33882 | 214.017 | 218.647
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 167.3323 | 13.8191 | 25.1047 | 1000.09 | 0.0049575 | 0.33832 | 210.95 | 215.362
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 165.1456 | 13.5457 | 25.0993 | 1000.08 | 0.0049528 | 0.33828 | 207.919 | 212.022
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 162.5645 | 13.3854 | 25.1012 | 1000.08 | 0.004949 | 0.33768 | 204.897 | 208.805
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 160.0384 | 13.2456 | 25.0956 | 1000.07 | 0.0049465 | 0.33783 | 201.896 | 205.63
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 158.0393 | 12.9347 | 25.0865 | 1000.06 | 0.0049422 | 0.33817 | 198.896 | 202.284
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 155.8878 | 12.6363 | 25.086 | 1000.07 | 0.0049371 | 0.33769 | 195.879 | 198.933
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 153.1301 | 12.5319 | 25.0889 | 1000.07 | 0.0049343 | 0.33706 | 192.843 | 195.758
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 150.8629 | 12.2729 | 25.0873 | 1000.07 | 0.0049294 | 0.33662 | 189.816 | 192.437
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 148.9533 | 11.9248 | 25.0778 | 1000.07 | 0.0049231 | 0.33675 | 186.819 | 189.056
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 146.6975 | 11.6588 | 25.0749 | 1000.08 | 0.0049157 | 0.3363 | 183.781 | 185.716
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 144.5275 | 11.3485 | 25.0756 | 999.765 | 0.004912 | 0.3358 | 180.702 | 182.291
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 142.2666 | 11.1263 | 25.0634 | 1000.07 | 0.0049077 | 0.33633 | 177.682 | 179.014
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 139.6251 | 10.9973 | 25.0592 | 999.756 | 0.0049023 | 0.33612 | 174.635 | 175.803
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 137.6543 | 10.6328 | 25.0619 | 1000.07 | 0.0048982 | 0.33542 | 171.612 | 172.38
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 135.5059 | 10.3549 | 25.0566 | 1000.06 | 0.0048932 | 0.33533 | 168.604 | 169.059
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 132.922 | 10.2209 | 25.0511 | 1000.06 | 0.0048916 | 0.33555 | 165.558 | 165.845
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 130.897 | 9.9198 | 25.0391 | 1000.06 | 0.0048899 | 0.33634 | 162.523 | 162.474
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 128.3831 | 9.7426 | 25.0388 | 1000.05 | 0.0048847 | 0.33581 | 159.489 | 159.228
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 125.8894 | 9.5926 | 25.0305 | 999.739 | 0.0048803 | 0.33604 | 156.486 | 156.041
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 124.1869 | 9.1858 | 25.0224 | 1000.05 | 0.0048762 | 0.33627 | 153.517 | 152.63
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 122.0946 | 8.8698 | 25.0252 | 1000.05 | 0.0048728 | 0.33564 | 150.5 | 149.262
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 119.4636 | 8.756 | 25.0126 | 1000.06 | 0.0048686 | 0.33622 | 147.438 | 146.052
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 117.2543 | 8.4744 | 25.0148 | 1000.04 | 0.0048619 | 0.33531 | 144.425 | 142.722
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 115.0055 | 8.2094 | 25.0171 | 1000.04 | 0.0048556 | 0.33448 | 141.41 | 139.407
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 113.032 | 7.8775 | 25.012 | 1000.04 | 0.0048567 | 0.3349 | 138.377 | 136.008
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 111.0562 | 7.5373 | 25.0044 | 1000.04 | 0.0048541 | 0.33527 | 135.321 | 132.576
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 108.5114 | 7.4059 | 24.9914 | 999.732 | 0.004852 | 0.33619 | 132.278 | 129.368
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 105.9863 | 7.2504 | 24.99 | 1000.05 | 0.0048533 | 0.33631 | 129.242 | 126.143
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 104.1734 | 6.8468 | 24.9914 | 1000.02 | 0.0048499 | 0.33585 | 126.223 | 122.684
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 102.1607 | 6.5247 | 24.995 | 1000.03 | 0.004859 | 0.3364 | 123.184 | 119.291
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 99.7296 | 6.3724 | 24.9932 | 1000.03 | 0.004882 | 0.33904 | 120.108 | 116.031
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 97.6337 | 6.1309 | 24.9944 | 1000.04 | 0.0049142 | 0.34244 | 117.078 | 112.728
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 95.3121 | 5.9459 | 25.0024 | 1000.02 | 0.0049418 | 0.34467 | 114.064 | 109.499
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 92.9148 | 5.7686 | 25.0036 | 999.727 | 0.0049555 | 0.34603 | 111.046 | 106.272
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 90.9924 | 5.4132 | 25.004 | 1000.01 | 0.0049602 | 0.34653 | 108.038 | 102.873
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 88.8145 | 5.1206 | 25.0118 | 1000.03 | 0.0049611 | 0.34586 | 105.01 | 99.5179
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 86.2895 | 4.9408 | 25.0214 | 1000.03 | 0.0049579 | 0.34453 | 101.985 | 96.2799
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 84.1152 | 4.6524 | 25.0243 | 1000.03 | 0.0049499 | 0.34343 | 98.9817 | 92.9529
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 81.761 | 4.4272 | 25.0247 | 999.71 | 0.0049402 | 0.34244 | 95.9715 | 89.6814
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 79.4079 | 4.1923 | 25.0299 | 1000.02 | 0.0049339 | 0.34122 | 92.9401 | 86.3807
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 77.5842 | 3.7736 | 25.0401 | 1000.02 | 0.0049272 | 0.33955 | 89.9127 | 82.8988
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 75.0918 | 3.575 | 25.0453 | 1000.03 | 0.0049203 | 0.3384 | 86.8336 | 79.5854
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 72.3097 | 3.4895 | 25.0479 | 1000.01 | 0.0049172 | 0.3378 | 83.7757 | 76.4086
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 70.4795 | 3.0983 | 25.048 | 1000 | 0.0049138 | 0.33735 | 80.7652 | 72.9715
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 68.4585 | 2.7642 | 25.0469 | 1000 | 0.00491 | 0.33698 | 77.7264 | 69.563
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 65.8744 | 2.602 | 25.0565 | 1000 | 0.004907 | 0.33577 | 74.6506 | 66.2909
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 63.636 | 2.3378 | 25.0586 | 999.992 | 0.0049045 | 0.33533 | 71.5966 | 62.9376
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 61.1417 | 2.1736 | 25.0527 | 999.997 | 0.0049038 | 0.33565 | 68.5651 | 59.7084
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 58.707 | 1.9862 | 25.0583 | 1000 | 0.0048997 | 0.33469 | 65.5511 | 56.4735
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 57.0421 | 1.5501 | 25.0598 | 1000 | 0.0048955 | 0.33415 | 62.5571 | 53.0063
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 55.0072 | 1.2275 | 25.0571 | 999.997 | 0.0048929 | 0.33409 | 59.52 | 49.6131
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 52.3701 | 1.0915 | 25.0601 | 999.998 | 0.0048861 | 0.33324 | 56.4431 | 46.3666
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 50.2479 | 0.79365 | 25.0647 | 1000.01 | 0.0048776 | 0.33217 | 53.3982 | 42.9867
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 48.0795 | 0.51994 | 25.068 | 999.993 | 0.004872 | 0.33146 | 50.3717 | 39.6526
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 45.8721 | 0.25846 | 25.0627 | 1000.01 | 0.0048691 | 0.33176 | 47.3351 | 36.3208
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 43.923 | -0.10059 | 25.0605 | 999.987 | 0.0048621 | 0.33146 | 44.2663 | 32.8557
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 41.3318 | -0.24441 | 25.0675 | 999.988 | 0.0048575 | 0.3305 | 41.1988 | 29.6107
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 38.7542 | -0.38131 | 25.0776 | 999.987 | 0.0048534 | 0.3294 | 38.1633 | 26.4049
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 37.1023 | -0.82102 | 25.073 | 999.997 | 0.0048472 | 0.3294 | 35.1414 | 22.9064
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 34.9541 | -1.1118 | 25.0647 | 999.984 | 0.0048441 | 0.3299 | 32.0762 | 19.5175
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 32.3838 | -1.2613 | 25.0714 | 999.984 | 0.0048387 | 0.32891 | 29.0006 | 16.2587
Populus deltoides | LI1C |  | RACiR 500 to 0 no VPD control | 30.7437 | -1.7043 | 25.0644 | 988.841 | 0.0048825 | 0.33227 | 26.0314 | 12.9445
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 334.1402 | 28.47 | 25.3235 | 1000.05 | 0.0054525 | 0.35296 | 420.63 | 434.261
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 329.4442 | 28.0869 | 25.3109 | 1000.06 | 0.0054435 | 0.35304 | 414.738 | 427.902
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 325.7979 | 27.2595 | 25.3035 | 999.742 | 0.0054315 | 0.3524 | 408.812 | 421.058
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 321.028 | 26.7516 | 25.3023 | 1000.05 | 0.0054209 | 0.3513 | 402.763 | 414.415
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 314.989 | 26.6978 | 25.2955 | 1000.04 | 0.0054097 | 0.35059 | 396.631 | 408.144
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 310.1805 | 26.2629 | 25.2844 | 1000.04 | 0.0053984 | 0.35034 | 390.543 | 401.535
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 305.5161 | 25.7818 | 25.2728 | 1000.04 | 0.0053871 | 0.35005 | 384.477 | 394.899
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 300.3882 | 25.3917 | 25.2674 | 999.739 | 0.0053748 | 0.34906 | 378.357 | 388.302
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 296.0634 | 24.7706 | 25.2572 | 1000.05 | 0.0053636 | 0.34872 | 372.236 | 381.472
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 290.9272 | 24.497 | 25.232 | 1000.05 | 0.00535 | 0.34944 | 366.06 | 374.933
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 284.9749 | 24.4503 | 25.2127 | 1000.06 | 0.0053324 | 0.34928 | 359.899 | 368.641
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 280.0644 | 24.0364 | 25.2018 | 1000.05 | 0.0053195 | 0.34874 | 353.823 | 362.063
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 275.1195 | 23.6483 | 25.194 | 1000.03 | 0.0053093 | 0.34818 | 347.788 | 355.554
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 270.3458 | 23.2474 | 25.1773 | 1000.04 | 0.0053001 | 0.34857 | 341.692 | 348.972
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 265.7632 | 22.726 | 25.1653 | 1000.04 | 0.0052877 | 0.34828 | 335.581 | 342.254
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 260.8899 | 22.2985 | 25.1545 | 999.737 | 0.005276 | 0.34793 | 329.455 | 335.614
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 255.9764 | 21.9525 | 25.1298 | 1000.03 | 0.0052662 | 0.34895 | 323.258 | 328.984
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 251.3924 | 21.3911 | 25.1224 | 1000.04 | 0.005257 | 0.3485 | 317.061 | 322.138
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 246.1641 | 21.0124 | 25.1216 | 999.735 | 0.0052446 | 0.34725 | 310.881 | 315.493
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 240.4311 | 20.9162 | 25.1043 | 1000.02 | 0.0052345 | 0.34759 | 304.711 | 309.144
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 235.4948 | 20.5925 | 25.0842 | 1000.01 | 0.0052264 | 0.34835 | 298.615 | 302.641
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 231.0793 | 20.0694 | 25.0708 | 1000 | 0.0052194 | 0.34858 | 292.575 | 295.993
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 225.8458 | 19.7668 | 25.0648 | 999.998 | 0.0052076 | 0.34785 | 286.506 | 289.538
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 221.5183 | 19.1962 | 25.054 | 1000 | 0.0051984 | 0.34783 | 280.458 | 282.835
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 216.9346 | 18.751 | 25.0395 | 999.998 | 0.0051929 | 0.34833 | 274.428 | 276.277
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 211.6931 | 18.5196 | 25.0276 | 1000.01 | 0.0051873 | 0.3486 | 268.391 | 269.928
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 208.0168 | 17.794 | 25.0063 | 999.996 | 0.0051812 | 0.34971 | 262.384 | 263.111
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 204.0497 | 17.123 | 24.9876 | 999.983 | 0.0051754 | 0.35069 | 256.273 | 256.244
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 198.3606 | 16.9743 | 24.9837 | 999.999 | 0.0051713 | 0.35047 | 250.102 | 249.843
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 193.4315 | 16.6304 | 24.9726 | 999.966 | 0.0051682 | 0.351 | 244.032 | 243.348
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 189.2317 | 16.2721 | 24.9529 | 999.962 | 0.0052031 | 0.35641 | 238.006 | 236.892
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 185.3284 | 16.0177 | 24.9528 | 999.971 | 0.0053015 | 0.36707 | 231.978 | 230.553
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 182.7681 | 15.2928 | 24.9742 | 999.99 | 0.0054309 | 0.37933 | 226.01 | 223.807
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 179.7737 | 14.5872 | 24.9923 | 999.982 | 0.0055425 | 0.39021 | 219.996 | 217.028
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 174.834 | 14.4078 | 25.0087 | 1000.02 | 0.0056129 | 0.39659 | 213.914 | 210.7
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 169.9365 | 14.0038 | 25.0293 | 1000.03 | 0.005634 | 0.39686 | 207.905 | 204.208
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 164.823 | 13.5645 | 25.0487 | 999.974 | 0.0056229 | 0.39333 | 201.931 | 197.711
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 159.5874 | 13.1139 | 25.0626 | 1000 | 0.005594 | 0.3884 | 195.911 | 191.151
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 155.1427 | 12.3582 | 25.0721 | 1000.01 | 0.0055609 | 0.38364 | 189.856 | 184.247
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 149.6821 | 12.0255 | 25.0797 | 1000.01 | 0.0055356 | 0.37983 | 183.772 | 177.744
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 143.7341 | 11.8946 | 25.0953 | 999.959 | 0.005515 | 0.3759 | 177.735 | 171.491
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 138.9818 | 11.3514 | 25.1143 | 1000.01 | 0.0054944 | 0.37173 | 171.815 | 164.941
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 134.3633 | 10.8185 | 25.1211 | 999.971 | 0.0054777 | 0.36912 | 165.9 | 158.41
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 129.7496 | 10.2609 | 25.1311 | 1000 | 0.0054623 | 0.36646 | 159.908 | 151.775
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 125.4029 | 9.6156 | 25.1379 | 999.999 | 0.005449 | 0.36425 | 153.89 | 145.026
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 120.7986 | 9.0887 | 25.1407 | 999.986 | 0.0054399 | 0.36283 | 147.862 | 138.387
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 115.7976 | 8.6812 | 25.1477 | 999.984 | 0.0054287 | 0.36093 | 141.787 | 131.82
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 111.7386 | 7.9437 | 25.1529 | 999.94 | 0.0054162 | 0.35917 | 135.721 | 124.926
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 106.9328 | 7.4734 | 25.164 | 999.982 | 0.0054152 | 0.35782 | 129.609 | 118.266
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 101.0771 | 7.3892 | 25.1661 | 999.985 | 0.0054106 | 0.35707 | 123.472 | 111.961
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 96.3406 | 6.9563 | 25.1592 | 1000.29 | 0.0054057 | 0.35709 | 117.439 | 105.41
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 92.007 | 6.3959 | 25.1533 | 999.939 | 0.0054021 | 0.35715 | 111.454 | 98.7826
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 87.1257 | 6.0073 | 25.1529 | 999.973 | 0.0054006 | 0.35687 | 105.416 | 92.2767
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 83.124 | 5.2925 | 25.161 | 999.944 | 0.005396 | 0.35557 | 99.3934 | 85.4536
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 78.3914 | 4.835 | 25.1667 | 999.965 | 0.0053925 | 0.35456 | 93.3336 | 78.8505
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 73.1482 | 4.5768 | 25.1671 | 999.972 | 0.0053922 | 0.35433 | 87.2857 | 72.4685
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 69.2822 | 3.8367 | 25.1714 | 999.933 | 0.005386 | 0.35333 | 81.2885 | 65.6433
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 65.4221 | 3.0762 | 25.1698 | 999.968 | 0.0053818 | 0.35294 | 75.2101 | 58.7205
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 60.1478 | 2.8006 | 25.1604 | 999.966 | 0.0053823 | 0.35359 | 69.0514 | 52.2052
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 55.2941 | 2.3922 | 25.1604 | 999.93 | 0.0053796 | 0.35314 | 62.9689 | 45.6304
Populus deltoides | LI1C |  | RACiR 500 to 0 2.5 min | 50.5315 | 1.9705 | 25.1654 | 999.962 | 0.0053761 | 0.35225 | 56.9457 | 39.1032
Populus deltoides | LI1C |  | RACiR 500 to 0 | 333.4336 | 31.474 | 24.977 | 1000.11 | 0.0064734 | 0.28089 | 454.564 | 479.763
Populus deltoides | LI1C |  | RACiR 500 to 0 | 332.1542 | 30.9851 | 24.9904 | 1000.1 | 0.0064684 | 0.28008 | 451.837 | 476.531
Populus deltoides | LI1C |  | RACiR 500 to 0 | 329.7887 | 30.873 | 24.9941 | 1000.11 | 0.0064745 | 0.28018 | 448.975 | 473.501
Populus deltoides | LI1C |  | RACiR 500 to 0 | 328.138 | 30.5424 | 24.986 | 1000.11 | 0.0064665 | 0.28018 | 446.083 | 470.255
Populus deltoides | LI1C |  | RACiR 500 to 0 | 326.2671 | 30.2982 | 24.984 | 999.79 | 0.0064715 | 0.28051 | 443.154 | 467.03
Populus deltoides | LI1C |  | RACiR 500 to 0 | 323.9261 | 30.128 | 24.9865 | 1000.1 | 0.0064715 | 0.28034 | 440.217 | 463.883
Populus deltoides | LI1C |  | RACiR 500 to 0 | 322.3279 | 29.7338 | 24.9857 | 1000.09 | 0.0064664 | 0.28013 | 437.236 | 460.484
Populus deltoides | LI1C |  | RACiR 500 to 0 | 320.3005 | 29.4764 | 24.9877 | 1000.09 | 0.0064711 | 0.28025 | 434.189 | 457.123
Populus deltoides | LI1C |  | RACiR 500 to 0 | 317.4228 | 29.374 | 24.9928 | 999.787 | 0.0064644 | 0.27973 | 431.081 | 453.883
Populus deltoides | LI1C |  | RACiR 500 to 0 | 315.0596 | 29.1925 | 24.9958 | 999.786 | 0.0064677 | 0.27975 | 428.003 | 450.575
Populus deltoides | LI1C |  | RACiR 500 to 0 | 312.4426 | 29.1056 | 24.9934 | 1000.08 | 0.0064693 | 0.27988 | 424.96 | 447.405
Populus deltoides | LI1C |  | RACiR 500 to 0 | 309.942 | 28.9922 | 24.9885 | 1000.09 | 0.0064682 | 0.28003 | 421.927 | 444.223
Populus deltoides | LI1C |  | RACiR 500 to 0 | 308.0623 | 28.6628 | 24.9911 | 999.789 | 0.0064672 | 0.27986 | 418.879 | 440.804
Populus deltoides | LI1C |  | RACiR 500 to 0 | 305.7013 | 28.4383 | 24.9989 | 1000.09 | 0.0064664 | 0.27949 | 415.8 | 437.471
Populus deltoides | LI1C |  | RACiR 500 to 0 | 303.2212 | 28.322 | 24.9959 | 1000.08 | 0.0064716 | 0.27982 | 412.716 | 434.214
Populus deltoides | LI1C |  | RACiR 500 to 0 | 301.2754 | 27.9897 | 24.9943 | 1000.08 | 0.0064674 | 0.27967 | 409.575 | 430.709
Populus deltoides | LI1C |  | RACiR 500 to 0 | 298.9258 | 27.7907 | 24.9969 | 999.771 | 0.0064695 | 0.27972 | 406.435 | 427.329
Populus deltoides | LI1C |  | RACiR 500 to 0 | 296.2934 | 27.6948 | 24.9913 | 1000.09 | 0.0064708 | 0.28003 | 403.277 | 424.027
Populus deltoides | LI1C |  | RACiR 500 to 0 | 293.942 | 27.4326 | 24.999 | 1000.09 | 0.0064673 | 0.27951 | 400.118 | 420.573
Populus deltoides | LI1C |  | RACiR 500 to 0 | 291.0506 | 27.3698 | 25.0056 | 999.782 | 0.0064702 | 0.27932 | 397.008 | 417.352
Populus deltoides | LI1C |  | RACiR 500 to 0 | 287.9995 | 27.3848 | 25.0025 | 1000.09 | 0.0064675 | 0.27931 | 393.939 | 414.266
Populus deltoides | LI1C |  | RACiR 500 to 0 | 286.201 | 27.0724 | 24.9993 | 1000.09 | 0.0064686 | 0.2795 | 390.898 | 410.869
Populus deltoides | LI1C |  | RACiR 500 to 0 | 284.3137 | 26.7728 | 24.9969 | 1000.09 | 0.0064704 | 0.27973 | 387.798 | 407.431
Populus deltoides | LI1C |  | RACiR 500 to 0 | 281.3299 | 26.7153 | 25.0058 | 1000.09 | 0.0064698 | 0.2793 | 384.695 | 404.228
Populus deltoides | LI1C |  | RACiR 500 to 0 | 278.8704 | 26.5136 | 25.0102 | 1000.08 | 0.0064685 | 0.27895 | 381.574 | 400.862
Populus deltoides | LI1C |  | RACiR 500 to 0 | 276.462 | 26.3792 | 25.0002 | 1000.08 | 0.0064672 | 0.27931 | 378.49 | 397.61
Populus deltoides | LI1C |  | RACiR 500 to 0 | 274.0352 | 26.2051 | 25.0037 | 999.779 | 0.0064681 | 0.27914 | 375.437 | 394.335
Populus deltoides | LI1C |  | RACiR 500 to 0 | 272.0983 | 25.8471 | 25.0119 | 1000.08 | 0.0064651 | 0.27867 | 372.327 | 390.836
Populus deltoides | LI1C |  | RACiR 500 to 0 | 269.7604 | 25.6883 | 25.0043 | 1000.08 | 0.0064687 | 0.27922 | 369.163 | 387.464
Populus deltoides | LI1C |  | RACiR 500 to 0 | 267.018 | 25.6184 | 24.9969 | 1000.09 | 0.0064652 | 0.27938 | 366.041 | 384.232
Populus deltoides | LI1C |  | RACiR 500 to 0 | 265.0678 | 25.3186 | 24.9972 | 1000.09 | 0.006466 | 0.27938 | 362.961 | 380.815
Populus deltoides | LI1C |  | RACiR 500 to 0 | 262.8822 | 25.0593 | 25.0042 | 1000.09 | 0.0064689 | 0.27915 | 359.867 | 377.416
Populus deltoides | LI1C |  | RACiR 500 to 0 | 259.8234 | 25.0306 | 25.0078 | 999.785 | 0.0064653 | 0.27879 | 356.757 | 374.24
Populus deltoides | LI1C |  | RACiR 500 to 0 | 257.7128 | 24.7804 | 25.0076 | 999.788 | 0.0064655 | 0.2788 | 353.684 | 370.878
Populus deltoides | LI1C |  | RACiR 500 to 0 | 255.2986 | 24.6404 | 25.0037 | 1000.09 | 0.0064682 | 0.27909 | 350.609 | 367.622
Populus deltoides | LI1C |  | RACiR 500 to 0 | 252.2966 | 24.6579 | 24.9981 | 1000.08 | 0.0064665 | 0.27919 | 347.564 | 364.555
Populus deltoides | LI1C |  | RACiR 500 to 0 | 250.6412 | 24.2935 | 24.9966 | 999.782 | 0.0064638 | 0.27909 | 344.579 | 361.166
Populus deltoides | LI1C |  | RACiR 500 to 0 | 249.0323 | 23.9192 | 24.9958 | 1000.09 | 0.006466 | 0.27924 | 341.528 | 357.699
Populus deltoides | LI1C |  | RACiR 500 to 0 | 246.1095 | 23.8776 | 24.9964 | 1000.09 | 0.0064639 | 0.27911 | 338.426 | 354.516
Populus deltoides | LI1C |  | RACiR 500 to 0 | 243.7662 | 23.678 | 25.0004 | 1000.09 | 0.0064636 | 0.27889 | 335.378 | 351.226
Populus deltoides | LI1C |  | RACiR 500 to 0 | 241.1479 | 23.4935 | 25.0042 | 999.778 | 0.006444 | 0.27789 | 332.34 | 347.966
Populus deltoides | LI1C |  | RACiR 500 to 0 | 238.5507 | 23.3368 | 25.0067 | 1000.09 | 0.00643 | 0.27727 | 329.3 | 344.715
Populus deltoides | LI1C |  | RACiR 500 to 0 | 236.4662 | 23.0925 | 25.0091 | 1000.09 | 0.0064298 | 0.27732 | 326.259 | 341.391
Populus deltoides | LI1C |  | RACiR 500 to 0 | 234.6345 | 22.8132 | 25.0075 | 1000.08 | 0.0064354 | 0.27783 | 323.208 | 338.03
Populus deltoides | LI1C |  | RACiR 500 to 0 | 232.3126 | 22.6921 | 25.0032 | 1000.08 | 0.0064453 | 0.27853 | 320.175 | 334.831
Populus deltoides | LI1C |  | RACiR 500 to 0 | 230.7454 | 22.3007 | 25.0013 | 1000.07 | 0.0064434 | 0.27855 | 317.141 | 331.37
Populus deltoides | LI1C |  | RACiR 500 to 0 | 228.8586 | 21.9686 | 25.0055 | 1000.07 | 0.0064433 | 0.27839 | 314.053 | 327.908
Populus deltoides | LI1C |  | RACiR 500 to 0 | 226.1733 | 21.8634 | 25.0083 | 1000.07 | 0.0064456 | 0.27842 | 310.917 | 324.625
Populus deltoides | LI1C |  | RACiR 500 to 0 | 224.0927 | 21.6014 | 25.0063 | 1000.08 | 0.0064461 | 0.27855 | 307.797 | 321.202
Populus deltoides | LI1C |  | RACiR 500 to 0 | 221.4089 | 21.5103 | 25.0075 | 1000.08 | 0.0064485 | 0.27861 | 304.702 | 317.978
Populus deltoides | LI1C |  | RACiR 500 to 0 | 218.5871 | 21.4646 | 25.0136 | 1000.09 | 0.006453 | 0.27857 | 301.667 | 314.858
Populus deltoides | LI1C |  | RACiR 500 to 0 | 216.771 | 21.1247 | 25.0193 | 1000.07 | 0.0064536 | 0.27828 | 298.66 | 311.47
Populus deltoides | LI1C |  | RACiR 500 to 0 | 214.5341 | 20.9077 | 25.0142 | 1000.07 | 0.0064472 | 0.27816 | 295.61 | 308.169
Populus deltoides | LI1C |  | RACiR 500 to 0 | 212.1737 | 20.7682 | 25.0084 | 1000.06 | 0.0064524 | 0.27865 | 292.549 | 304.923
Populus deltoides | LI1C |  | RACiR 500 to 0 | 210.4769 | 20.3952 | 25.0119 | 1000.06 | 0.0064493 | 0.27835 | 289.54 | 301.506
Populus deltoides | LI1C |  | RACiR 500 to 0 | 208.0358 | 20.2504 | 25.0161 | 1000.07 | 0.0064525 | 0.27831 | 286.53 | 298.308
Populus deltoides | LI1C |  | RACiR 500 to 0 | 205.4764 | 20.1528 | 25.0127 | 1000.06 | 0.0064533 | 0.27848 | 283.509 | 295.147
Populus deltoides | LI1C |  | RACiR 500 to 0 | 203.7038 | 19.8158 | 25.0057 | 1000.07 | 0.006443 | 0.27835 | 280.497 | 291.769
Populus deltoides | LI1C |  | RACiR 500 to 0 | 201.6673 | 19.5729 | 25.0106 | 1000.07 | 0.0064537 | 0.27861 | 277.469 | 288.449
Populus deltoides | LI1C |  | RACiR 500 to 0 | 199.2482 | 19.3828 | 25.0158 | 999.767 | 0.0064519 | 0.27825 | 274.4 | 285.152
Populus deltoides | LI1C |  | RACiR 500 to 0 | 197.69 | 18.9698 | 25.0196 | 1000.07 | 0.0064524 | 0.27815 | 271.331 | 281.631
Populus deltoides | LI1C |  | RACiR 500 to 0 | 195.5413 | 18.7301 | 25.0208 | 1000.05 | 0.0064549 | 0.27819 | 268.25 | 278.265
Populus deltoides | LI1C |  | RACiR 500 to 0 | 192.4744 | 18.7308 | 25.0182 | 1000.05 | 0.00645 | 0.27802 | 265.154 | 275.134
Populus deltoides | LI1C |  | RACiR 500 to 0 | 190.5425 | 18.4326 | 25.0196 | 1000.06 | 0.0064501 | 0.278 | 262.098 | 271.739
Populus deltoides | LI1C |  | RACiR 500 to 0 | 188.501 | 18.1747 | 25.0188 | 1000.06 | 0.006451 | 0.2781 | 259.045 | 268.388
Populus deltoides | LI1C |  | RACiR 500 to 0 | 186.0536 | 18.0652 | 25.0105 | 999.744 | 0.0064552 | 0.27865 | 256.004 | 265.197
Populus deltoides | LI1C |  | RACiR 500 to 0 | 184.3315 | 17.7161 | 25.0111 | 1000.06 | 0.006454 | 0.27855 | 253 | 261.806
Populus deltoides | LI1C |  | RACiR 500 to 0 | 182.3815 | 17.4122 | 25.0156 | 999.759 | 0.0064542 | 0.27831 | 249.959 | 258.419
Populus deltoides | LI1C |  | RACiR 500 to 0 | 179.9072 | 17.2886 | 25.0119 | 1000.06 | 0.0064554 | 0.27854 | 246.921 | 255.219
Populus deltoides | LI1C |  | RACiR 500 to 0 | 177.8592 | 17.0397 | 25.0115 | 1000.06 | 0.0064568 | 0.27865 | 243.896 | 251.906
Populus deltoides | LI1C |  | RACiR 500 to 0 | 175.6106 | 16.8457 | 25.0116 | 1000.05 | 0.0064579 | 0.27868 | 240.885 | 248.66
Populus deltoides | LI1C |  | RACiR 500 to 0 | 173.2336 | 16.6861 | 25.0121 | 1000.06 | 0.0064543 | 0.27851 | 237.909 | 245.489
Populus deltoides | LI1C |  | RACiR 500 to 0 | 171.6345 | 16.325 | 25.0121 | 1000.06 | 0.0064594 | 0.27876 | 234.907 | 242.084
Populus deltoides | LI1C |  | RACiR 500 to 0 | 169.6242 | 16.059 | 25.009 | 1000.05 | 0.0064595 | 0.27889 | 231.855 | 238.727
Populus deltoides | LI1C |  | RACiR 500 to 0 | 167.1985 | 15.9073 | 25.0093 | 1000.04 | 0.0064611 | 0.27896 | 228.808 | 235.487
Populus deltoides | LI1C |  | RACiR 500 to 0 | 165.6733 | 15.5205 | 25.0049 | 1000.05 | 0.0064642 | 0.2793 | 225.772 | 232.023
Populus deltoides | LI1C |  | RACiR 500 to 0 | 163.4824 | 15.2837 | 25.0017 | 999.741 | 0.0064584 | 0.27919 | 222.692 | 228.67
Populus deltoides | LI1C |  | RACiR 500 to 0 | 160.8841 | 15.1716 | 25.0059 | 1000.05 | 0.0064643 | 0.27927 | 219.613 | 225.436
Populus deltoides | LI1C |  | RACiR 500 to 0 | 159.0721 | 14.8347 | 25.0027 | 1000.05 | 0.0064604 | 0.2792 | 216.55 | 221.996
Populus deltoides | LI1C |  | RACiR 500 to 0 | 156.939 | 14.6019 | 24.9992 | 1000.05 | 0.0064608 | 0.27939 | 213.485 | 218.658
Populus deltoides | LI1C |  | RACiR 500 to 0 | 154.1233 | 14.5743 | 24.9984 | 1000.05 | 0.0064649 | 0.27959 | 210.465 | 215.571
Populus deltoides | LI1C |  | RACiR 500 to 0 | 152.1942 | 14.2728 | 25.0049 | 1000.04 | 0.0064622 | 0.27922 | 207.473 | 212.24
Populus deltoides | LI1C |  | RACiR 500 to 0 | 150.5362 | 13.8734 | 25.0058 | 1000.05 | 0.0064651 | 0.27935 | 204.305 | 208.629
Populus deltoides | LI1C |  | RACiR 500 to 0 | 148.1306 | 13.7175 | 25.0012 | 1000.05 | 0.0064636 | 0.27946 | 201.257 | 205.385
Populus deltoides | LI1C |  | RACiR 500 to 0 | 145.8295 | 13.5196 | 25.0052 | 1000.05 | 0.006465 | 0.27931 | 198.214 | 202.105
Populus deltoides | LI1C |  | RACiR 500 to 0 | 143.3198 | 13.4196 | 25.0016 | 1000.04 | 0.0064701 | 0.27969 | 195.215 | 198.968
Populus deltoides | LI1C |  | RACiR 500 to 0 | 141.2357 | 13.1926 | 25.0028 | 999.735 | 0.00647 | 0.27966 | 192.267 | 195.755
Populus deltoides | LI1C |  | RACiR 500 to 0 | 140.1059 | 12.6916 | 25.0036 | 1000.04 | 0.0064718 | 0.27974 | 189.286 | 192.232
Populus deltoides | LI1C |  | RACiR 500 to 0 | 138.2518 | 12.3751 | 25.003 | 1000.04 | 0.0064711 | 0.27975 | 186.242 | 188.831
Populus deltoides | LI1C |  | RACiR 500 to 0 | 135.8178 | 12.2116 | 25.0036 | 1000.03 | 0.0064714 | 0.27971 | 183.165 | 185.55
Populus deltoides | LI1C |  | RACiR 500 to 0 | 134.0194 | 11.8762 | 25.0017 | 1000.03 | 0.006471 | 0.27978 | 180.101 | 182.11
Populus deltoides | LI1C |  | RACiR 500 to 0 | 131.8536 | 11.6384 | 25.0049 | 1000.03 | 0.0064728 | 0.27971 | 177.036 | 178.767
Populus deltoides | LI1C |  | RACiR 500 to 0 | 129.3782 | 11.4991 | 25.003 | 1000.04 | 0.0064743 | 0.2799 | 173.97 | 175.522
Populus deltoides | LI1C |  | RACiR 500 to 0 | 127.2711 | 11.2606 | 25.0012 | 1000.04 | 0.0064745 | 0.27999 | 170.937 | 172.211
Populus deltoides | LI1C |  | RACiR 500 to 0 | 124.9227 | 11.0867 | 25.0058 | 1000.04 | 0.0064772 | 0.27988 | 167.924 | 168.985
Populus deltoides | LI1C |  | RACiR 500 to 0 | 122.744 | 10.8732 | 25.0046 | 999.725 | 0.0064756 | 0.27986 | 164.924 | 165.732
Populus deltoides | LI1C |  | RACiR 500 to 0 | 121.0424 | 10.5265 | 25.0022 | 1000.05 | 0.0064769 | 0.28001 | 161.908 | 162.329
Populus deltoides | LI1C |  | RACiR 500 to 0 | 118.8966 | 10.2914 | 25.0052 | 1000.04 | 0.0064769 | 0.27993 | 158.874 | 159.02
Populus deltoides | LI1C |  | RACiR 500 to 0 | 116.5733 | 10.1215 | 25.002 | 999.735 | 0.0064804 | 0.28026 | 155.837 | 155.774
Populus deltoides | LI1C |  | RACiR 500 to 0 | 114.8772 | 9.769 | 24.9986 | 1000.03 | 0.0064794 | 0.28033 | 152.819 | 152.363
Populus deltoides | LI1C |  | RACiR 500 to 0 | 112.456 | 9.6144 | 25.0019 | 1000.03 | 0.00648 | 0.2802 | 149.801 | 149.151
Populus deltoides | LI1C |  | RACiR 500 to 0 | 110.0061 | 9.4833 | 25.003 | 1000.04 | 0.0064851 | 0.28038 | 146.799 | 145.98
Populus deltoides | LI1C |  | RACiR 500 to 0 | 108.5073 | 9.0769 | 25.0029 | 1000.03 | 0.0064807 | 0.28018 | 143.822 | 142.555
Populus deltoides | LI1C |  | RACiR 500 to 0 | 106.0552 | 8.9532 | 24.9986 | 1000.03 | 0.0064851 | 0.28057 | 140.819 | 139.391
Populus deltoides | LI1C |  | RACiR 500 to 0 | 104.3351 | 8.6255 | 24.9961 | 1000.04 | 0.0064873 | 0.28079 | 137.851 | 136.056
Populus deltoides | LI1C |  | RACiR 500 to 0 | 103.5133 | 7.9993 | 25.0074 | 999.736 | 0.0064863 | 0.2803 | 134.809 | 132.344
Populus deltoides | LI1C |  | RACiR 500 to 0 | 101.0773 | 7.8207 | 25.0145 | 1000.03 | 0.006492 | 0.2802 | 131.685 | 129.003
Populus deltoides | LI1C |  | RACiR 500 to 0 | 98.4604 | 7.706 | 25.0101 | 1000.03 | 0.0064865 | 0.28015 | 128.596 | 125.756
Populus deltoides | LI1C |  | RACiR 500 to 0 | 96.6446 | 7.399 | 24.9987 | 1000.04 | 0.0064934 | 0.28104 | 125.538 | 122.353
Populus deltoides | LI1C |  | RACiR 500 to 0 | 94.3147 | 7.2182 | 24.996 | 999.727 | 0.0064975 | 0.28134 | 122.475 | 119.071
Populus deltoides | LI1C |  | RACiR 500 to 0 | 91.8714 | 7.0542 | 25.0063 | 1000.03 | 0.006493 | 0.28067 | 119.449 | 115.838
Populus deltoides | LI1C |  | RACiR 500 to 0 | 90.0701 | 6.7418 | 25.0048 | 1000.04 | 0.0064981 | 0.28099 | 116.451 | 112.49
Populus deltoides | LI1C |  | RACiR 500 to 0 | 87.8659 | 6.5332 | 25.0043 | 1000.03 | 0.0064998 | 0.28109 | 113.436 | 109.227
Populus deltoides | LI1C |  | RACiR 500 to 0 | 85.5691 | 6.3475 | 25.0083 | 1000.03 | 0.0064993 | 0.28089 | 110.435 | 106
Populus deltoides | LI1C |  | RACiR 500 to 0 | 84.0261 | 5.9597 | 25.0068 | 1000.04 | 0.0065005 | 0.28104 | 107.446 | 102.584
Populus deltoides | LI1C |  | RACiR 500 to 0 | 81.8033 | 5.7465 | 25.01 | 1000.03 | 0.0064988 | 0.28081 | 104.422 | 99.3056
Populus deltoides | LI1C |  | RACiR 500 to 0 | 79.4176 | 5.594 | 25.012 | 1000.03 | 0.0065036 | 0.28093 | 101.423 | 96.1177
Populus deltoides | LI1C |  | RACiR 500 to 0 | 77.6652 | 5.271 | 25.0062 | 1000.04 | 0.0065017 | 0.28109 | 98.4542 | 92.7843
Populus deltoides | LI1C |  | RACiR 500 to 0 | 75.8922 | 4.9495 | 25.0051 | 1000.03 | 0.0065031 | 0.28121 | 95.4744 | 89.4434
Populus deltoides | LI1C |  | RACiR 500 to 0 | 73.8072 | 4.7036 | 25.0038 | 1000.04 | 0.0064964 | 0.28096 | 92.4708 | 86.1546
Populus deltoides | LI1C |  | RACiR 500 to 0 | 72.0213 | 4.3467 | 25.0022 | 999.723 | 0.0064692 | 0.27987 | 89.4182 | 82.6986
Populus deltoides | LI1C |  | RACiR 500 to 0 | 69.9908 | 4.061 | 25.0126 | 1000.03 | 0.0064562 | 0.27902 | 86.354 | 79.3111
Populus deltoides | LI1C |  | RACiR 500 to 0 | 67.7338 | 3.8625 | 25.0098 | 1000.04 | 0.0064594 | 0.27951 | 83.2977 | 76.0138
Populus deltoides | LI1C |  | RACiR 500 to 0 | 66.0189 | 3.519 | 25.0077 | 1000.03 | 0.0064725 | 0.28042 | 80.256 | 72.589
Populus deltoides | LI1C |  | RACiR 500 to 0 | 63.8223 | 3.2967 | 25.0166 | 1000.04 | 0.0064806 | 0.28046 | 77.2059 | 69.2795
Populus deltoides | LI1C |  | RACiR 500 to 0 | 61.3399 | 3.1649 | 25.0121 | 1000.04 | 0.0064814 | 0.28078 | 74.1759 | 66.0767
Populus deltoides | LI1C |  | RACiR 500 to 0 | 59.3027 | 2.9167 | 25.0097 | 1000.03 | 0.0064864 | 0.28117 | 71.1812 | 62.7952
Populus deltoides | LI1C |  | RACiR 500 to 0 | 57.2952 | 2.6586 | 25.0115 | 1000.03 | 0.0064876 | 0.28113 | 68.2 | 59.5158
Populus deltoides | LI1C |  | RACiR 500 to 0 | 55.4685 | 2.3411 | 25.0118 | 1000.04 | 0.0064882 | 0.28116 | 65.1872 | 56.1457
Populus deltoides | LI1C |  | RACiR 500 to 0 | 53.7956 | 1.9721 | 25.0156 | 1000.03 | 0.0064884 | 0.28097 | 62.1533 | 52.7033
Populus deltoides | LI1C |  | RACiR 500 to 0 | 51.4248 | 1.8059 | 25.0112 | 1000.03 | 0.0064893 | 0.28126 | 59.1209 | 49.4623
Populus deltoides | LI1C |  | RACiR 500 to 0 | 49.0972 | 1.6266 | 25.0125 | 999.717 | 0.0064939 | 0.28146 | 56.0896 | 46.2144
Populus deltoides | LI1C |  | RACiR 500 to 0 | 47.4586 | 1.2544 | 25.0134 | 1000.03 | 0.006496 | 0.28151 | 53.0744 | 42.7875
Populus deltoides | LI1C |  | RACiR 500 to 0 | 45.4925 | 0.97971 | 25.0055 | 1000.02 | 0.006493 | 0.28173 | 50.0713 | 39.4655
Populus deltoides | LI1C |  | RACiR 500 to 0 | 43.2599 | 0.77611 | 25.0071 | 1000.02 | 0.006497 | 0.28181 | 47.055 | 36.212
Populus deltoides | LI1C |  | RACiR 500 to 0 | 41.7634 | 0.36151 | 25.0105 | 1000.03 | 0.006494 | 0.28149 | 44.0374 | 32.7329
Populus deltoides | LI1C |  | RACiR 500 to 0 | 39.6209 | 0.11935 | 25.0142 | 999.724 | 0.0064984 | 0.28155 | 40.9749 | 29.3925
Populus deltoides | LI1C |  | RACiR 500 to 0 | 37.0636 | -0.0016598 | 25.0145 | 1000.04 | 0.0065021 | 0.28173 | 37.923 | 26.185
Populus deltoides | LI1C |  | RACiR 500 to 0 | 35.2073 | -0.31127 | 25.0033 | 1000.04 | 0.0064928 | 0.28182 | 34.9105 | 22.8042
Populus deltoides | LI1C |  | RACiR 500 to 0 | 33.0589 | -0.53068 | 24.9985 | 1000.04 | 0.0065001 | 0.28234 | 31.9274 | 19.5771
Populus deltoides | LI1C |  | RACiR 500 to 0 | 30.7726 | -0.709 | 25.0077 | 1000.04 | 0.0064972 | 0.28169 | 28.9443 | 16.3799
Populus deltoides | LI1C |  | RACiR 500 to 0 | 29.1486 | -1.0768 | 25.0107 | 1000.03 | 0.0064908 | 0.28125 | 25.9557 | 12.9702
Populus deltoides | LI1C |  | RACiR 500 to 0 | 29.2095 | -1.9393 | 25.0106 | 1000.04 | 0.0064986 | 0.28161 | 22.9246 | 9.0422
Populus deltoides | LI1C |  | RACiR 500 to 0 | 25.8457 | -1.8845 | 25.01 | 1000.04 | 0.0064962 | 0.28148 | 19.676 | 5.806
Populus deltoides | LI1C |  | RACiR 300 to 800 | 221.8606 | 20.0239 | 24.98 | 999.835 | 0.0070343 | 0.30634 | 293.139 | 323.091
Populus deltoides | LI1C |  | RACiR 300 to 800 | 223.6305 | 20.2955 | 24.9764 | 1000.14 | 0.007031 | 0.3064 | 295.83 | 326.112
Populus deltoides | LI1C |  | RACiR 300 to 800 | 224.5556 | 20.8511 | 24.9797 | 1000.13 | 0.007034 | 0.30635 | 298.624 | 329.485
Populus deltoides | LI1C |  | RACiR 300 to 800 | 226.0276 | 21.2468 | 24.9824 | 999.826 | 0.0070282 | 0.3059 | 301.546 | 332.847
Populus deltoides | LI1C |  | RACiR 300 to 800 | 228.4513 | 21.4256 | 24.9787 | 999.84 | 0.0070308 | 0.30626 | 304.528 | 336.044
Populus deltoides | LI1C |  | RACiR 300 to 800 | 230.3496 | 21.7319 | 24.9834 | 999.831 | 0.0070341 | 0.30615 | 307.514 | 339.369
Populus deltoides | LI1C |  | RACiR 300 to 800 | 232.4664 | 21.9562 | 24.9851 | 999.838 | 0.0070276 | 0.30581 | 310.503 | 342.657
Populus deltoides | LI1C |  | RACiR 300 to 800 | 234.7697 | 22.207 | 24.9808 | 1000.15 | 0.0070387 | 0.30662 | 313.489 | 345.904
Populus deltoides | LI1C |  | RACiR 300 to 800 | 236.6451 | 22.5193 | 24.9797 | 1000.14 | 0.0070353 | 0.30651 | 316.463 | 349.244
Populus deltoides | LI1C |  | RACiR 300 to 800 | 238.3663 | 22.8904 | 24.9833 | 1000.15 | 0.0070363 | 0.30636 | 319.49 | 352.684
Populus deltoides | LI1C |  | RACiR 300 to 800 | 240.9303 | 23.067 | 24.9795 | 999.839 | 0.0070395 | 0.3067 | 322.608 | 356.01
Populus deltoides | LI1C |  | RACiR 300 to 800 | 242.8862 | 23.3934 | 24.9749 | 999.845 | 0.0070327 | 0.30664 | 325.696 | 359.491
Populus deltoides | LI1C |  | RACiR 300 to 800 | 245.0211 | 23.7061 | 24.9723 | 999.835 | 0.0070395 | 0.3071 | 328.792 | 362.917
Populus deltoides | LI1C |  | RACiR 300 to 800 | 247.9722 | 23.7425 | 24.9684 | 1000.16 | 0.0070361 | 0.30717 | 331.909 | 366.131
Populus deltoides | LI1C |  | RACiR 300 to 800 | 249.685 | 24.1403 | 24.9732 | 1000.15 | 0.0070387 | 0.30707 | 335.001 | 369.651
Populus deltoides | LI1C |  | RACiR 300 to 800 | 252.0391 | 24.3565 | 24.976 | 999.854 | 0.0070436 | 0.30716 | 338.102 | 372.997
Populus deltoides | LI1C |  | RACiR 300 to 800 | 255.0997 | 24.3649 | 24.9701 | 1000.16 | 0.0070385 | 0.30731 | 341.212 | 376.195
Populus deltoides | LI1C |  | RACiR 300 to 800 | 257.0917 | 24.7038 | 24.9723 | 1000.15 | 0.0070468 | 0.30759 | 344.293 | 379.624
Populus deltoides | LI1C |  | RACiR 300 to 800 | 259.2936 | 24.9525 | 24.9736 | 1000.16 | 0.007046 | 0.30746 | 347.401 | 383.028
Populus deltoides | LI1C |  | RACiR 300 to 800 | 261.9021 | 25.091 | 24.9752 | 999.854 | 0.0070487 | 0.30761 | 350.485 | 386.299
Populus deltoides | LI1C |  | RACiR 300 to 800 | 263.9949 | 25.3837 | 24.9759 | 999.862 | 0.0070513 | 0.30772 | 353.56 | 389.693
Populus deltoides | LI1C |  | RACiR 300 to 800 | 266.824 | 25.4223 | 24.9735 | 1000.17 | 0.0070432 | 0.30749 | 356.641 | 392.892
Populus deltoides | LI1C |  | RACiR 300 to 800 | 269.5486 | 25.5297 | 24.9799 | 999.867 | 0.0070543 | 0.30775 | 359.717 | 396.089
Populus deltoides | LI1C |  | RACiR 300 to 800 | 271.0716 | 25.9657 | 24.9809 | 1000.17 | 0.0070506 | 0.30749 | 362.784 | 399.639
Populus deltoides | LI1C |  | RACiR 300 to 800 | 273.264 | 26.2465 | 24.9754 | 1000.17 | 0.0070501 | 0.30779 | 365.858 | 403.047
Populus deltoides | LI1C |  | RACiR 300 to 800 | 276.0065 | 26.3756 | 24.9767 | 999.857 | 0.00706 | 0.30827 | 368.952 | 406.289
Populus deltoides | LI1C |  | RACiR 300 to 800 | 278.0787 | 26.6634 | 24.9759 | 999.863 | 0.0070552 | 0.30815 | 372.049 | 409.753
Populus deltoides | LI1C |  | RACiR 300 to 800 | 280.3431 | 26.9385 | 24.9754 | 999.875 | 0.0070615 | 0.30848 | 375.173 | 413.168
Populus deltoides | LI1C |  | RACiR 300 to 800 | 282.979 | 27.0793 | 24.9755 | 1000.17 | 0.0070599 | 0.30837 | 378.364 | 416.545
Populus deltoides | LI1C |  | RACiR 300 to 800 | 285.3166 | 27.3304 | 24.9716 | 999.861 | 0.0070588 | 0.30856 | 381.517 | 419.998
Populus deltoides | LI1C |  | RACiR 300 to 800 | 288.025 | 27.4611 | 24.9742 | 999.861 | 0.0070659 | 0.30881 | 384.647 | 423.284
Populus deltoides | LI1C |  | RACiR 300 to 800 | 291.1508 | 27.5318 | 24.9751 | 1000.16 | 0.0070904 | 0.3099 | 387.759 | 426.557
Populus deltoides | LI1C |  | RACiR 300 to 800 | 293.7296 | 27.8747 | 24.976 | 1000.17 | 0.0071481 | 0.31223 | 390.841 | 430.021
Populus deltoides | LI1C |  | RACiR 300 to 800 | 296.223 | 28.0585 | 24.9786 | 999.867 | 0.0071665 | 0.31242 | 393.95 | 433.363
Populus deltoides | LI1C |  | RACiR 300 to 800 | 298.7486 | 28.1463 | 24.9724 | 999.867 | 0.0071514 | 0.31159 | 397.066 | 436.597
Populus deltoides | LI1C |  | RACiR 300 to 800 | 300.952 | 28.3507 | 24.9702 | 999.879 | 0.0071416 | 0.31096 | 400.173 | 439.938
Populus deltoides | LI1C |  | RACiR 300 to 800 | 303.4097 | 28.4371 | 24.9757 | 999.877 | 0.0071275 | 0.31 | 403.265 | 443.203
Populus deltoides | LI1C |  | RACiR 300 to 800 | 306.154 | 28.5748 | 24.9681 | 999.881 | 0.0071337 | 0.31062 | 406.328 | 446.4
Populus deltoides | LI1C |  | RACiR 300 to 800 | 307.8961 | 28.9402 | 24.9653 | 999.869 | 0.0071279 | 0.31038 | 409.375 | 449.867
Populus deltoides | LI1C |  | RACiR 300 to 800 | 309.8622 | 29.2725 | 24.9659 | 999.867 | 0.0071276 | 0.31033 | 412.486 | 453.361
Populus deltoides | LI1C |  | RACiR 300 to 800 | 312.6946 | 29.3877 | 24.9599 | 1000.18 | 0.0071304 | 0.3107 | 415.637 | 456.641
Populus deltoides | LI1C |  | RACiR 300 to 800 | 314.6911 | 29.685 | 24.9597 | 999.869 | 0.0071217 | 0.31037 | 418.749 | 460.146
Populus deltoides | LI1C |  | RACiR 300 to 800 | 317.0226 | 29.9622 | 24.9622 | 999.867 | 0.0071331 | 0.31082 | 421.898 | 463.575
Populus deltoides | LI1C |  | RACiR 300 to 800 | 320.072 | 29.9823 | 24.9625 | 1000.18 | 0.0071344 | 0.31085 | 425.077 | 466.81
Populus deltoides | LI1C |  | RACiR 300 to 800 | 322.4362 | 30.2123 | 24.9599 | 999.887 | 0.0071331 | 0.31105 | 428.173 | 470.222
Populus deltoides | LI1C |  | RACiR 300 to 800 | 325.0846 | 30.4093 | 24.9561 | 1000.18 | 0.007144 | 0.31177 | 431.288 | 473.533
Populus deltoides | LI1C |  | RACiR 300 to 800 | 328.0105 | 30.4018 | 24.9589 | 1000.19 | 0.0071358 | 0.31122 | 434.438 | 476.739
Populus deltoides | LI1C |  | RACiR 300 to 800 | 330.0092 | 30.7097 | 24.9627 | 999.885 | 0.0071362 | 0.3111 | 437.524 | 480.183
Populus deltoides | LI1C |  | RACiR 300 to 800 | 332.1409 | 31.0164 | 24.9659 | 999.877 | 0.0071451 | 0.31136 | 440.623 | 483.603
Populus deltoides | LI1C |  | RACiR 300 to 800 | 335.0084 | 31.0523 | 24.9698 | 999.884 | 0.0071428 | 0.31108 | 443.77 | 486.853
Populus deltoides | LI1C |  | RACiR 300 to 800 | 337.4981 | 31.276 | 24.9665 | 999.889 | 0.007149 | 0.31157 | 446.878 | 490.215
Populus deltoides | LI1C |  | RACiR 300 to 800 | 340.1802 | 31.3988 | 24.9656 | 999.895 | 0.0071491 | 0.31166 | 449.99 | 493.505
Populus deltoides | LI1C |  | RACiR 300 to 800 | 342.8211 | 31.4946 | 24.9733 | 999.894 | 0.0071508 | 0.31145 | 453.081 | 496.732
Populus deltoides | LI1C |  | RACiR 300 to 800 | 344.9593 | 31.7789 | 24.9775 | 999.889 | 0.0071551 | 0.3115 | 456.181 | 500.148
Populus deltoides | LI1C |  | RACiR 300 to 800 | 347.5699 | 31.9454 | 24.9742 | 1000.2 | 0.0071556 | 0.31176 | 459.299 | 503.489
Populus deltoides | LI1C |  | RACiR 300 to 800 | 350.462 | 32.0629 | 24.9678 | 1000.2 | 0.0071619 | 0.3124 | 462.412 | 506.754
Populus deltoides | LI1C |  | RACiR 300 to 800 | 352.1605 | 32.4836 | 24.9702 | 999.899 | 0.007161 | 0.31226 | 465.565 | 510.386
Populus deltoides | LI1C |  | RACiR 300 to 800 | 354.6781 | 32.6565 | 24.9789 | 999.896 | 0.0071667 | 0.31211 | 468.764 | 513.788
Populus deltoides | LI1C |  | RACiR 300 to 800 | 357.7347 | 32.695 | 24.9759 | 999.897 | 0.007166 | 0.31228 | 471.953 | 517.069
Populus deltoides | LI1C |  | RACiR 300 to 800 | 360.1058 | 32.9444 | 24.9789 | 999.898 | 0.0071695 | 0.31237 | 475.159 | 520.563
Populus deltoides | LI1C |  | RACiR 300 to 800 | 362.9016 | 33.0488 | 24.9842 | 1000.22 | 0.0071762 | 0.31246 | 478.332 | 523.869
Populus deltoides | LI1C |  | RACiR 300 to 800 | 366.1491 | 32.9726 | 24.9844 | 999.905 | 0.0071712 | 0.31233 | 481.448 | 526.992
Populus deltoides | LI1C |  | RACiR 300 to 800 | 368.4594 | 33.2509 | 24.9853 | 999.904 | 0.0071791 | 0.31271 | 484.582 | 530.419
Populus deltoides | LI1C |  | RACiR 300 to 800 | 371.0427 | 33.4536 | 24.98 | 999.917 | 0.0071821 | 0.31314 | 487.724 | 533.806
Populus deltoides | LI1C |  | RACiR 300 to 800 | 373.7644 | 33.5482 | 24.9851 | 999.912 | 0.0071833 | 0.31298 | 490.876 | 537.101
Populus deltoides | LI1C |  | RACiR 300 to 800 | 376.0626 | 33.7853 | 24.9877 | 999.911 | 0.0071835 | 0.31288 | 494.032 | 540.541
Populus deltoides | LI1C |  | RACiR 300 to 800 | 378.9323 | 33.8804 | 24.9861 | 999.905 | 0.0071874 | 0.31322 | 497.15 | 543.81
Populus deltoides | LI1C |  | RACiR 300 to 800 | 381.7744 | 33.958 | 24.9894 | 999.918 | 0.0071944 | 0.31343 | 500.24 | 546.998
Populus deltoides | LI1C |  | RACiR 300 to 800 | 384.016 | 34.2018 | 24.9874 | 999.916 | 0.0071888 | 0.31333 | 503.353 | 550.431
Populus deltoides | LI1C |  | RACiR 300 to 800 | 386.596 | 34.4193 | 24.9867 | 999.917 | 0.0071993 | 0.31395 | 506.47 | 553.787
Populus deltoides | LI1C |  | RACiR 300 to 800 | 389.5076 | 34.5005 | 24.9907 | 999.916 | 0.0072057 | 0.31412 | 509.653 | 557.086
Populus deltoides | LI1C |  | RACiR 300 to 800 | 391.3182 | 34.8933 | 24.9945 | 1000.22 | 0.0072046 | 0.31393 | 512.843 | 560.737
Populus deltoides | LI1C |  | RACiR 300 to 800 | 394.1345 | 35.0292 | 24.9952 | 999.908 | 0.0072114 | 0.31424 | 516.052 | 564.096
Populus deltoides | LI1C |  | RACiR 300 to 800 | 397.5064 | 34.9033 | 25.0031 | 999.907 | 0.0072076 | 0.31372 | 519.296 | 567.304
Populus deltoides | LI1C |  | RACiR 300 to 800 | 400.167 | 35.1015 | 25.0061 | 1000.22 | 0.0072181 | 0.31415 | 522.5 | 570.724
Populus deltoides | LI1C |  | RACiR 300 to 800 | 403.2258 | 35.138 | 25.0075 | 999.912 | 0.0072217 | 0.31427 | 525.706 | 573.988
Populus deltoides | LI1C |  | RACiR 300 to 800 | 406.4101 | 35.0693 | 25.0092 | 999.9 | 0.0072144 | 0.31392 | 528.873 | 577.184
Populus deltoides | LI1C |  | RACiR 300 to 800 | 408.3219 | 35.5003 | 25.0088 | 1000.22 | 0.0072253 | 0.31451 | 531.999 | 580.754
Populus deltoides | LI1C |  | RACiR 300 to 800 | 411.2429 | 35.6101 | 25.0047 | 999.914 | 0.0072279 | 0.31487 | 535.201 | 584.107
Populus deltoides | LI1C |  | RACiR 300 to 800 | 414.6082 | 35.5586 | 25.001 | 999.92 | 0.0072289 | 0.31512 | 538.381 | 587.282
Populus deltoides | LI1C |  | RACiR 300 to 800 | 416.9317 | 35.7703 | 25.006 | 1000.22 | 0.0072275 | 0.3148 | 541.563 | 590.724
Populus deltoides | LI1C |  | RACiR 300 to 800 | 419.773 | 35.8396 | 25.0068 | 1000.22 | 0.0072264 | 0.3147 | 544.731 | 594.003
Populus deltoides | LI1C |  | RACiR 300 to 800 | 423.0505 | 35.802 | 25.0006 | 999.917 | 0.0072259 | 0.31497 | 547.854 | 597.127
Populus deltoides | LI1C |  | RACiR 300 to 800 | 425.4261 | 36.0487 | 25.0001 | 999.915 | 0.0072284 | 0.31515 | 551.011 | 600.575
Populus deltoides | LI1C |  | RACiR 300 to 800 | 428.0201 | 36.2728 | 25.0041 | 999.931 | 0.0072432 | 0.31567 | 554.196 | 604.031
Populus deltoides | LI1C |  | RACiR 300 to 800 | 431.2772 | 36.3336 | 25.0002 | 999.922 | 0.0072599 | 0.31654 | 557.397 | 607.354
Populus deltoides | LI1C |  | RACiR 300 to 800 | 433.4376 | 36.6766 | 24.9946 | 1000.22 | 0.0072645 | 0.31685 | 560.586 | 610.93
Populus deltoides | LI1C |  | RACiR 300 to 800 | 436.0083 | 36.8198 | 24.9929 | 999.913 | 0.0072576 | 0.31645 | 563.825 | 614.346
Populus deltoides | LI1C |  | RACiR 300 to 800 | 439.3248 | 36.8183 | 24.9906 | 999.918 | 0.007265 | 0.3168 | 567.083 | 617.647
Populus deltoides | LI1C |  | RACiR 300 to 800 | 442.6887 | 36.957 | 24.9896 | 1000.21 | 0.0073047 | 0.31849 | 570.342 | 621.144
Populus deltoides | LI1C |  | RACiR 300 to 800 | 446.3133 | 36.9258 | 24.9902 | 999.912 | 0.0073404 | 0.3196 | 573.561 | 624.366
Populus deltoides | LI1C |  | RACiR 300 to 800 | 449.4872 | 36.7902 | 24.9855 | 999.91 | 0.0073278 | 0.31871 | 576.721 | 627.41
Populus deltoides | LI1C |  | RACiR 300 to 800 | 451.6933 | 36.9696 | 24.9796 | 1000.22 | 0.007309 | 0.31787 | 579.857 | 630.807
Populus deltoides | LI1C |  | RACiR 300 to 800 | 453.9776 | 37.2281 | 24.9789 | 1000.22 | 0.0073101 | 0.31777 | 583.06 | 634.268
Populus deltoides | LI1C |  | RACiR 300 to 800 | 456.8381 | 37.267 | 24.9787 | 999.916 | 0.0073021 | 0.31727 | 586.305 | 637.609
Populus deltoides | LI1C |  | RACiR 300 to 800 | 459.2236 | 37.4939 | 24.9761 | 999.919 | 0.0072977 | 0.31715 | 589.509 | 641.088
Populus deltoides | LI1C |  | RACiR 300 to 800 | 462.5991 | 37.4307 | 24.9739 | 999.918 | 0.0072992 | 0.31725 | 592.723 | 644.248
Populus deltoides | LI1C |  | RACiR 300 to 800 | 466.1832 | 37.285 | 24.9646 | 999.924 | 0.0072888 | 0.31728 | 595.898 | 647.378
Populus deltoides | LI1C |  | RACiR 300 to 800 | 468.3578 | 37.6191 | 24.9661 | 1000.23 | 0.0072996 | 0.31773 | 599.017 | 650.838
Populus deltoides | LI1C |  | RACiR 300 to 800 | 471.0642 | 37.7718 | 24.9698 | 999.915 | 0.0073057 | 0.3178 | 602.254 | 654.247
Populus deltoides | LI1C |  | RACiR 300 to 800 | 474.2616 | 37.7153 | 24.9705 | 999.909 | 0.0072962 | 0.3174 | 605.499 | 657.544
Populus deltoides | LI1C |  | RACiR 300 to 800 | 476.4752 | 38.0702 | 24.9692 | 999.913 | 0.0073048 | 0.31784 | 608.722 | 661.119
Populus deltoides | LI1C |  | RACiR 300 to 800 | 479.0521 | 38.2494 | 24.9675 | 999.917 | 0.0073016 | 0.31773 | 611.97 | 664.594
Populus deltoides | LI1C |  | RACiR 300 to 800 | 482.5709 | 38.114 | 24.9704 | 999.919 | 0.0072983 | 0.31748 | 615.239 | 667.783
Populus deltoides | LI1C |  | RACiR 300 to 800 | 485.364 | 38.2569 | 24.9714 | 1000.23 | 0.0073029 | 0.31765 | 618.492 | 671.2
Populus deltoides | LI1C |  | RACiR 300 to 800 | 488.789 | 38.1907 | 24.966 | 1000.22 | 0.0072989 | 0.31774 | 621.738 | 674.464
Populus deltoides | LI1C |  | RACiR 300 to 800 | 492.4942 | 38.1199 | 24.9568 | 999.915 | 0.0073084 | 0.31868 | 624.925 | 677.588
Populus deltoides | LI1C |  | RACiR 300 to 800 | 494.791 | 38.3432 | 24.9615 | 999.917 | 0.007308 | 0.31842 | 628.094 | 681.035
Populus deltoides | LI1C |  | RACiR 300 to 800 | 497.2771 | 38.5105 | 24.9683 | 1000.22 | 0.0073053 | 0.31803 | 631.33 | 684.516
Populus deltoides | LI1C |  | RACiR 300 to 800 | 500.8007 | 38.4752 | 24.9647 | 999.912 | 0.0073128 | 0.31856 | 634.613 | 687.76
Populus deltoides | LI1C |  | RACiR 300 to 800 | 503.6366 | 38.5464 | 24.9648 | 999.911 | 0.0073069 | 0.31826 | 637.857 | 691.154
Populus deltoides | LI1C |  | RACiR 300 to 800 | 506.6121 | 38.5635 | 24.9726 | 1000.22 | 0.0073096 | 0.31801 | 641.066 | 694.408
Populus deltoides | LI1C |  | RACiR 300 to 800 | 509.7476 | 38.5168 | 24.9826 | 1000.23 | 0.0073139 | 0.31771 | 644.264 | 697.578
Populus deltoides | LI1C |  | RACiR 300 to 800 | 512.6469 | 38.6221 | 24.9758 | 999.927 | 0.0073103 | 0.31794 | 647.458 | 700.979
Populus deltoides | LI1C |  | RACiR 300 to 800 | 515.5171 | 38.8282 | 24.9662 | 1000.22 | 0.007318 | 0.31878 | 650.706 | 704.435
Populus deltoides | LI1C |  | RACiR 300 to 800 | 518.5473 | 38.8557 | 24.9696 | 999.909 | 0.0073153 | 0.3185 | 654.01 | 707.816
Populus deltoides | LI1C |  | RACiR 300 to 800 | 521.3704 | 38.9511 | 24.9735 | 999.924 | 0.007313 | 0.3183 | 657.282 | 711.271
Populus deltoides | LI1C |  | RACiR 300 to 800 | 525.3451 | 38.7685 | 24.9724 | 999.932 | 0.0073238 | 0.31889 | 660.537 | 714.341
Populus deltoides | LI1C |  | RACiR 300 to 800 | 528.0229 | 38.8898 | 24.9702 | 999.914 | 0.0073186 | 0.31873 | 663.721 | 717.707
Populus deltoides | LI1C |  | RACiR 300 to 800 | 529.7573 | 39.363 | 24.9737 | 999.918 | 0.0073227 | 0.31878 | 666.983 | 721.486
Populus deltoides | LI1C |  | RACiR 300 to 800 | 533.7552 | 39.2009 | 24.9708 | 1000.22 | 0.0073277 | 0.31923 | 670.372 | 724.747
Populus deltoides | LI1C |  | RACiR 300 to 800 | 537.1853 | 39.0382 | 24.9849 | 999.913 | 0.007327 | 0.31853 | 673.67 | 727.94
Populus deltoides | LI1C |  | RACiR 300 to 800 | 539.77 | 39.2134 | 24.9919 | 999.911 | 0.0073328 | 0.31843 | 676.923 | 731.399
Populus deltoides | LI1C |  | RACiR 300 to 800 | 543.206 | 39.139 | 24.9873 | 1000.23 | 0.0073269 | 0.31844 | 680.19 | 734.677
Populus deltoides | LI1C |  | RACiR 300 to 800 | 546.8918 | 39.0281 | 24.9897 | 999.924 | 0.0073349 | 0.31872 | 683.504 | 737.88
Populus deltoides | LI1C |  | RACiR 300 to 800 | 549.967 | 39.0789 | 24.989 | 999.918 | 0.0073368 | 0.31881 | 686.775 | 741.245
Populus deltoides | LI1C |  | RACiR 300 to 800 | 553.1511 | 39.0321 | 24.9937 | 999.917 | 0.0073355 | 0.31861 | 689.971 | 744.493
Populus deltoides | LI1C |  | RACiR 300 to 800 | 556.461 | 38.9805 | 25.0046 | 999.912 | 0.0073475 | 0.31868 | 693.187 | 747.652
Populus deltoides | LI1C |  | RACiR 300 to 800 | 559.3343 | 39.0574 | 25.0022 | 1000.23 | 0.0073395 | 0.31846 | 696.45 | 751.072
Populus deltoides | LI1C |  | RACiR 300 to 800 | 562.4979 | 39.1666 | 24.9934 | 999.931 | 0.0073446 | 0.31917 | 699.734 | 754.484
Populus deltoides | LI1C |  | RACiR 300 to 800 | 565.6181 | 39.191 | 24.9964 | 999.913 | 0.0073471 | 0.31911 | 703.034 | 757.84
Populus deltoides | LI1C |  | RACiR 300 to 800 | 568.1891 | 39.347 | 25.0032 | 999.916 | 0.0073439 | 0.31869 | 706.34 | 761.388
Populus deltoides | LI1C |  | RACiR 300 to 800 | 571.441 | 39.4066 | 24.9995 | 999.906 | 0.0073497 | 0.31916 | 709.662 | 764.782
Populus deltoides | LI1C |  | RACiR 300 to 800 | 575.2999 | 39.2779 | 24.9869 | 999.911 | 0.0073445 | 0.31956 | 713.018 | 768.085
Populus deltoides | LI1C |  | RACiR 300 to 800 | 578.2303 | 39.369 | 24.9933 | 999.906 | 0.0073481 | 0.31946 | 716.357 | 771.544
Populus deltoides | LI1C |  | RACiR 300 to 800 | 581.7946 | 39.2562 | 24.9962 | 1000.22 | 0.0073494 | 0.31939 | 719.679 | 774.78
Populus deltoides | LI1C |  | RACiR 300 to 800 | 585.4828 | 39.1174 | 24.9952 | 1000.23 | 0.0073501 | 0.31949 | 722.971 | 777.996
Populus deltoides | LI1C |  | RACiR 300 to 800 | 588.3522 | 39.2682 | 24.9925 | 999.91 | 0.0073535 | 0.31977 | 726.27 | 781.473
Populus deltoides | LI1C |  | RACiR 300 to 800 | 591.4141 | 39.3171 | 24.9887 | 1000.21 | 0.0073487 | 0.31975 | 729.554 | 784.885
Populus deltoides | LI1C |  | RACiR 300 to 800 | 594.5659 | 39.3051 | 24.9992 | 1000.22 | 0.0073569 | 0.31967 | 732.795 | 788.136
Populus deltoides | LI1C |  | RACiR 300 to 800 | 597.7625 | 39.3352 | 24.9983 | 999.908 | 0.0073602 | 0.31986 | 736.084 | 791.484
Populus deltoides | LI1C |  | RACiR 0 to 500 | 103.8008 | 5.8907 | 24.9636 | 1000.01 | 0.0065306 | 0.28525 | 127.095 | 139.079
Populus deltoides | LI1C |  | RACiR 0 to 500 | 106.2594 | 6.0374 | 24.9609 | 1000.02 | 0.0065376 | 0.28571 | 130.095 | 142.261
Populus deltoides | LI1C |  | RACiR 0 to 500 | 108.2018 | 6.3301 | 24.9595 | 1000.01 | 0.0065434 | 0.28601 | 133.095 | 145.608
Populus deltoides | LI1C |  | RACiR 0 to 500 | 110.9276 | 6.4224 | 24.959 | 1000.03 | 0.0065764 | 0.28749 | 136.092 | 148.741
Populus deltoides | LI1C |  | RACiR 0 to 500 | 113.6421 | 6.4809 | 24.9612 | 1000.03 | 0.0065943 | 0.28798 | 139.04 | 151.789
Populus deltoides | LI1C |  | RACiR 0 to 500 | 115.3283 | 6.796 | 24.9629 | 1000.03 | 0.0065747 | 0.2868 | 141.974 | 155.094
Populus deltoides | LI1C |  | RACiR 0 to 500 | 117.3616 | 7.0428 | 24.9578 | 1000.03 | 0.0065755 | 0.28695 | 144.91 | 158.309
Populus deltoides | LI1C |  | RACiR 0 to 500 | 119.6485 | 7.2051 | 24.9569 | 999.712 | 0.0065724 | 0.28676 | 147.84 | 161.446
Populus deltoides | LI1C |  | RACiR 0 to 500 | 121.1343 | 7.5973 | 24.9596 | 1000.03 | 0.0065685 | 0.2865 | 150.769 | 164.819
Populus deltoides | LI1C |  | RACiR 0 to 500 | 123.3685 | 7.7909 | 24.9594 | 1000.02 | 0.006572 | 0.28665 | 153.724 | 168.004
Populus deltoides | LI1C |  | RACiR 0 to 500 | 125.9157 | 7.8741 | 24.9611 | 1000.03 | 0.0065686 | 0.28641 | 156.648 | 171.063
Populus deltoides | LI1C |  | RACiR 0 to 500 | 127.681 | 8.2059 | 24.9568 | 1000.03 | 0.0065742 | 0.28693 | 159.571 | 174.357
Populus deltoides | LI1C |  | RACiR 0 to 500 | 129.8418 | 8.4131 | 24.9589 | 1000.03 | 0.0065744 | 0.28682 | 162.525 | 177.562
Populus deltoides | LI1C |  | RACiR 0 to 500 | 132.1938 | 8.5499 | 24.9679 | 1000.04 | 0.0065745 | 0.28644 | 165.459 | 180.682
Populus deltoides | LI1C |  | RACiR 0 to 500 | 133.6481 | 8.9712 | 24.967 | 1000.03 | 0.0065789 | 0.28671 | 168.404 | 184.086
Populus deltoides | LI1C |  | RACiR 0 to 500 | 135.7006 | 9.2096 | 24.9657 | 999.72 | 0.0065703 | 0.28646 | 171.372 | 187.353
Populus deltoides | LI1C |  | RACiR 0 to 500 | 137.9794 | 9.4216 | 24.9641 | 1000.04 | 0.0065808 | 0.28707 | 174.379 | 190.606
Populus deltoides | LI1C |  | RACiR 0 to 500 | 139.4903 | 9.8445 | 24.9632 | 1000.03 | 0.0065832 | 0.2872 | 177.4 | 194.094
Populus deltoides | LI1C |  | RACiR 0 to 500 | 141.8274 | 10.0151 | 24.9647 | 1000.03 | 0.0065795 | 0.28699 | 180.419 | 197.338
Populus deltoides | LI1C |  | RACiR 0 to 500 | 144.4781 | 10.1129 | 24.9593 | 1000.04 | 0.0065879 | 0.28766 | 183.388 | 200.439
Populus deltoides | LI1C |  | RACiR 0 to 500 | 146.1248 | 10.4603 | 24.9672 | 1000.03 | 0.0065858 | 0.28724 | 186.353 | 203.812
Populus deltoides | LI1C |  | RACiR 0 to 500 | 148.1232 | 10.7412 | 24.9691 | 1000.04 | 0.0065916 | 0.2875 | 189.352 | 207.129
Populus deltoides | LI1C |  | RACiR 0 to 500 | 150.3495 | 10.9661 | 24.9625 | 1000.04 | 0.0065971 | 0.28808 | 192.339 | 210.378
Populus deltoides | LI1C |  | RACiR 0 to 500 | 151.8562 | 11.3649 | 24.9671 | 1000.03 | 0.006589 | 0.28757 | 195.355 | 213.864
Populus deltoides | LI1C |  | RACiR 0 to 500 | 154.3836 | 11.5194 | 24.9666 | 1000.05 | 0.0065998 | 0.28819 | 198.394 | 217.086
Populus deltoides | LI1C |  | RACiR 0 to 500 | 157.1311 | 11.5737 | 24.9721 | 999.736 | 0.0066003 | 0.28801 | 201.424 | 220.217
Populus deltoides | LI1C |  | RACiR 0 to 500 | 158.5253 | 12.0196 | 24.9741 | 1000.06 | 0.0066025 | 0.28806 | 204.41 | 223.7
Populus deltoides | LI1C |  | RACiR 0 to 500 | 160.4889 | 12.3342 | 24.9721 | 999.736 | 0.0066106 | 0.28857 | 207.446 | 227.086
Populus deltoides | LI1C |  | RACiR 0 to 500 | 163.1346 | 12.425 | 24.9735 | 1000.06 | 0.0066061 | 0.2884 | 210.497 | 230.296
Populus deltoides | LI1C |  | RACiR 0 to 500 | 164.7449 | 12.8313 | 24.9784 | 1000.05 | 0.0066143 | 0.28863 | 213.535 | 233.773
Populus deltoides | LI1C |  | RACiR 0 to 500 | 166.5902 | 13.1553 | 24.9796 | 999.722 | 0.0066153 | 0.28858 | 216.568 | 237.172
Populus deltoides | LI1C |  | RACiR 0 to 500 | 169.1982 | 13.2963 | 24.9741 | 1000.03 | 0.0066199 | 0.28913 | 219.636 | 240.449
Populus deltoides | LI1C |  | RACiR 0 to 500 | 171.5203 | 13.6192 | 24.9752 | 1000.04 | 0.0066804 | 0.29181 | 222.705 | 243.882
Populus deltoides | LI1C |  | RACiR 0 to 500 | 174.3939 | 13.7067 | 24.9761 | 999.733 | 0.0067145 | 0.29291 | 225.776 | 247.1
Populus deltoides | LI1C |  | RACiR 0 to 500 | 176.6297 | 13.8725 | 24.9736 | 1000.04 | 0.0067025 | 0.29208 | 228.777 | 250.309
Populus deltoides | LI1C |  | RACiR 0 to 500 | 177.8877 | 14.3247 | 24.9754 | 1000.03 | 0.0066953 | 0.29133 | 231.765 | 253.789
Populus deltoides | LI1C |  | RACiR 0 to 500 | 179.5548 | 14.678 | 24.9764 | 1000.04 | 0.0066834 | 0.29059 | 234.832 | 257.275
Populus deltoides | LI1C |  | RACiR 0 to 500 | 181.8991 | 14.8777 | 24.9761 | 1000.04 | 0.006687 | 0.2907 | 237.907 | 260.575
Populus deltoides | LI1C |  | RACiR 0 to 500 | 183.3942 | 15.285 | 24.9815 | 1000.04 | 0.0066831 | 0.29015 | 240.961 | 264.082
Populus deltoides | LI1C |  | RACiR 0 to 500 | 185.8184 | 15.4334 | 24.9869 | 999.741 | 0.0066763 | 0.28962 | 244.061 | 267.397
Populus deltoides | LI1C |  | RACiR 0 to 500 | 188.7982 | 15.4718 | 24.9813 | 1000.04 | 0.0066859 | 0.29035 | 247.104 | 270.502
Populus deltoides | LI1C |  | RACiR 0 to 500 | 190.7999 | 15.7617 | 24.9731 | 1000.05 | 0.0066816 | 0.29057 | 250.114 | 273.864
Populus deltoides | LI1C |  | RACiR 0 to 500 | 192.7406 | 16.0776 | 24.9734 | 1000.04 | 0.006688 | 0.29098 | 253.117 | 277.228
Populus deltoides | LI1C |  | RACiR 0 to 500 | 194.794 | 16.3436 | 24.9755 | 1000.05 | 0.0066926 | 0.29113 | 256.115 | 280.529
Populus deltoides | LI1C |  | RACiR 0 to 500 | 196.3403 | 16.7429 | 24.9769 | 1000.03 | 0.0066871 | 0.2908 | 259.153 | 284.025
Populus deltoides | LI1C |  | RACiR 0 to 500 | 198.3355 | 17.0675 | 24.9734 | 1000.04 | 0.0066911 | 0.29118 | 262.243 | 287.481
Populus deltoides | LI1C |  | RACiR 0 to 500 | 200.8691 | 17.2207 | 24.9736 | 1000.04 | 0.0066955 | 0.29141 | 265.319 | 290.758
Populus deltoides | LI1C |  | RACiR 0 to 500 | 202.7603 | 17.5417 | 24.9796 | 1000.05 | 0.0067018 | 0.29149 | 268.356 | 294.162
Populus deltoides | LI1C |  | RACiR 0 to 500 | 205.4719 | 17.6439 | 24.976 | 1000.05 | 0.006701 | 0.29174 | 271.428 | 297.393
Populus deltoides | LI1C |  | RACiR 0 to 500 | 207.9885 | 17.7897 | 24.9805 | 1000.05 | 0.0067082 | 0.29194 | 274.467 | 300.612
Populus deltoides | LI1C |  | RACiR 0 to 500 | 209.4304 | 18.2154 | 24.9907 | 999.746 | 0.0067124 | 0.29169 | 277.481 | 304.1
Populus deltoides | LI1C |  | RACiR 0 to 500 | 211.4148 | 18.5071 | 24.9893 | 1000.05 | 0.006709 | 0.29166 | 280.528 | 307.503
Populus deltoides | LI1C |  | RACiR 0 to 500 | 213.8838 | 18.7095 | 24.9846 | 1000.04 | 0.0067177 | 0.29233 | 283.602 | 310.808
Populus deltoides | LI1C |  | RACiR 0 to 500 | 215.6557 | 19.1125 | 24.9793 | 999.738 | 0.0067199 | 0.29275 | 286.707 | 314.377
Populus deltoides | LI1C |  | RACiR 0 to 500 | 218.1387 | 19.2782 | 24.9856 | 1000.05 | 0.0067244 | 0.29276 | 289.822 | 317.702
Populus deltoides | LI1C |  | RACiR 0 to 500 | 221.1169 | 19.2969 | 24.9907 | 1000.05 | 0.0067307 | 0.29284 | 292.92 | 320.852
Populus deltoides | LI1C |  | RACiR 0 to 500 | 223.1929 | 19.5682 | 24.9843 | 1000.05 | 0.0067237 | 0.2929 | 295.959 | 324.243
Populus deltoides | LI1C |  | RACiR 0 to 500 | 225.2067 | 19.8881 | 24.9847 | 1000.03 | 0.0067378 | 0.29362 | 298.954 | 327.583
Populus deltoides | LI1C |  | RACiR 0 to 500 | 227.6662 | 20.059 | 24.9872 | 1000.04 | 0.0067427 | 0.29375 | 302.031 | 330.871
Populus deltoides | LI1C |  | RACiR 0 to 500 | 229.3867 | 20.4329 | 24.9861 | 999.73 | 0.0067363 | 0.29357 | 305.121 | 334.404
Populus deltoides | LI1C |  | RACiR 0 to 500 | 231.5349 | 20.7066 | 24.9881 | 1000.04 | 0.006743 | 0.29383 | 308.2 | 337.788
Populus deltoides | LI1C |  | RACiR 0 to 500 | 234.0989 | 20.85 | 24.9889 | 1000.03 | 0.0067443 | 0.29393 | 311.29 | 341.081
Populus deltoides | LI1C |  | RACiR 0 to 500 | 236.0359 | 21.1898 | 24.9891 | 999.745 | 0.0067494 | 0.29424 | 314.365 | 344.54
Populus deltoides | LI1C |  | RACiR 0 to 500 | 238.7667 | 21.3 | 24.9876 | 1000.04 | 0.0067534 | 0.29453 | 317.462 | 347.789
Populus deltoides | LI1C |  | RACiR 0 to 500 | 241.2518 | 21.448 | 24.9926 | 1000.05 | 0.0067559 | 0.29451 | 320.522 | 351.051
Populus deltoides | LI1C |  | RACiR 0 to 500 | 242.898 | 21.8234 | 25.0019 | 999.734 | 0.006759 | 0.29424 | 323.573 | 354.519
Populus deltoides | LI1C |  | RACiR 0 to 500 | 245.0803 | 22.0779 | 24.9999 | 1000.06 | 0.0067597 | 0.29445 | 326.623 | 357.887
Populus deltoides | LI1C |  | RACiR 0 to 500 | 247.5343 | 22.2791 | 24.999 | 1000.05 | 0.0067688 | 0.29501 | 329.677 | 361.176
Populus deltoides | LI1C |  | RACiR 0 to 500 | 249.3398 | 22.6574 | 24.9998 | 1000.05 | 0.0067727 | 0.29517 | 332.778 | 364.706
Populus deltoides | LI1C |  | RACiR 0 to 500 | 252.1614 | 22.7885 | 25.001 | 1000.05 | 0.0067924 | 0.29602 | 335.891 | 368.017
Populus deltoides | LI1C |  | RACiR 0 to 500 | 255.4603 | 22.7778 | 25.0027 | 1000.05 | 0.0068154 | 0.29691 | 339 | 371.151
Populus deltoides | LI1C |  | RACiR 0 to 500 | 257.2756 | 23.1256 | 24.9982 | 1000.05 | 0.0068114 | 0.29681 | 342.066 | 374.623
Populus deltoides | LI1C |  | RACiR 0 to 500 | 259.309 | 23.4135 | 24.9976 | 1000.05 | 0.0068112 | 0.29679 | 345.133 | 378.022
Populus deltoides | LI1C |  | RACiR 0 to 500 | 261.9446 | 23.5036 | 25.001 | 1000.06 | 0.0068114 | 0.29659 | 348.198 | 381.217
Populus deltoides | LI1C |  | RACiR 0 to 500 | 263.9845 | 23.9181 | 25.0024 | 999.751 | 0.0068489 | 0.29809 | 351.291 | 384.799
Populus deltoides | LI1C |  | RACiR 0 to 500 | 266.8015 | 24.1574 | 24.9997 | 1000.06 | 0.0068984 | 0.30003 | 354.46 | 388.263
Populus deltoides | LI1C |  | RACiR 0 to 500 | 269.857 | 24.1584 | 24.9899 | 1000.06 | 0.0068996 | 0.29999 | 357.604 | 391.457
Populus deltoides | LI1C |  | RACiR 0 to 500 | 271.7273 | 24.4017 | 24.9883 | 1000.05 | 0.0068795 | 0.29872 | 360.697 | 394.841
Populus deltoides | LI1C |  | RACiR 0 to 500 | 274.6549 | 24.3795 | 24.9887 | 1000.05 | 0.0068699 | 0.29802 | 363.82 | 397.983
Populus deltoides | LI1C |  | RACiR 0 to 500 | 277.4521 | 24.4243 | 24.9803 | 999.743 | 0.006866 | 0.29806 | 366.817 | 401.051
Populus deltoides | LI1C |  | RACiR 0 to 500 | 278.9209 | 24.8698 | 24.9717 | 1000.06 | 0.0068595 | 0.29806 | 369.825 | 404.559
Populus deltoides | LI1C |  | RACiR 0 to 500 | 280.658 | 25.215 | 24.9724 | 1000.05 | 0.0068508 | 0.2976 | 372.908 | 408.059
Populus deltoides | LI1C |  | RACiR 0 to 500 | 283.0628 | 25.416 | 24.9738 | 1000.06 | 0.0068548 | 0.2977 | 376.027 | 411.395
Populus deltoides | LI1C |  | RACiR 0 to 500 | 284.8653 | 25.8061 | 24.969 | 1000.06 | 0.00685 | 0.2977 | 379.191 | 415.016
Populus deltoides | LI1C |  | RACiR 0 to 500 | 287.6645 | 25.9106 | 24.9697 | 1000.04 | 0.0068541 | 0.29788 | 382.355 | 418.321
Populus deltoides | LI1C |  | RACiR 0 to 500 | 291.4153 | 25.7465 | 24.9642 | 1000.05 | 0.0068553 | 0.29822 | 385.531 | 421.369
Populus deltoides | LI1C |  | RACiR 0 to 500 | 293.2745 | 26.0868 | 24.9593 | 1000.06 | 0.0068482 | 0.29812 | 388.61 | 424.88
Populus deltoides | LI1C |  | RACiR 0 to 500 | 296.1563 | 26.1803 | 24.9563 | 1000.06 | 0.0068587 | 0.2987 | 391.7 | 428.071
Populus deltoides | LI1C |  | RACiR 0 to 500 | 298.4785 | 26.3757 | 24.9547 | 1000.06 | 0.0068554 | 0.29858 | 394.773 | 431.383
Populus deltoides | LI1C |  | RACiR 0 to 500 | 299.2096 | 26.9601 | 24.9592 | 1000.06 | 0.0068312 | 0.29743 | 397.859 | 435.136
Populus deltoides | LI1C |  | RACiR 0 to 500 | 301.6693 | 27.1229 | 24.9656 | 999.762 | 0.0068236 | 0.29699 | 401.069 | 438.512
Populus deltoides | LI1C |  | RACiR 0 to 500 | 304.9885 | 27.0843 | 24.9638 | 1000.08 | 0.0068215 | 0.29713 | 404.282 | 441.743
Populus deltoides | LI1C |  | RACiR 0 to 500 | 307.5154 | 27.3286 | 24.9582 | 1000.07 | 0.006828 | 0.29789 | 407.45 | 445.214
Populus deltoides | LI1C |  | RACiR 0 to 500 | 310.2743 | 27.4409 | 24.9612 | 999.767 | 0.006837 | 0.29824 | 410.545 | 448.453
Populus deltoides | LI1C |  | RACiR 0 to 500 | 312.9323 | 27.53 | 24.9641 | 1000.08 | 0.0068335 | 0.298 | 413.644 | 451.718
Populus deltoides | LI1C |  | RACiR 0 to 500 | 314.9854 | 27.8662 | 24.9635 | 1000.09 | 0.006842 | 0.2985 | 416.727 | 455.164
Populus deltoides | LI1C |  | RACiR 0 to 500 | 317.4518 | 28.0567 | 24.9609 | 1000.09 | 0.0068437 | 0.29874 | 419.813 | 458.494
Populus deltoides | LI1C |  | RACiR 0 to 500 | 320.0889 | 28.2079 | 24.9594 | 1000.1 | 0.006847 | 0.29906 | 422.916 | 461.802
Populus deltoides | LI1C |  | RACiR 0 to 500 | 321.9584 | 28.547 | 24.9679 | 1000.1 | 0.0068507 | 0.29882 | 426.064 | 465.323
Populus deltoides | LI1C |  | RACiR 0 to 500 | 324.6412 | 28.6225 | 24.9718 | 1000.1 | 0.0068443 | 0.29839 | 429.211 | 468.633
Populus deltoides | LI1C |  | RACiR 0 to 500 | 327.8959 | 28.6435 | 24.9647 | 1000.1 | 0.0068541 | 0.29923 | 432.326 | 471.776
Populus deltoides | LI1C |  | RACiR 0 to 500 | 330.243 | 28.8729 | 24.9583 | 999.793 | 0.0068528 | 0.29947 | 435.418 | 475.146
Populus deltoides | LI1C |  | RACiR 0 to 500 | 332.5679 | 29.0722 | 24.9614 | 1000.09 | 0.0068525 | 0.29937 | 438.503 | 478.507
Populus deltoides | LI1C |  | RACiR 0 to 500 | 334.892 | 29.2775 | 24.9687 | 1000.1 | 0.0068588 | 0.29936 | 441.587 | 481.821
Populus deltoides | LI1C |  | RACiR 0 to 500 | 336.7156 | 29.6416 | 24.9688 | 1000.1 | 0.0068561 | 0.29926 | 444.717 | 485.376
Populus deltoides | LI1C | 3 | Standard | 278.8384 | 24.5559 | 24.9718 | 999.858 | 0.0064791 | 0.28315 | 373.022 | 399.994
Populus deltoides | LI1C | 3 | Standard | 243.0916 | 22.0135 | 25.0321 | 999.837 | 0.0066091 | 0.28892 | 325.79 | 349.96
Populus deltoides | LI1C | 3 | Standard | 209.4007 | 19.0684 | 24.9886 | 1000.18 | 0.0067546 | 0.29746 | 279.05 | 300
Populus deltoides | LI1C | 3 | Standard | 176.9982 | 15.642 | 24.9715 | 1000.15 | 0.0069909 | 0.30598 | 232.801 | 250.065
Populus deltoides | LI1C | 3 | Standard | 144.6738 | 11.9623 | 25.0325 | 1000.13 | 0.0071612 | 0.31318 | 186.658 | 199.954
Populus deltoides | LI1C | 3 | Standard | 112.9792 | 8.0703 | 25.0038 | 1000.12 | 0.0073068 | 0.32208 | 140.916 | 150.016
Populus deltoides | LI1C | 3 | Standard | 81.0626 | 4.038 | 24.9757 | 999.772 | 0.0075391 | 0.33094 | 95.2724 | 100.029
Populus deltoides | LI1C | 3 | Standard | 49.2734 | -0.17747 | 25.0258 | 1000.11 | 0.0076981 | 0.33865 | 49.8763 | 50.0828
Populus deltoides | LI1C | 3 | Standard | 289.5843 | 26.1082 | 25.0031 | 1000.14 | 0.0080357 | 0.35322 | 371.014 | 400.104
Populus deltoides | LI1C | 3 | Standard | 365.1461 | 32.4235 | 24.9979 | 999.907 | 0.0082325 | 0.36304 | 463.858 | 500.095
Populus deltoides | LI1C | 3 | Standard | 447.4804 | 36.8466 | 25.0221 | 999.922 | 0.0083534 | 0.36938 | 558.608 | 600.121
Populus deltoides | LI1C | 3 | Standard | 535.9625 | 39.5561 | 25.0278 | 999.946 | 0.0084223 | 0.37443 | 655.001 | 700.06
Populus deltoides | LI1C | 3 | Standard | 629.986 | 40.7051 | 24.9858 | 1000.28 | 0.0086247 | 0.37998 | 752.805 | 800.011
Populus deltoides | LI1C | 3 | Standard | 723.838 | 41.6887 | 25.0506 | 999.923 | 0.0086558 | 0.38138 | 851.018 | 900.066
Populus deltoides | LI1C | 3 | Standard | 820.5408 | 42.011 | 25.0041 | 999.89 | 0.0087526 | 0.38497 | 949.793 | 1000.11
Populus deltoides | LI1C | 3 | Standard | 915.3471 | 42.5402 | 25.0104 | 1000.23 | 0.0087026 | 0.38395 | 1048.4 | 1100.06
Populus deltoides | LI1C | 3 | Standard | 1012.4646 | 42.5809 | 25.028 | 1000.23 | 0.0087105 | 0.38505 | 1147.48 | 1200.05
Populus deltoides | LI1C | 3 | Standard | 1110.5702 | 42.5021 | 24.9972 | 999.879 | 0.0088134 | 0.38892 | 1246.56 | 1300.05
Populus deltoides | LI1C | 3 | Standard | 1205.3827 | 42.8075 | 25.0303 | 1000.19 | 0.0087396 | 0.38496 | 1345.54 | 1400.1
Populus deltoides | LI1C | 3 | Standard | 1302.1856 | 42.8399 | 25.0145 | 999.863 | 0.0087612 | 0.38527 | 1444.61 | 1500.12
Populus deltoides | LI1C | 3 | Standard | 1592.2264 | 43.0469 | 25.0093 | 999.803 | 0.0087121 | 0.38469 | 1741.88 | 1800.08
Populus deltoides | LI1C | 3 | Standard | 1785.3658 | 43.1669 | 25.0173 | 1000.15 | 0.0087271 | 0.38408 | 1940.08 | 2000.19
Populus deltoides | LI1C | 2 | Standard | 239.5897 | 17.455 | 25.034 | 999.71 | 0.0030151 | 0.12968 | 381.415 | 400.016
Populus deltoides | LI1C | 2 | Standard | 210.9844 | 14.9274 | 24.9961 | 999.987 | 0.0029383 | 0.12767 | 334.177 | 350.086
Populus deltoides | LI1C | 2 | Standard | 185.298 | 12.4554 | 25.0087 | 999.961 | 0.0029941 | 0.12995 | 286.583 | 299.896
Populus deltoides | LI1C | 2 | Standard | 158.4087 | 9.9789 | 24.9925 | 999.939 | 0.0030142 | 0.13079 | 239.287 | 249.988
Populus deltoides | LI1C | 2 | Standard | 130.5502 | 7.8066 | 24.9976 | 1000.24 | 0.0031431 | 0.13618 | 191.592 | 200.001
Populus deltoides | LI1C | 2 | Standard | 104.8614 | 5.2929 | 24.9981 | 1000.24 | 0.0033365 | 0.14501 | 144.228 | 150.003
Populus deltoides | LI1C | 2 | Standard | 76.5905 | 2.7927 | 25.0058 | 999.872 | 0.003505 | 0.15237 | 96.9155 | 100.048
Populus deltoides | LI1C | 2 | Standard | 49.8244 | -0.15444 | 24.9937 | 1000.24 | 0.0037038 | 0.16139 | 50.0131 | 50.0439
Populus deltoides | LI1C | 2 | Standard | 255.8539 | 20.9765 | 24.9774 | 999.895 | 0.0042318 | 0.18336 | 377.567 | 400.146
Populus deltoides | LI1C | 2 | Standard | 321.9239 | 27.5816 | 24.9953 | 999.994 | 0.0045362 | 0.19799 | 470.305 | 500.022
Populus deltoides | LI1C | 2 | Standard | 398.5303 | 33.4704 | 24.9845 | 1000.03 | 0.0049615 | 0.21698 | 563.786 | 600.053
Populus deltoides | LI1C | 2 | Standard | 484.2448 | 37.8515 | 24.988 | 1000.05 | 0.0054031 | 0.2348 | 658.598 | 700.013
Populus deltoides | LI1C | 2 | Standard | 575.5659 | 40.5003 | 25.0137 | 1000.06 | 0.0056503 | 0.24632 | 755.254 | 800.016
Populus deltoides | LI1C | 2 | Standard | 671.755 | 42.1075 | 24.9922 | 999.734 | 0.0058659 | 0.25699 | 853.005 | 900.118
Populus deltoides | LI1C | 2 | Standard | 770.6589 | 42.8605 | 25.0137 | 1000.03 | 0.0060398 | 0.2658 | 951.475 | 1000.08
Populus deltoides | LI1C | 2 | Standard | 867.3255 | 43.3419 | 25.0076 | 1000.02 | 0.0061729 | 0.26945 | 1050.14 | 1099.97
Populus deltoides | LI1C | 2 | Standard | 964.4652 | 43.6461 | 25.0129 | 999.993 | 0.0062313 | 0.2718 | 1149.28 | 1200.09
Populus deltoides | LI1C | 2 | Standard | 1064.2278 | 43.7173 | 25.001 | 999.985 | 0.0063461 | 0.27708 | 1248.49 | 1300.13
Populus deltoides | LI1C | 2 | Standard | 1159.9514 | 43.8398 | 25.022 | 999.95 | 0.0062708 | 0.27522 | 1347.81 | 1400.1
Populus deltoides | LI1C | 2 | Standard | 1258.4253 | 43.8125 | 25.0006 | 999.953 | 0.0063465 | 0.27786 | 1447 | 1499.99
Populus deltoides | LI1C | 2 | Standard | 1551.2823 | 43.6581 | 24.9812 | 1000.24 | 0.006393 | 0.27936 | 1745.28 | 1800.1
Populus deltoides | LI1C | 2 | Standard | 1745.8343 | 43.7665 | 25.0163 | 999.884 | 0.0064442 | 0.28118 | 1943.77 | 2000.06
Populus deltoides | LI1C | 1 | Standard | 302.0312 | 28.1344 | 25.0538 | 999.866 | 0.010941 | 0.48249 | 367.938 | 400.094
Populus deltoides | LI1C | 1 | Standard | 262.7607 | 24.7506 | 25.0547 | 1000.11 | 0.01062 | 0.47144 | 321.845 | 350.015
Populus deltoides | LI1C | 1 | Standard | 224.4977 | 21.0973 | 25.0618 | 999.78 | 0.010414 | 0.46 | 276.027 | 299.999
Populus deltoides | LI1C | 1 | Standard | 188.2362 | 16.9416 | 25.039 | 1000.02 | 0.010178 | 0.44882 | 230.734 | 250.027
Populus deltoides | LI1C | 1 | Standard | 152.1268 | 12.8554 | 25.0379 | 1000.01 | 0.0098885 | 0.43727 | 185.342 | 200.031
Populus deltoides | LI1C | 1 | Standard | 117.6924 | 8.4199 | 25.0391 | 999.967 | 0.0097397 | 0.4304 | 140.173 | 149.956
Populus deltoides | LI1C | 1 | Standard | 83.4268 | 4.0828 | 25.0173 | 999.935 | 0.0096451 | 0.42537 | 95.0484 | 100.049
Populus deltoides | LI1C | 1 | Standard | 50.4513 | -0.60705 | 25.0377 | 999.944 | 0.009499 | 0.42087 | 50.1441 | 50.0134
Populus deltoides | LI1C | 1 | Standard | 299.0183 | 26.527 | 24.9982 | 999.987 | 0.0095381 | 0.41884 | 369.97 | 400.028
Populus deltoides | LI1C | 1 | Standard | 377.1268 | 32.0445 | 25.0671 | 1000.05 | 0.0094949 | 0.41718 | 463.506 | 499.95
Populus deltoides | LI1C | 1 | Standard | 460.8224 | 35.8817 | 25.0543 | 999.765 | 0.0094102 | 0.41471 | 558.914 | 600.054
Populus deltoides | LI1C | 1 | Standard | 548.1648 | 37.9289 | 25.0355 | 1000.07 | 0.0091597 | 0.40254 | 656.09 | 700.036
Populus deltoides | LI1C | 1 | Standard | 639.0705 | 38.9417 | 25.0328 | 1000.07 | 0.0089547 | 0.39171 | 754.414 | 800.107
Populus deltoides | LI1C | 1 | Standard | 731.6244 | 39.4044 | 25.0454 | 1000.05 | 0.0086537 | 0.38038 | 853.246 | 900.028
Populus deltoides | LI1C | 1 | Standard | 824.9145 | 39.741 | 25.0112 | 1000.05 | 0.0084537 | 0.37073 | 952.375 | 1000.16
Populus deltoides | LI1C | 1 | Standard | 919.1492 | 39.6938 | 25.0395 | 1000.03 | 0.0082858 | 0.36109 | 1051.69 | 1100.1
Populus deltoides | LI1C | 1 | Standard | 1012.3123 | 39.9323 | 25.0175 | 1000.03 | 0.0080448 | 0.35092 | 1150.9 | 1200.09
Populus deltoides | LI1C | 1 | Standard | 1106.8044 | 39.9482 | 25.0475 | 1000.03 | 0.0078446 | 0.34294 | 1250.25 | 1300
Populus deltoides | LI1C | 1 | Standard | 1200.6824 | 40.1779 | 25.0175 | 1000.02 | 0.007662 | 0.33538 | 1349.61 | 1400.14
Populus deltoides | LI1C | 1 | Standard | 1293.4232 | 40.3833 | 25.0344 | 1000.03 | 0.0074804 | 0.32587 | 1448.82 | 1500.04
Populus deltoides | LI1C | 1 | Standard | 1578.5527 | 40.7884 | 25.0111 | 1000.04 | 0.007183 | 0.31334 | 1746.84 | 1800.17
Populus deltoides | LI1C | 1 | Standard | 1767.3601 | 40.9902 | 25.0084 | 1000.06 | 0.0069944 | 0.30292 | 1945.54 | 2000.14
Populus deltoides | LI2A | 2 | Standard | 273.3268 | 26.4016 | 24.9697 | 1000.1 | 0.0066099 | 0.29169 | 371.142 | 399.997
Populus deltoides | LI2A | 2 | Standard | 235.7721 | 23.716 | 24.9961 | 999.809 | 0.0065766 | 0.28911 | 324.173 | 350.02
Populus deltoides | LI2A | 2 | Standard | 201.453 | 20.4462 | 24.9987 | 999.75 | 0.0065886 | 0.28837 | 277.831 | 300.109
Populus deltoides | LI2A | 2 | Standard | 169.7498 | 16.6401 | 24.9936 | 1000.02 | 0.0065953 | 0.2894 | 231.824 | 249.992
Populus deltoides | LI2A | 2 | Standard | 138.5224 | 12.7094 | 25.0065 | 1000.02 | 0.0066092 | 0.28997 | 186.051 | 199.99
Populus deltoides | LI2A | 2 | Standard | 108.2003 | 8.6211 | 24.9916 | 999.721 | 0.0066667 | 0.2931 | 140.441 | 149.998
Populus deltoides | LI2A | 2 | Standard | 77.8057 | 4.4967 | 24.9926 | 1000.03 | 0.0067267 | 0.29652 | 94.9298 | 100.065
Populus deltoides | LI2A | 2 | Standard | 47.2553 | 0.30294 | 25 | 1000.01 | 0.0067736 | 0.29762 | 49.3728 | 50.0102
Populus deltoides | LI2A | 2 | Standard | 270.2121 | 28.1224 | 24.9792 | 999.911 | 0.0069984 | 0.30646 | 369.28 | 399.986
Populus deltoides | LI2A | 2 | Standard | 344.304 | 34.2075 | 24.985 | 1000.01 | 0.0071606 | 0.31383 | 462.507 | 500.022
Populus deltoides | LI2A | 2 | Standard | 420.6647 | 39.343 | 24.9891 | 999.902 | 0.0071539 | 0.315 | 556.662 | 599.988
Populus deltoides | LI2A | 2 | Standard | 505.1834 | 42.629 | 24.9941 | 999.819 | 0.0072128 | 0.31747 | 652.612 | 699.95
Populus deltoides | LI2A | 2 | Standard | 594.7536 | 44.6512 | 25.0068 | 999.852 | 0.0072635 | 0.31917 | 749.951 | 800.049
Populus deltoides | LI2A | 2 | Standard | 686.064 | 45.8097 | 24.9948 | 1000.12 | 0.007239 | 0.31697 | 848.109 | 900.054
Populus deltoides | LI2A | 2 | Standard | 779.9519 | 46.3202 | 25.0197 | 999.74 | 0.0072167 | 0.31486 | 946.854 | 1000.01
Populus deltoides | LI2A | 2 | Standard | 873.0131 | 46.8583 | 25.0092 | 999.8 | 0.0071329 | 0.31099 | 1045.69 | 1100.01
Populus deltoides | LI2A | 2 | Standard | 966.3394 | 47.3002 | 25.0198 | 1000.19 | 0.0070585 | 0.3071 | 1144.62 | 1200
Populus deltoides | LI2A | 2 | Standard | 1059.0067 | 47.7844 | 25.0147 | 999.842 | 0.006929 | 0.30203 | 1243.63 | 1300.03
Populus deltoides | LI2A | 2 | Standard | 1152.044 | 48.2087 | 25.0043 | 999.773 | 0.006814 | 0.2975 | 1342.66 | 1400.02
Populus deltoides | LI2A | 2 | Standard | 1246.7696 | 48.1802 | 25.0083 | 1000.12 | 0.0067147 | 0.29283 | 1442.13 | 1500
Populus deltoides | LI2A | 2 | Standard | 1536.6271 | 48.3049 | 24.9979 | 999.744 | 0.0066683 | 0.291 | 1740.17 | 1800.08
Populus deltoides | LI2A | 2 | Standard | 1729.8205 | 48.2289 | 24.994 | 999.769 | 0.0066089 | 0.28859 | 1938.95 | 1999.99
Populus deltoides | LI2A | 1 | Standard | 281.8584 | 31.2663 | 25.0373 | 1000.22 | 0.0091186 | 0.40972 | 365.373 | 399.969
Populus deltoides | LI2A | 1 | Standard | 246.3636 | 27.7603 | 24.9855 | 999.88 | 0.0092785 | 0.41661 | 319.296 | 350.019
Populus deltoides | LI2A | 1 | Standard | 211.5585 | 23.739 | 25.0246 | 1000.19 | 0.0093276 | 0.41828 | 273.723 | 300.016
Populus deltoides | LI2A | 1 | Standard | 178.0149 | 19.3928 | 25.0125 | 1000.18 | 0.0093981 | 0.42207 | 228.487 | 250.027
Populus deltoides | LI2A | 1 | Standard | 145.354 | 14.7407 | 25.0303 | 999.835 | 0.0094922 | 0.42666 | 183.562 | 200.045
Populus deltoides | LI2A | 1 | Standard | 112.7508 | 10.0011 | 24.997 | 1000.07 | 0.009617 | 0.43214 | 138.692 | 150.027
Populus deltoides | LI2A | 1 | Standard | 81.0139 | 4.8929 | 25.0253 | 999.783 | 0.0096839 | 0.43558 | 94.1944 | 100
Populus deltoides | LI2A | 1 | Standard | 48.8682 | -0.15202 | 25.0262 | 1000.06 | 0.0097915 | 0.44025 | 49.6181 | 49.9519
Populus deltoides | LI2A | 1 | Standard | 292.8559 | 29.7355 | 25.0146 | 999.769 | 0.010043 | 0.44798 | 366.626 | 400.043
Populus deltoides | LI2A | 1 | Standard | 368.685 | 36.6996 | 25.0079 | 999.88 | 0.010109 | 0.45401 | 458.732 | 500.07
Populus deltoides | LI2A | 1 | Standard | 450.1805 | 41.7757 | 25.0344 | 999.901 | 0.010146 | 0.45738 | 552.64 | 600.022
Populus deltoides | LI2A | 1 | Standard | 538.327 | 44.6638 | 25.0011 | 999.929 | 0.010212 | 0.4595 | 648.718 | 700.004
Populus deltoides | LI2A | 1 | Standard | 629.6476 | 46.5331 | 25.0481 | 1000.26 | 0.010263 | 0.46113 | 745.866 | 800.054
Populus deltoides | LI2A | 1 | Standard | 722.4493 | 48.1194 | 25.0165 | 999.867 | 0.010317 | 0.46526 | 843.233 | 900.048
Populus deltoides | LI2A | 1 | Standard | 816.5645 | 48.8189 | 24.9769 | 999.963 | 0.010399 | 0.46436 | 941.38 | 999.99
Populus deltoides | LI2A | 1 | Standard | 911.6345 | 49.3955 | 25.031 | 1000.27 | 0.01034 | 0.46394 | 1039.85 | 1100
Populus deltoides | LI2A | 1 | Standard | 1007.8387 | 49.6719 | 25.0064 | 1000.3 | 0.010344 | 0.46494 | 1138.55 | 1200
Populus deltoides | LI2A | 1 | Standard | 1104.287 | 50.0476 | 24.9948 | 999.97 | 0.010282 | 0.46669 | 1237.32 | 1300.09
Populus deltoides | LI2A | 1 | Standard | 1199.9268 | 50.2118 | 24.9858 | 999.94 | 0.010336 | 0.46532 | 1336 | 1400.03
Populus deltoides | LI2A | 1 | Standard | 1296.3673 | 50.2406 | 24.9995 | 999.946 | 0.010338 | 0.46444 | 1434.94 | 1500.01
Populus deltoides | LI2A | 1 | Standard | 1586.5065 | 50.2244 | 25.0021 | 1000.11 | 0.010383 | 0.46386 | 1731.92 | 1800.12
Populus deltoides | LI2A | 1 | Standard | 1778.5004 | 50.6367 | 24.9865 | 999.882 | 0.010298 | 0.46209 | 1929.4 | 1999.91
Populus deltoides | LI2A | 3 | Standard | 262.5489 | 28.5852 | 24.9557 | 999.902 | 0.0065292 | 0.2877 | 369.074 | 400.068
Populus deltoides | LI2A | 3 | Standard | 229.423 | 25.1906 | 25.0017 | 999.953 | 0.0066523 | 0.28985 | 322.667 | 350.005
Populus deltoides | LI2A | 3 | Standard | 198.3366 | 21.52 | 24.9871 | 999.867 | 0.0067456 | 0.29537 | 276.617 | 300.003
Populus deltoides | LI2A | 3 | Standard | 167.4392 | 17.6303 | 25.0016 | 999.857 | 0.0067935 | 0.29963 | 230.794 | 249.993
Populus deltoides | LI2A | 3 | Standard | 137.4119 | 13.5082 | 24.9964 | 1000.11 | 0.0069416 | 0.30599 | 185.217 | 200.011
Populus deltoides | LI2A | 3 | Standard | 107.0542 | 9.2455 | 24.9924 | 999.741 | 0.0070063 | 0.30893 | 139.781 | 150.006
Populus deltoides | LI2A | 3 | Standard | 77.1443 | 4.8501 | 24.9923 | 1000.13 | 0.0071213 | 0.31476 | 94.4951 | 100.018
Populus deltoides | LI2A | 3 | Standard | 47.0973 | 0.35123 | 24.9879 | 1000.05 | 0.0072548 | 0.32059 | 49.2834 | 49.9922
Populus deltoides | LI2A | 3 | Standard | 273.5892 | 29.0022 | 24.9423 | 1000.03 | 0.0075029 | 0.33178 | 368.262 | 400.029
Populus deltoides | LI2A | 3 | Standard | 350.556 | 34.6218 | 25.0041 | 1000.1 | 0.0077471 | 0.33952 | 461.796 | 499.994
Populus deltoides | LI2A | 3 | Standard | 430.289 | 39.6227 | 24.9813 | 999.812 | 0.0078143 | 0.34571 | 556.048 | 600.015
Populus deltoides | LI2A | 3 | Standard | 516.6658 | 42.6561 | 24.9987 | 1000.04 | 0.0078976 | 0.34873 | 652.22 | 700.026
Populus deltoides | LI2A | 3 | Standard | 607.7068 | 44.5174 | 24.9886 | 999.88 | 0.0079559 | 0.35184 | 749.581 | 800.057
Populus deltoides | LI2A | 3 | Standard | 701.2796 | 45.5818 | 24.9929 | 999.797 | 0.0080364 | 0.35399 | 847.627 | 900.012
Populus deltoides | LI2A | 3 | Standard | 795.7571 | 46.2063 | 24.9945 | 1000.12 | 0.0080285 | 0.3537 | 946.164 | 999.961
Populus deltoides | LI2A | 3 | Standard | 890.2416 | 46.7752 | 25.0043 | 999.85 | 0.0080272 | 0.35268 | 1044.89 | 1100.05
Populus deltoides | LI2A | 3 | Standard | 984.6185 | 47.2121 | 25.0002 | 999.767 | 0.0079644 | 0.34988 | 1143.78 | 1200.1
Populus deltoides | LI2A | 3 | Standard | 1079.5794 | 47.4117 | 24.9979 | 999.859 | 0.0079202 | 0.34716 | 1242.64 | 1299.89
Populus deltoides | LI2A | 3 | Standard | 1173.5456 | 47.7746 | 25.0234 | 999.758 | 0.0078353 | 0.34271 | 1341.7 | 1399.98
Populus deltoides | LI2A | 3 | Standard | 1268.9693 | 47.7643 | 25.0096 | 999.743 | 0.0077353 | 0.33832 | 1441.13 | 1500.04
Populus deltoides | LI2A | 3 | Standard | 1558.1513 | 47.7313 | 25.011 | 999.756 | 0.0076305 | 0.33347 | 1739.01 | 1800.01
Populus deltoides | LI2A | 3 | Standard | 1750.3912 | 47.8262 | 25.0155 | 1000.07 | 0.0075547 | 0.33001 | 1937.53 | 1999.99
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 347.094 | 30.7394 | 25.0543 | 1000.04 | 0.0073164 | 0.37571 | 436.541 | 461.783
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 343.9913 | 30.6987 | 25.0632 | 1000.05 | 0.0073152 | 0.37486 | 433.47 | 458.654
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 341.7764 | 30.3236 | 25.072 | 999.99 | 0.0073123 | 0.3739 | 430.427 | 455.223
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 339.5318 | 30.0296 | 25.0671 | 1000.06 | 0.0073121 | 0.37406 | 427.306 | 451.79
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 336.4117 | 29.9833 | 25.0725 | 1000.04 | 0.007309 | 0.3733 | 424.178 | 448.598
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 333.3455 | 29.8609 | 25.0858 | 1000.07 | 0.0073056 | 0.37197 | 421.032 | 445.315
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 330.6471 | 29.6704 | 25.0924 | 999.724 | 0.0073054 | 0.37134 | 417.911 | 441.986
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 328.0914 | 29.4797 | 25.0912 | 1000.06 | 0.0073063 | 0.37137 | 414.779 | 438.647
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 325.7228 | 29.2127 | 25.086 | 1000.01 | 0.007304 | 0.37143 | 411.621 | 435.206
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 322.8379 | 29.121 | 25.0837 | 999.995 | 0.0073017 | 0.37129 | 408.461 | 431.935
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 319.6089 | 29.1231 | 25.0905 | 1000.1 | 0.007302 | 0.37072 | 405.307 | 428.766
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 316.8147 | 28.9903 | 25.095 | 1000.06 | 0.0073032 | 0.37038 | 402.175 | 425.49
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 314.1666 | 28.8004 | 25.1013 | 1000.01 | 0.0073047 | 0.36991 | 399.066 | 422.174
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 311.6263 | 28.5629 | 25.1076 | 1000 | 0.0073041 | 0.36927 | 395.974 | 418.826
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 309.1006 | 28.3297 | 25.108 | 1000 | 0.0073018 | 0.36896 | 392.829 | 415.434
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 306.1043 | 28.2856 | 25.1058 | 1000.1 | 0.0072997 | 0.36884 | 389.678 | 412.22
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 303.3644 | 28.1813 | 25.1025 | 1000.03 | 0.0072995 | 0.36889 | 386.586 | 409.006
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 301.0602 | 27.835 | 25.1087 | 999.732 | 0.0072948 | 0.36795 | 383.494 | 405.552
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 298.5555 | 27.5574 | 25.1172 | 1000.14 | 0.0072934 | 0.36715 | 380.356 | 402.12
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 295.4684 | 27.5071 | 25.1215 | 999.999 | 0.0072914 | 0.3666 | 377.191 | 398.887
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 292.7965 | 27.3293 | 25.1215 | 1000.03 | 0.0072881 | 0.36626 | 374.049 | 395.552
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 290.2783 | 27.1072 | 25.1236 | 1000.02 | 0.0072849 | 0.36576 | 370.976 | 392.241
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 287.4353 | 26.9783 | 25.1272 | 999.744 | 0.0072809 | 0.36508 | 367.867 | 388.983
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 285.223 | 26.6524 | 25.1229 | 1000.05 | 0.0072761 | 0.36488 | 364.749 | 385.524
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 282.5299 | 26.4912 | 25.1226 | 1000.08 | 0.0072754 | 0.36468 | 361.597 | 382.192
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 279.476 | 26.4452 | 25.1243 | 999.7 | 0.0072709 | 0.3641 | 358.477 | 379.009
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 277.0253 | 26.222 | 25.1196 | 1000.09 | 0.0072642 | 0.36379 | 355.419 | 375.712
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 274.605 | 25.9904 | 25.114 | 1000.02 | 0.0072572 | 0.36351 | 352.361 | 372.404
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 271.9829 | 25.8132 | 25.1085 | 1000.07 | 0.0072492 | 0.36311 | 349.273 | 369.119
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 269.6031 | 25.5401 | 25.1076 | 1000.04 | 0.0072438 | 0.36265 | 346.18 | 365.737
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 266.9198 | 25.3898 | 25.1078 | 1000.06 | 0.00724 | 0.36226 | 343.102 | 362.491
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 264.0577 | 25.2917 | 25.1089 | 1000.04 | 0.0072354 | 0.3617 | 340.024 | 359.296
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 262.0721 | 24.8822 | 25.1086 | 999.991 | 0.0072288 | 0.36108 | 336.979 | 355.827
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 259.6669 | 24.6087 | 25.1071 | 1000.06 | 0.0072226 | 0.36058 | 333.859 | 352.415
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 256.6136 | 24.565 | 25.1066 | 1000.09 | 0.0072173 | 0.3601 | 330.716 | 349.21
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 254.1117 | 24.3428 | 25.1084 | 999.973 | 0.0072141 | 0.35962 | 327.638 | 345.895
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 251.7411 | 24.0656 | 25.1098 | 1000.04 | 0.007208 | 0.35891 | 324.581 | 342.543
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 249.1727 | 23.8957 | 25.097 | 1000.13 | 0.0071994 | 0.35896 | 321.464 | 339.235
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 246.915 | 23.6068 | 25.0902 | 1000.04 | 0.0071935 | 0.35886 | 318.364 | 335.832
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 244.2394 | 23.4253 | 25.0958 | 1000.02 | 0.0071898 | 0.35807 | 315.278 | 332.547
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 241.3845 | 23.3374 | 25.0917 | 999.69 | 0.0071814 | 0.35757 | 312.21 | 329.372
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 239.0957 | 23.0833 | 25.0821 | 1000.06 | 0.0071728 | 0.35744 | 309.177 | 326.069
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 236.765 | 22.8253 | 25.077 | 1000.05 | 0.007164 | 0.35701 | 306.147 | 322.763
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 234.0863 | 22.6774 | 25.0732 | 1000.01 | 0.0071575 | 0.35665 | 303.061 | 319.511
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 231.8027 | 22.3893 | 25.0709 | 1000 | 0.0071521 | 0.35626 | 299.987 | 316.133
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 229.461 | 22.1465 | 25.0649 | 1000.06 | 0.0071453 | 0.35599 | 296.955 | 312.84
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 226.8508 | 21.9758 | 25.0611 | 1000.03 | 0.0071359 | 0.35543 | 293.908 | 309.605
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 224.5695 | 21.6467 | 25.0651 | 1000.09 | 0.0071271 | 0.35438 | 290.835 | 306.188
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 222.2295 | 21.3541 | 25.0675 | 1000.04 | 0.0071218 | 0.3537 | 287.737 | 302.779
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 219.4988 | 21.2521 | 25.0592 | 1000.02 | 0.0071178 | 0.35381 | 284.636 | 299.558
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 217.2112 | 21.0173 | 25.0465 | 1000.05 | 0.0071114 | 0.35404 | 281.585 | 296.257
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 214.8373 | 20.7594 | 25.045 | 999.995 | 0.0071034 | 0.3534 | 278.538 | 292.934
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 212.3121 | 20.552 | 25.0429 | 999.715 | 0.0070957 | 0.35284 | 275.465 | 289.636
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 210.2709 | 20.1931 | 25.0372 | 1000.03 | 0.0070882 | 0.35259 | 272.394 | 286.192
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 207.6815 | 20.0013 | 25.0378 | 1000.06 | 0.0070815 | 0.35191 | 269.317 | 282.905
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 204.8831 | 19.9222 | 25.0329 | 1000.01 | 0.0070773 | 0.35176 | 266.259 | 279.751
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 202.4635 | 19.6893 | 25.0313 | 1000.01 | 0.0070709 | 0.35127 | 263.202 | 276.446
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 200.2081 | 19.3959 | 25.0294 | 999.983 | 0.0070607 | 0.3505 | 260.183 | 273.116
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 197.9083 | 19.1554 | 25.0198 | 999.68 | 0.0070527 | 0.35041 | 257.153 | 269.828
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 195.7017 | 18.8602 | 25.0162 | 999.985 | 0.0070463 | 0.35007 | 254.105 | 266.47
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 193.3107 | 18.6297 | 25.0136 | 999.992 | 0.0070408 | 0.34966 | 251.065 | 263.183
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 190.6714 | 18.4909 | 25.0089 | 1000.05 | 0.0070333 | 0.34928 | 248.028 | 259.99
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 188.6678 | 18.1361 | 25.0035 | 1000.04 | 0.0070261 | 0.34902 | 244.995 | 256.587
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 186.6358 | 17.7763 | 25.0004 | 1000.04 | 0.0070186 | 0.34855 | 241.951 | 253.166
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 183.9187 | 17.6231 | 25.003 | 1000.03 | 0.0070106 | 0.34764 | 238.875 | 249.92
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 181.4424 | 17.4007 | 25.0051 | 999.993 | 0.0070022 | 0.34677 | 235.835 | 246.643
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 179.3588 | 17.0925 | 25.0034 | 1000.02 | 0.0069992 | 0.34653 | 232.847 | 243.33
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 177.0352 | 16.8364 | 25.0017 | 999.963 | 0.0069934 | 0.34611 | 229.79 | 240.001
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 174.9972 | 16.468 | 25.0004 | 999.984 | 0.0069847 | 0.34545 | 226.729 | 236.557
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 172.5044 | 16.2649 | 25.0014 | 999.997 | 0.0069798 | 0.34487 | 223.676 | 233.284
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 169.675 | 16.1795 | 25.0015 | 1000.04 | 0.0069749 | 0.34441 | 220.604 | 230.111
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 167.4892 | 15.8963 | 24.999 | 1000.05 | 0.0069686 | 0.34396 | 217.606 | 226.815
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 165.3456 | 15.6005 | 24.9964 | 999.699 | 0.0069603 | 0.34337 | 214.63 | 223.527
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 163.1269 | 15.332 | 24.9937 | 999.997 | 0.0069528 | 0.34286 | 211.645 | 220.257
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 161.0091 | 15.0205 | 24.9929 | 1000 | 0.0069467 | 0.34238 | 208.63 | 216.916
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 158.9431 | 14.7248 | 24.9927 | 1000.03 | 0.00695 | 0.34257 | 205.624 | 213.601
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 156.5905 | 14.5532 | 24.9964 | 999.989 | 0.0069676 | 0.34354 | 202.595 | 210.389
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 154.7527 | 14.1751 | 25.0101 | 999.968 | 0.0069806 | 0.34362 | 199.605 | 207.011
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 152.746 | 13.8252 | 25.0236 | 999.995 | 0.0069883 | 0.34327 | 196.577 | 203.62
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 150.0894 | 13.6696 | 25.0396 | 999.977 | 0.0069927 | 0.34247 | 193.512 | 200.385
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 147.8251 | 13.3649 | 25.0603 | 1000.04 | 0.0069948 | 0.34124 | 190.458 | 197.013
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 145.622 | 13.057 | 25.0756 | 999.998 | 0.0069956 | 0.34028 | 187.417 | 193.65
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 143.1619 | 12.8354 | 25.0939 | 1000.08 | 0.0069984 | 0.33922 | 184.378 | 190.375
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 141.1412 | 12.4498 | 25.1181 | 999.96 | 0.0070027 | 0.33794 | 181.322 | 186.92
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 138.6424 | 12.2418 | 25.1376 | 999.977 | 0.0070086 | 0.33712 | 178.25 | 183.626
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 135.8857 | 12.1507 | 25.1474 | 1000.05 | 0.007014 | 0.33683 | 175.198 | 180.469
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 133.822 | 11.844 | 25.1543 | 999.979 | 0.0070155 | 0.33645 | 172.216 | 177.167
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 131.7367 | 11.5298 | 25.1664 | 999.947 | 0.0070178 | 0.33579 | 169.218 | 173.841
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 129.4564 | 11.2713 | 25.1788 | 1000.04 | 0.0070197 | 0.33513 | 166.185 | 170.535
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 127.4199 | 10.9476 | 25.1836 | 1000.03 | 0.0070202 | 0.33478 | 163.167 | 167.179
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 125.1464 | 10.703 | 25.1882 | 999.974 | 0.0070195 | 0.33434 | 160.154 | 163.907
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 122.6868 | 10.5122 | 25.1976 | 1000.04 | 0.0070195 | 0.33368 | 157.134 | 160.682
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 120.8448 | 10.119 | 25.2058 | 1000.02 | 0.0070215 | 0.33328 | 154.103 | 157.244
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 118.8307 | 9.7798 | 25.2078 | 1000.03 | 0.0070212 | 0.33304 | 151.041 | 153.828
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 116.393 | 9.5908 | 25.2081 | 999.986 | 0.0070185 | 0.33273 | 148.006 | 150.589
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 114.0981 | 9.3501 | 25.2139 | 1000.04 | 0.0070184 | 0.3323 | 144.972 | 147.3
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 111.8728 | 9.097 | 25.2122 | 1000 | 0.0070162 | 0.33215 | 141.942 | 144.002
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 109.7016 | 8.8309 | 25.2074 | 1000.01 | 0.0070125 | 0.33207 | 138.92 | 140.699
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 107.9482 | 8.4137 | 25.2065 | 1000.02 | 0.0070059 | 0.33163 | 135.895 | 137.242
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 105.6736 | 8.1656 | 25.2091 | 999.986 | 0.0070036 | 0.33121 | 132.849 | 133.933
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 103.0872 | 8.0188 | 25.2098 | 999.964 | 0.0070009 | 0.33086 | 129.787 | 130.709
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 100.8925 | 7.7575 | 25.2093 | 999.972 | 0.0069969 | 0.33051 | 126.774 | 127.42
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 98.8208 | 7.4659 | 25.212 | 999.982 | 0.0069937 | 0.33004 | 123.801 | 124.141
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 96.8541 | 7.1446 | 25.2051 | 1000 | 0.0069879 | 0.32995 | 120.813 | 120.817
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 94.8455 | 6.8128 | 25.2039 | 999.964 | 0.0069799 | 0.32934 | 117.785 | 117.442
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 92.5079 | 6.5877 | 25.2089 | 999.916 | 0.006975 | 0.32858 | 114.757 | 114.174
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 89.9821 | 6.433 | 25.2041 | 1000.03 | 0.0069674 | 0.32822 | 111.723 | 110.97
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 88.2822 | 6.0206 | 25.1985 | 1000.02 | 0.0069605 | 0.32797 | 108.73 | 107.55
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 86.4311 | 5.6445 | 25.1961 | 1000.03 | 0.006952 | 0.32741 | 105.711 | 104.14
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 83.8881 | 5.4857 | 25.1919 | 999.958 | 0.0069462 | 0.32711 | 102.639 | 100.894
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 81.7643 | 5.2034 | 25.1887 | 1000.01 | 0.0069395 | 0.32674 | 99.6156 | 97.5735
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 79.7101 | 4.9052 | 25.1842 | 1000 | 0.0069311 | 0.32632 | 96.6148 | 94.2597
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 77.6169 | 4.6177 | 25.1775 | 1000.01 | 0.0069254 | 0.32621 | 93.5899 | 90.9327
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 75.8537 | 4.2134 | 25.165 | 1000 | 0.0069169 | 0.32625 | 90.5324 | 87.4561
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 73.5585 | 3.9778 | 25.1579 | 999.954 | 0.0069091 | 0.32598 | 87.4677 | 84.1411
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 71.0467 | 3.8201 | 25.1605 | 1000.01 | 0.0069002 | 0.32507 | 84.4484 | 80.9489
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 69.0182 | 3.5101 | 25.1566 | 1000.02 | 0.0068878 | 0.32433 | 81.4385 | 77.6138
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 67.0076 | 3.2026 | 25.1494 | 1000.03 | 0.0068786 | 0.32399 | 78.4365 | 74.2901
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 64.8831 | 2.9287 | 25.1453 | 999.958 | 0.0068699 | 0.32349 | 75.4263 | 70.9919
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 63.0914 | 2.5438 | 25.1418 | 999.977 | 0.0068592 | 0.32285 | 72.41 | 67.5754
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 60.9249 | 2.2788 | 25.1344 | 1000 | 0.0068513 | 0.3226 | 69.3722 | 64.2584
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 58.4787 | 2.1069 | 25.122 | 1000.05 | 0.006841 | 0.32251 | 66.3353 | 61.0353
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 56.7485 | 1.7122 | 25.1099 | 1000.05 | 0.0068283 | 0.32222 | 63.3347 | 57.6253
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 55.0615 | 1.2882 | 25.1075 | 999.969 | 0.0068184 | 0.32155 | 60.2909 | 54.1433
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 52.6049 | 1.1041 | 25.1047 | 999.927 | 0.0068101 | 0.321 | 57.2092 | 50.8631
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 50.2917 | 0.88139 | 25.1002 | 999.99 | 0.0068008 | 0.3205 | 54.1498 | 47.566
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 48.385 | 0.53933 | 25.0947 | 1000 | 0.0067934 | 0.32021 | 51.1249 | 44.185
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 46.214 | 0.27289 | 25.0825 | 1000 | 0.006788 | 0.32041 | 48.0643 | 40.8444
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 44.3825 | -0.10697 | 25.0719 | 1000 | 0.0067766 | 0.32011 | 44.9944 | 37.3794
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 42.1569 | -0.36048 | 25.0676 | 999.967 | 0.0067657 | 0.31945 | 41.9188 | 34.0366
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 39.6839 | -0.52645 | 25.0648 | 999.983 | 0.0067568 | 0.31887 | 38.8651 | 30.803
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 37.7423 | -0.85467 | 25.0532 | 999.945 | 0.006746 | 0.31869 | 35.8392 | 27.4346
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 35.6454 | -1.1365 | 25.04 | 999.957 | 0.0067375 | 0.31874 | 32.8031 | 24.1028
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 33.5354 | -1.4121 | 25.0335 | 999.992 | 0.0067303 | 0.31848 | 29.7705 | 20.7819
Populus deltoides | LI2A |  | RACiR 500 to 0 no VPD control | 31.7043 | -1.7786 | 25.0295 | 999.964 | 0.0067212 | 0.31797 | 26.7283 | 17.3591
Populus deltoides | LI2A |  | RACiR 800 to 300 | 47.0984 | 596.1659 | 1000.04 | 25.0143 | 0.010413 | 0.46213 | 712.828 | 756.793
Populus deltoides | LI2A |  | RACiR 800 to 300 | 47.0168 | 594.907 | 999.653 | 25.0139 | 0.010413 | 0.46218 | 711.349 | 755.219
Populus deltoides | LI2A |  | RACiR 800 to 300 | 47.0914 | 593.009 | 1000.01 | 25.0204 | 0.010411 | 0.46167 | 709.7 | 753.636
Populus deltoides | LI2A |  | RACiR 800 to 300 | 47.049 | 591.5378 | 999.977 | 25.0263 | 0.01041 | 0.46117 | 708.229 | 752.107
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.8392 | 590.3957 | 1000.06 | 25.024 | 0.01041 | 0.46131 | 706.565 | 750.22
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.779 | 589.1903 | 999.996 | 25.0162 | 0.010408 | 0.4618 | 705.074 | 748.661
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.6766 | 587.8849 | 1000.02 | 25.0098 | 0.010407 | 0.46213 | 703.428 | 746.896
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.5026 | 586.8174 | 1000 | 25.0052 | 0.010404 | 0.4623 | 701.91 | 745.196
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.4749 | 585.0891 | 1000.07 | 25.0123 | 0.010403 | 0.46166 | 700.239 | 743.487
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.5136 | 583.4027 | 999.989 | 25.0175 | 0.010403 | 0.46123 | 698.709 | 741.98
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.3945 | 581.9588 | 1000.03 | 25.0172 | 0.010404 | 0.46128 | 696.959 | 740.1
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.4351 | 580.407 | 1000 | 25.0156 | 0.010403 | 0.46142 | 695.425 | 738.599
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.4575 | 578.7811 | 1000.01 | 25.0133 | 0.010405 | 0.46174 | 693.733 | 736.912
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.2871 | 577.7397 | 1000.01 | 25.0112 | 0.010407 | 0.46203 | 692.223 | 735.221
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.2681 | 576.121 | 1000.05 | 25.0112 | 0.010405 | 0.46195 | 690.544 | 733.516
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.2381 | 574.7447 | 1000.07 | 25.0087 | 0.010407 | 0.46228 | 688.991 | 731.92
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.0897 | 573.3714 | 999.956 | 25.0095 | 0.010408 | 0.4623 | 687.258 | 730.028
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.1752 | 571.5528 | 1000.01 | 25.0118 | 0.010406 | 0.46191 | 685.68 | 728.525
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.2937 | 569.6411 | 1000.06 | 25.0102 | 0.010408 | 0.46218 | 683.92 | 726.867
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.1851 | 568.3348 | 1000.01 | 25.0123 | 0.010407 | 0.46191 | 682.411 | 725.239
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.2586 | 566.5148 | 999.689 | 25.011 | 0.010405 | 0.46188 | 680.717 | 723.606
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.3286 | 564.8858 | 1000.08 | 25.0095 | 0.010406 | 0.46205 | 679.163 | 722.108
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.0891 | 563.7008 | 1000.08 | 25.0134 | 0.010406 | 0.46174 | 677.502 | 720.195
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.1454 | 561.9978 | 1000.06 | 25.0143 | 0.010403 | 0.46152 | 675.936 | 718.677
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.1457 | 560.3708 | 1000.07 | 25.0119 | 0.010402 | 0.46172 | 674.221 | 716.947
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.9626 | 559.2401 | 999.696 | 25.0121 | 0.010401 | 0.4616 | 672.691 | 715.222
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.0038 | 557.506 | 1000.07 | 25.0079 | 0.010401 | 0.46193 | 670.927 | 713.488
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.0009 | 556.0894 | 1000.04 | 25.0032 | 0.010401 | 0.46228 | 669.384 | 711.932
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.8686 | 554.6594 | 1000.04 | 25.0039 | 0.0104 | 0.46213 | 667.668 | 710.07
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.9471 | 552.9145 | 1000.06 | 25.0074 | 0.010401 | 0.46192 | 666.108 | 708.579
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.0258 | 551.0106 | 1000.05 | 25.0119 | 0.010401 | 0.46155 | 664.425 | 706.962
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.0184 | 549.4308 | 1000.02 | 25.0167 | 0.010402 | 0.4613 | 662.858 | 705.375
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.0796 | 547.6839 | 1000.04 | 25.0172 | 0.010403 | 0.46141 | 661.179 | 703.748
Populus deltoides | LI2A |  | RACiR 800 to 300 | 46.0778 | 546.2453 | 1000.08 | 25.0171 | 0.010404 | 0.46151 | 659.679 | 702.234
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.917 | 544.9607 | 1000.04 | 25.017 | 0.010405 | 0.46157 | 658 | 700.379
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.8301 | 543.6762 | 1000.05 | 25.0176 | 0.010403 | 0.46142 | 656.529 | 698.815
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.8137 | 542.0646 | 999.766 | 25.0148 | 0.010401 | 0.46149 | 654.827 | 697.081
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.6484 | 540.9491 | 1000.09 | 25.0116 | 0.010399 | 0.46158 | 653.298 | 695.381
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.6553 | 539.2105 | 1000.04 | 25.0131 | 0.010398 | 0.46139 | 651.58 | 693.653
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.723 | 537.5625 | 1000.02 | 25.0115 | 0.010396 | 0.46138 | 650.044 | 692.17
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.6196 | 536.081 | 999.993 | 25.012 | 0.010396 | 0.46126 | 648.329 | 690.337
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.6949 | 534.3864 | 1000.03 | 25.0115 | 0.010393 | 0.46115 | 646.784 | 688.86
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.7499 | 532.6497 | 999.772 | 25.009 | 0.010391 | 0.46119 | 645.118 | 687.234
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.6151 | 531.4362 | 1000.07 | 25.0074 | 0.010388 | 0.46109 | 643.6 | 685.572
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.5415 | 530.0089 | 1000.07 | 25.0052 | 0.010389 | 0.46129 | 641.933 | 683.816
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.4666 | 528.6911 | 1000.04 | 25.0046 | 0.010388 | 0.46125 | 640.43 | 682.221
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.264 | 527.4907 | 999.983 | 25.0028 | 0.010389 | 0.46142 | 638.716 | 680.294
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.2481 | 526.0003 | 1000.12 | 25.0053 | 0.010394 | 0.46149 | 637.148 | 678.705
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.337 | 524.2247 | 1000.08 | 25.0045 | 0.010407 | 0.46208 | 635.399 | 677.04
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.2574 | 522.9303 | 1000.05 | 25.0065 | 0.010419 | 0.4624 | 633.838 | 675.394
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.328 | 521.0761 | 999.734 | 25.008 | 0.010426 | 0.46243 | 632.096 | 673.721
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.3772 | 519.4223 | 999.915 | 25.0063 | 0.010424 | 0.46227 | 630.551 | 672.208
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.2386 | 517.7837 | 1000.05 | 25.0117 | 0.010411 | 0.4611 | 628.839 | 670.339
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.2381 | 516.234 | 1000.1 | 25.0099 | 0.010402 | 0.46079 | 627.317 | 668.8
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.2439 | 514.5549 | 999.773 | 25.0043 | 0.010392 | 0.4607 | 625.62 | 667.088
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.0828 | 513.3242 | 1000.04 | 25.0026 | 0.010382 | 0.46036 | 624.078 | 665.367
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.0863 | 511.6712 | 1000.04 | 24.9964 | 0.01037 | 0.46036 | 622.38 | 663.653
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.0795 | 510.2089 | 1000.02 | 24.9921 | 0.010361 | 0.46041 | 620.847 | 662.1
Populus deltoides | LI2A |  | RACiR 800 to 300 | 44.9057 | 508.8137 | 999.792 | 24.9958 | 0.010352 | 0.45984 | 619.164 | 660.223
Populus deltoides | LI2A |  | RACiR 800 to 300 | 44.9598 | 507.1799 | 1000.04 | 24.9984 | 0.010349 | 0.45972 | 617.64 | 658.746
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.1048 | 505.2072 | 1000.07 | 25.0008 | 0.010353 | 0.45985 | 615.914 | 657.152
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.0547 | 503.8206 | 1000.06 | 25.0026 | 0.010355 | 0.45987 | 614.384 | 655.557
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.1146 | 502.0804 | 1000.1 | 25.0016 | 0.010356 | 0.46004 | 612.695 | 653.921
Populus deltoides | LI2A |  | RACiR 800 to 300 | 45.1149 | 500.6369 | 1000.06 | 25.0043 | 0.010361 | 0.46017 | 611.193 | 652.409
Populus deltoides | LI2A |  | RACiR 800 to 300 | 44.895 | 499.5478 | 1000.01 | 25.0034 | 0.010363 | 0.4604 | 609.542 | 650.529
Populus deltoides | LI2A |  | RACiR 800 to 300 | 44.8021 | 498.2832 | 999.776 | 25.0054 | 0.010365 | 0.46036 | 608.056 | 648.942
Populus deltoides | LI2A |  | RACiR 800 to 300 | 44.779 | 496.64 | 1000.08 | 25.0081 | 0.010365 | 0.46023 | 606.357 | 647.206
Populus deltoides | LI2A |  | RACiR 800 to 300 | 44.6059 | 495.6149 | 1000.04 | 25.0058 | 0.010366 | 0.46048 | 604.869 | 645.534
Populus deltoides | LI2A |  | RACiR 800 to 300 | 44.596 | 494.0351 | 1000 | 25.0008 | 0.010364 | 0.46076 | 603.163 | 643.81
Populus deltoides | LI2A |  | RACiR 800 to 300 | 44.6855 | 492.3449 | 999.786 | 25.001 | 0.010366 | 0.46087 | 601.606 | 642.327
Populus deltoides | LI2A |  | RACiR 800 to 300 | 44.5356 | 491.1901 | 1000.09 | 24.9944 | 0.010367 | 0.46145 | 599.96 | 640.523
Populus deltoides | LI2A |  | RACiR 800 to 300 | 44.6057 | 489.5335 | 1000.07 | 24.9924 | 0.010364 | 0.46146 | 598.412 | 639.036
Populus deltoides | LI2A |  | RACiR 800 to 300 | 44.6545 | 487.7014 | 1000.12 | 24.9978 | 0.010364 | 0.46107 | 596.74 | 637.398
Populus deltoides | LI2A |  | RACiR 800 to 300 | 44.461 | 486.6363 | 999.825 | 24.9983 | 0.010363 | 0.46092 | 595.26 | 635.711
Populus deltoides | LI2A |  | RACiR 800 to 300 | 44.3897 | 485.0295 | 1000.01 | 25.004 | 0.010361 | 0.46042 | 593.578 | 633.948
Populus deltoides | LI2A |  | RACiR 800 to 300 | 44.32 | 483.6503 | 1000.1 | 25.0071 | 0.010363 | 0.46024 | 592.058 | 632.341
Populus deltoides | LI2A |  | RACiR 800 to 300 | 44.0991 | 482.4422 | 999.73 | 25.0072 | 0.010362 | 0.4602 | 590.346 | 630.397
Populus deltoides | LI2A |  | RACiR 800 to 300 | 44.0941 | 480.95 | 1000.1 | 25.0067 | 0.01036 | 0.46015 | 588.819 | 628.857
Populus deltoides | LI2A |  | RACiR 800 to 300 | 44.1333 | 479.2931 | 1000.12 | 25.0011 | 0.01036 | 0.46055 | 587.115 | 627.178
Populus deltoides | LI2A |  | RACiR 800 to 300 | 44.0196 | 478.0926 | 1000.14 | 24.9984 | 0.01036 | 0.46071 | 585.598 | 625.536
Populus deltoides | LI2A |  | RACiR 800 to 300 | 44.0758 | 476.2516 | 1000.09 | 24.9992 | 0.010356 | 0.46047 | 583.892 | 623.878
Populus deltoides | LI2A |  | RACiR 800 to 300 | 44.1035 | 474.7743 | 999.753 | 24.9966 | 0.010359 | 0.46086 | 582.354 | 622.356
Populus deltoides | LI2A |  | RACiR 800 to 300 | 43.9434 | 473.435 | 1000.13 | 24.9994 | 0.01036 | 0.46074 | 580.661 | 620.491
Populus deltoides | LI2A |  | RACiR 800 to 300 | 43.9377 | 471.8721 | 1000.06 | 25.0039 | 0.010361 | 0.46045 | 579.119 | 618.935
Populus deltoides | LI2A |  | RACiR 800 to 300 | 43.9554 | 470.2235 | 999.756 | 25.0018 | 0.010362 | 0.46067 | 577.421 | 617.238
Populus deltoides | LI2A |  | RACiR 800 to 300 | 43.7722 | 469.1024 | 1000.08 | 25.0048 | 0.010362 | 0.46045 | 575.925 | 615.55
Populus deltoides | LI2A |  | RACiR 800 to 300 | 43.7743 | 467.4159 | 999.76 | 25.0035 | 0.01036 | 0.46041 | 574.212 | 613.826
Populus deltoides | LI2A |  | RACiR 800 to 300 | 43.7872 | 465.9441 | 999.878 | 25.0006 | 0.010361 | 0.4607 | 572.668 | 612.283
Populus deltoides | LI2A |  | RACiR 800 to 300 | 43.6293 | 464.6269 | 1000.05 | 25.0011 | 0.010361 | 0.46071 | 570.971 | 610.423
Populus deltoides | LI2A |  | RACiR 800 to 300 | 43.6725 | 463.0224 | 999.759 | 25.0017 | 0.01036 | 0.46065 | 569.439 | 608.921
Populus deltoides | LI2A |  | RACiR 800 to 300 | 43.7815 | 461.0831 | 1000.06 | 25.0051 | 0.010361 | 0.46046 | 567.74 | 607.315
Populus deltoides | LI2A |  | RACiR 800 to 300 | 43.6923 | 459.7321 | 1000.14 | 25.0081 | 0.01036 | 0.46013 | 566.238 | 605.714
Populus deltoides | LI2A |  | RACiR 800 to 300 | 43.7411 | 457.9139 | 999.785 | 25.0119 | 0.010359 | 0.45979 | 564.565 | 604.076
Populus deltoides | LI2A |  | RACiR 800 to 300 | 43.7451 | 456.4402 | 1000.12 | 25.0123 | 0.010359 | 0.45975 | 563.075 | 602.576
Populus deltoides | LI2A |  | RACiR 800 to 300 | 43.538 | 455.2478 | 1000.03 | 25.0147 | 0.010359 | 0.45958 | 561.44 | 600.725
Populus deltoides | LI2A |  | RACiR 800 to 300 | 43.4917 | 453.8569 | 1000.09 | 25.0163 | 0.010358 | 0.45944 | 559.948 | 599.177
Populus deltoides | LI2A |  | RACiR 800 to 300 | 43.4994 | 452.242 | 1000.13 | 25.0158 | 0.01036 | 0.45965 | 558.267 | 597.491
Populus deltoides | LI2A |  | RACiR 800 to 300 | 43.3103 | 451.264 | 999.741 | 25.0133 | 0.010362 | 0.45993 | 556.786 | 595.81
Populus deltoides | LI2A |  | RACiR 800 to 300 | 43.3016 | 449.6845 | 1000.17 | 25.0089 | 0.01036 | 0.46016 | 555.097 | 594.101
Populus deltoides | LI2A |  | RACiR 800 to 300 | 43.3397 | 448.1457 | 999.988 | 25.0062 | 0.01036 | 0.46034 | 553.565 | 592.594
Populus deltoides | LI2A |  | RACiR 800 to 300 | 43.1512 | 446.9054 | 1000.1 | 25.0062 | 0.010357 | 0.46014 | 551.925 | 590.752
Populus deltoides | LI2A |  | RACiR 800 to 300 | 43.1569 | 445.2975 | 1000.15 | 25.011 | 0.010354 | 0.45958 | 550.418 | 589.243
Populus deltoides | LI2A |  | RACiR 800 to 300 | 43.2101 | 443.4476 | 999.739 | 25.0146 | 0.010353 | 0.45926 | 548.717 | 587.581
Populus deltoides | LI2A |  | RACiR 800 to 300 | 43.0198 | 442.4717 | 1000.09 | 25.0134 | 0.010355 | 0.4595 | 547.247 | 585.912
Populus deltoides | LI2A |  | RACiR 800 to 300 | 42.9813 | 440.982 | 999.772 | 25.0103 | 0.010356 | 0.45978 | 545.574 | 584.19
Populus deltoides | LI2A |  | RACiR 800 to 300 | 42.9459 | 439.5986 | 1000.1 | 25.0094 | 0.010358 | 0.45997 | 544.041 | 582.608
Populus deltoides | LI2A |  | RACiR 800 to 300 | 42.7072 | 438.4948 | 999.842 | 25.0071 | 0.010358 | 0.46009 | 542.358 | 580.675
Populus deltoides | LI2A |  | RACiR 800 to 300 | 42.7198 | 436.907 | 1000.15 | 25.0089 | 0.010356 | 0.45984 | 540.817 | 579.136
Populus deltoides | LI2A |  | RACiR 800 to 300 | 42.7669 | 435.1569 | 999.828 | 25.0077 | 0.010356 | 0.45997 | 539.103 | 577.455
Populus deltoides | LI2A |  | RACiR 800 to 300 | 42.6554 | 433.935 | 1000.1 | 25.0066 | 0.010355 | 0.45997 | 537.607 | 575.836
Populus deltoides | LI2A |  | RACiR 800 to 300 | 42.7188 | 432.186 | 1000.16 | 25.0057 | 0.010355 | 0.46011 | 535.924 | 574.208
Populus deltoides | LI2A |  | RACiR 800 to 300 | 42.7053 | 430.748 | 1000.1 | 25.0073 | 0.010358 | 0.46017 | 534.414 | 572.672
Populus deltoides | LI2A |  | RACiR 800 to 300 | 42.488 | 429.5414 | 999.739 | 25.0123 | 0.010358 | 0.45987 | 532.769 | 570.802
Populus deltoides | LI2A |  | RACiR 800 to 300 | 42.4454 | 428.1574 | 1000.11 | 25.0132 | 0.010359 | 0.45989 | 531.257 | 569.239
Populus deltoides | LI2A |  | RACiR 800 to 300 | 42.3922 | 426.5802 | 999.722 | 25.0152 | 0.010358 | 0.45971 | 529.567 | 567.481
Populus deltoides | LI2A |  | RACiR 800 to 300 | 42.2283 | 425.4182 | 1000.13 | 25.0164 | 0.010357 | 0.45949 | 528.067 | 565.807
Populus deltoides | LI2A |  | RACiR 800 to 300 | 42.2262 | 423.7385 | 999.816 | 25.0174 | 0.010356 | 0.45942 | 526.36 | 564.088
Populus deltoides | LI2A |  | RACiR 800 to 300 | 42.2136 | 422.1713 | 1000.09 | 25.0221 | 0.010357 | 0.45902 | 524.818 | 562.519
Populus deltoides | LI2A |  | RACiR 800 to 300 | 42.0585 | 420.8919 | 1000.06 | 25.0214 | 0.010359 | 0.4592 | 523.131 | 560.661
Populus deltoides | LI2A |  | RACiR 800 to 300 | 42.0823 | 419.3691 | 1000.17 | 25.0195 | 0.010357 | 0.45927 | 521.607 | 559.153
Populus deltoides | LI2A |  | RACiR 800 to 300 | 42.1389 | 417.5818 | 999.792 | 25.0184 | 0.010355 | 0.45922 | 519.913 | 557.501
Populus deltoides | LI2A |  | RACiR 800 to 300 | 42.0885 | 416.3203 | 1000.16 | 25.0138 | 0.010357 | 0.45974 | 518.398 | 555.926
Populus deltoides | LI2A |  | RACiR 800 to 300 | 42.1017 | 414.6418 | 1000.13 | 25.0155 | 0.010355 | 0.4595 | 516.761 | 554.293
Populus deltoides | LI2A |  | RACiR 800 to 300 | 42.0868 | 413.2293 | 1000.13 | 25.0158 | 0.010355 | 0.45947 | 515.291 | 552.796
Populus deltoides | LI2A |  | RACiR 800 to 300 | 41.8791 | 412.1523 | 999.826 | 25.0122 | 0.010356 | 0.45981 | 513.658 | 550.945
Populus deltoides | LI2A |  | RACiR 800 to 300 | 41.7883 | 410.854 | 1000.15 | 25.0144 | 0.010354 | 0.45952 | 512.192 | 549.377
Populus deltoides | LI2A |  | RACiR 800 to 300 | 41.7167 | 409.3916 | 999.838 | 25.0125 | 0.010353 | 0.45959 | 510.523 | 547.622
Populus deltoides | LI2A |  | RACiR 800 to 300 | 41.5337 | 408.3202 | 999.752 | 25.0131 | 0.010355 | 0.45964 | 509.015 | 545.92
Populus deltoides | LI2A |  | RACiR 800 to 300 | 41.5034 | 406.7023 | 1000.12 | 25.0171 | 0.010354 | 0.4593 | 507.367 | 544.233
Populus deltoides | LI2A |  | RACiR 800 to 300 | 41.516 | 405.2039 | 999.793 | 25.0147 | 0.010352 | 0.45932 | 505.857 | 542.721
Populus deltoides | LI2A |  | RACiR 800 to 300 | 41.3765 | 403.8644 | 1000.1 | 25.0152 | 0.010349 | 0.45912 | 504.22 | 540.937
Populus deltoides | LI2A |  | RACiR 800 to 300 | 41.41 | 402.3114 | 999.763 | 25.0148 | 0.010348 | 0.45907 | 502.715 | 539.453
Populus deltoides | LI2A |  | RACiR 800 to 300 | 41.4121 | 400.6713 | 1000.13 | 25.0152 | 0.010348 | 0.45899 | 501.06 | 537.785
Populus deltoides | LI2A |  | RACiR 800 to 300 | 41.1996 | 399.61 | 1000.11 | 25.0194 | 0.010347 | 0.45864 | 499.582 | 536.088
Populus deltoides | LI2A |  | RACiR 800 to 300 | 41.123 | 398.1478 | 999.834 | 25.0196 | 0.010347 | 0.4587 | 497.906 | 534.323
Populus deltoides | LI2A |  | RACiR 800 to 300 | 41.0751 | 396.7948 | 999.969 | 25.0197 | 0.010349 | 0.45878 | 496.399 | 532.755
Populus deltoides | LI2A |  | RACiR 800 to 300 | 40.8724 | 395.5967 | 1000.11 | 25.018 | 0.010348 | 0.45881 | 494.721 | 530.859
Populus deltoides | LI2A |  | RACiR 800 to 300 | 40.8581 | 394.1346 | 999.825 | 25.0164 | 0.010346 | 0.45878 | 493.197 | 529.313
Populus deltoides | LI2A |  | RACiR 800 to 300 | 40.8998 | 392.3956 | 1000.17 | 25.0147 | 0.010344 | 0.45878 | 491.509 | 527.651
Populus deltoides | LI2A |  | RACiR 800 to 300 | 40.7475 | 391.3196 | 999.777 | 25.0108 | 0.010343 | 0.45901 | 490.023 | 526.001
Populus deltoides | LI2A |  | RACiR 800 to 300 | 40.7103 | 389.8007 | 1000.13 | 25.009 | 0.01034 | 0.459 | 488.386 | 524.319
Populus deltoides | LI2A |  | RACiR 800 to 300 | 40.7437 | 388.3236 | 999.812 | 25.0051 | 0.01034 | 0.45927 | 486.89 | 522.843
Populus deltoides | LI2A |  | RACiR 800 to 300 | 40.5399 | 387.2298 | 1000.18 | 25.0044 | 0.010342 | 0.45944 | 485.287 | 521.024
Populus deltoides | LI2A |  | RACiR 800 to 300 | 40.4873 | 385.8123 | 999.814 | 25.0086 | 0.01034 | 0.45899 | 483.817 | 519.495
Populus deltoides | LI2A |  | RACiR 800 to 300 | 40.4606 | 384.2846 | 1000.14 | 25.0081 | 0.010342 | 0.45914 | 482.166 | 517.804
Populus deltoides | LI2A |  | RACiR 800 to 300 | 40.1765 | 383.5893 | 999.817 | 25.0064 | 0.010348 | 0.4596 | 480.737 | 516.083
Populus deltoides | LI2A |  | RACiR 800 to 300 | 40.0757 | 382.2701 | 999.779 | 25.0093 | 0.010362 | 0.46 | 479.092 | 514.337
Populus deltoides | LI2A |  | RACiR 800 to 300 | 40.0194 | 381.0102 | 1000 | 25.0086 | 0.010374 | 0.46054 | 477.578 | 512.757
Populus deltoides | LI2A |  | RACiR 800 to 300 | 39.8355 | 379.8068 | 999.841 | 25.0092 | 0.010385 | 0.46078 | 475.901 | 510.891
Populus deltoides | LI2A |  | RACiR 800 to 300 | 39.863 | 378.2379 | 1000.17 | 25.0095 | 0.010387 | 0.46071 | 474.374 | 509.386
Populus deltoides | LI2A |  | RACiR 800 to 300 | 39.9771 | 376.25 | 999.75 | 25.0066 | 0.01038 | 0.46032 | 472.668 | 507.774
Populus deltoides | LI2A |  | RACiR 800 to 300 | 39.8747 | 374.9478 | 1000.13 | 25.0065 | 0.010373 | 0.45985 | 471.203 | 506.193
Populus deltoides | LI2A |  | RACiR 800 to 300 | 39.8419 | 373.3119 | 999.809 | 25.0059 | 0.01036 | 0.45919 | 469.586 | 504.526
Populus deltoides | LI2A |  | RACiR 800 to 300 | 39.8419 | 371.7355 | 1000.14 | 25.0077 | 0.010349 | 0.4585 | 468.111 | 503.034
Populus deltoides | LI2A |  | RACiR 800 to 300 | 39.6063 | 370.5458 | 999.899 | 25.0094 | 0.010339 | 0.4579 | 466.49 | 501.16
Populus deltoides | LI2A |  | RACiR 800 to 300 | 39.5131 | 369.2941 | 999.803 | 25.0073 | 0.010329 | 0.45772 | 465.033 | 499.6
Populus deltoides | LI2A |  | RACiR 800 to 300 | 39.4588 | 367.7927 | 1000.11 | 25.0082 | 0.010324 | 0.45752 | 463.416 | 497.912
Populus deltoides | LI2A |  | RACiR 800 to 300 | 39.2164 | 366.9129 | 999.798 | 25.0089 | 0.010324 | 0.45751 | 461.983 | 496.225
Populus deltoides | LI2A |  | RACiR 800 to 300 | 39.151 | 365.4763 | 1000.16 | 25.0069 | 0.010324 | 0.45767 | 460.335 | 494.503
Populus deltoides | LI2A |  | RACiR 800 to 300 | 39.1483 | 364.0751 | 999.905 | 25.0058 | 0.010326 | 0.45789 | 458.851 | 493.003
Populus deltoides | LI2A |  | RACiR 800 to 300 | 39.0064 | 362.7549 | 1000.17 | 25.0059 | 0.010324 | 0.45785 | 457.195 | 491.193
Populus deltoides | LI2A |  | RACiR 800 to 300 | 39.0165 | 361.1751 | 999.789 | 25.0116 | 0.010322 | 0.45739 | 455.695 | 489.694
Populus deltoides | LI2A |  | RACiR 800 to 300 | 38.9977 | 359.6268 | 1000.17 | 25.0109 | 0.010318 | 0.45739 | 454.066 | 488.032
Populus deltoides | LI2A |  | RACiR 800 to 300 | 38.7502 | 358.7202 | 999.836 | 25.0117 | 0.010313 | 0.45716 | 452.634 | 486.34
Populus deltoides | LI2A |  | RACiR 800 to 300 | 38.6205 | 357.4545 | 1000.14 | 25.0073 | 0.010308 | 0.45738 | 451.003 | 484.57
Populus deltoides | LI2A |  | RACiR 800 to 300 | 38.5008 | 356.4222 | 999.788 | 25.0003 | 0.01031 | 0.45806 | 449.544 | 482.977
Populus deltoides | LI2A |  | RACiR 800 to 300 | 38.2661 | 355.2936 | 1000.17 | 25.0025 | 0.01031 | 0.458 | 447.885 | 481.077
Populus deltoides | LI2A |  | RACiR 800 to 300 | 38.2262 | 353.9682 | 999.823 | 25.0009 | 0.010312 | 0.45828 | 446.385 | 479.527
Populus deltoides | LI2A |  | RACiR 800 to 300 | 38.2706 | 352.2484 | 1000.17 | 25.0021 | 0.010315 | 0.4584 | 444.702 | 477.876
Populus deltoides | LI2A |  | RACiR 800 to 300 | 38.1553 | 351.0139 | 999.824 | 25.005 | 0.010316 | 0.45828 | 443.21 | 476.258
Populus deltoides | LI2A |  | RACiR 800 to 300 | 38.1547 | 349.3887 | 1000.2 | 25.0041 | 0.010314 | 0.45826 | 441.55 | 474.589
Populus deltoides | LI2A |  | RACiR 800 to 300 | 38.1098 | 348.0393 | 1000.01 | 25.0063 | 0.010315 | 0.45813 | 440.098 | 473.077
Populus deltoides | LI2A |  | RACiR 800 to 300 | 37.8678 | 346.9534 | 1000.21 | 25.0082 | 0.010315 | 0.45797 | 438.485 | 471.21
Populus deltoides | LI2A |  | RACiR 800 to 300 | 37.8374 | 345.5697 | 999.802 | 25.0075 | 0.010314 | 0.45797 | 437.001 | 469.689
Populus deltoides | LI2A |  | RACiR 800 to 300 | 37.7767 | 344.1251 | 1000.13 | 25.0091 | 0.010315 | 0.45797 | 435.392 | 468.006
Populus deltoides | LI2A |  | RACiR 800 to 300 | 37.5552 | 343.1752 | 999.779 | 25.0095 | 0.010317 | 0.45805 | 433.916 | 466.298
Populus deltoides | LI2A |  | RACiR 800 to 300 | 37.4918 | 341.6741 | 1000.16 | 25.0124 | 0.010317 | 0.45789 | 432.272 | 464.584
Populus deltoides | LI2A |  | RACiR 800 to 300 | 37.4054 | 340.4301 | 999.836 | 25.0122 | 0.010318 | 0.45801 | 430.787 | 463
Populus deltoides | LI2A |  | RACiR 800 to 300 | 37.2093 | 339.1912 | 1000.15 | 25.0152 | 0.010318 | 0.45775 | 429.138 | 461.144
Populus deltoides | LI2A |  | RACiR 800 to 300 | 37.2112 | 337.728 | 1000.15 | 25.0139 | 0.010316 | 0.45773 | 427.647 | 459.644
Populus deltoides | LI2A |  | RACiR 800 to 300 | 37.2545 | 336.0642 | 999.8 | 25.0109 | 0.010316 | 0.45797 | 425.994 | 458.019
Populus deltoides | LI2A |  | RACiR 800 to 300 | 37.1712 | 334.75 | 1000.13 | 25.0152 | 0.010318 | 0.45769 | 424.523 | 456.453
Populus deltoides | LI2A |  | RACiR 800 to 300 | 37.158 | 333.1595 | 1000.11 | 25.0163 | 0.010315 | 0.45748 | 422.907 | 454.816
Populus deltoides | LI2A |  | RACiR 800 to 300 | 37.1197 | 331.8054 | 1000.12 | 25.0181 | 0.010316 | 0.45746 | 421.441 | 453.301
Populus deltoides | LI2A |  | RACiR 800 to 300 | 36.8918 | 330.6806 | 999.826 | 25.0199 | 0.010315 | 0.45724 | 419.829 | 451.454
Populus deltoides | LI2A |  | RACiR 800 to 300 | 36.783 | 329.5434 | 1000.11 | 25.0169 | 0.010315 | 0.45745 | 418.384 | 449.888
Populus deltoides | LI2A |  | RACiR 800 to 300 | 36.6244 | 328.3922 | 999.795 | 25.0124 | 0.010315 | 0.45785 | 416.779 | 448.111
Populus deltoides | LI2A |  | RACiR 800 to 300 | 36.4057 | 327.4234 | 999.86 | 25.0144 | 0.010315 | 0.45768 | 415.337 | 446.439
Populus deltoides | LI2A |  | RACiR 800 to 300 | 36.3371 | 325.9691 | 1000.13 | 25.013 | 0.010312 | 0.45753 | 413.725 | 444.75
Populus deltoides | LI2A |  | RACiR 800 to 300 | 36.2878 | 324.6514 | 999.822 | 25.0128 | 0.010312 | 0.45759 | 412.257 | 443.22
Populus deltoides | LI2A |  | RACiR 800 to 300 | 36.1182 | 323.4158 | 1000.18 | 25.0156 | 0.010312 | 0.45739 | 410.657 | 441.44
Populus deltoides | LI2A |  | RACiR 800 to 300 | 36.1322 | 321.9566 | 999.782 | 25.0147 | 0.010312 | 0.45743 | 409.188 | 439.977
Populus deltoides | LI2A |  | RACiR 800 to 300 | 36.0859 | 320.4402 | 1000.15 | 25.018 | 0.010313 | 0.45729 | 407.562 | 438.292
Populus deltoides | LI2A |  | RACiR 800 to 300 | 35.8745 | 319.4681 | 999.839 | 25.019 | 0.010313 | 0.45721 | 406.115 | 436.62
Populus deltoides | LI2A |  | RACiR 800 to 300 | 35.7363 | 318.1479 | 1000.2 | 25.0186 | 0.010311 | 0.45712 | 404.475 | 434.834
Populus deltoides | LI2A |  | RACiR 800 to 300 | 35.5829 | 317.1008 | 1000.03 | 25.0155 | 0.010312 | 0.45737 | 403.017 | 433.21
Populus deltoides | LI2A |  | RACiR 800 to 300 | 35.3102 | 316.0362 | 1000.12 | 25.0167 | 0.010307 | 0.457 | 401.392 | 431.304
Populus deltoides | LI2A |  | RACiR 800 to 300 | 35.2778 | 314.689 | 999.846 | 25.0137 | 0.010306 | 0.45712 | 399.919 | 429.788
Populus deltoides | LI2A |  | RACiR 800 to 300 | 35.2685 | 313.1247 | 1000.22 | 25.0125 | 0.010304 | 0.45709 | 398.302 | 428.147
Populus deltoides | LI2A |  | RACiR 800 to 300 | 35.1373 | 311.951 | 999.783 | 25.0117 | 0.010301 | 0.45695 | 396.836 | 426.537
Populus deltoides | LI2A |  | RACiR 800 to 300 | 35.1165 | 310.4372 | 1000.22 | 25.0087 | 0.010301 | 0.45716 | 395.202 | 424.871
Populus deltoides | LI2A |  | RACiR 800 to 300 | 35.0759 | 309.0631 | 999.811 | 25.01 | 0.010301 | 0.45703 | 393.731 | 423.348
Populus deltoides | LI2A |  | RACiR 800 to 300 | 34.8453 | 308 | 1000.2 | 25.0082 | 0.010301 | 0.45716 | 392.109 | 421.487
Populus deltoides | LI2A |  | RACiR 800 to 300 | 34.7687 | 306.7487 | 999.787 | 25.0096 | 0.010303 | 0.45718 | 390.657 | 419.95
Populus deltoides | LI2A |  | RACiR 800 to 300 | 34.7102 | 305.3007 | 1000.11 | 25.0071 | 0.010304 | 0.45744 | 389 | 418.221
Populus deltoides | LI2A |  | RACiR 800 to 300 | 34.4865 | 304.407 | 999.853 | 25.0052 | 0.010305 | 0.45765 | 387.553 | 416.539
Populus deltoides | LI2A |  | RACiR 800 to 300 | 34.3871 | 303.0016 | 1000.17 | 25.0074 | 0.010304 | 0.45748 | 385.927 | 414.806
Populus deltoides | LI2A |  | RACiR 800 to 300 | 34.3289 | 301.6816 | 999.957 | 25.0094 | 0.010306 | 0.45747 | 384.451 | 413.26
Populus deltoides | LI2A |  | RACiR 800 to 300 | 34.1004 | 300.5752 | 1000.11 | 25.0128 | 0.010306 | 0.45725 | 382.855 | 411.425
Populus deltoides | LI2A |  | RACiR 800 to 300 | 34.0776 | 299.1716 | 1000.13 | 25.014 | 0.010306 | 0.4572 | 381.378 | 409.917
Populus deltoides | LI2A |  | RACiR 800 to 300 | 34.1159 | 297.5417 | 999.805 | 25.0111 | 0.010307 | 0.45741 | 379.758 | 408.32
Populus deltoides | LI2A |  | RACiR 800 to 300 | 33.9911 | 296.4306 | 999.778 | 25.0075 | 0.010306 | 0.45761 | 378.308 | 406.735
Populus deltoides | LI2A |  | RACiR 800 to 300 | 33.9675 | 294.9578 | 1000.09 | 25.0049 | 0.010305 | 0.45781 | 376.714 | 405.108
Populus deltoides | LI2A |  | RACiR 800 to 300 | 33.9165 | 293.687 | 1000.1 | 25.0029 | 0.010305 | 0.45791 | 375.282 | 403.614
Populus deltoides | LI2A |  | RACiR 800 to 300 | 33.6704 | 292.6972 | 999.862 | 24.9997 | 0.010306 | 0.45825 | 373.667 | 401.742
Populus deltoides | LI2A |  | RACiR 800 to 300 | 33.5547 | 291.5187 | 1000.1 | 25.0016 | 0.010306 | 0.4581 | 372.232 | 400.183
Populus deltoides | LI2A |  | RACiR 800 to 300 | 33.439 | 290.1962 | 999.746 | 25.0012 | 0.010306 | 0.4581 | 370.624 | 398.445
Populus deltoides | LI2A |  | RACiR 800 to 300 | 33.2237 | 289.214 | 1000.12 | 25.0017 | 0.010304 | 0.45789 | 369.18 | 396.775
Populus deltoides | LI2A |  | RACiR 800 to 300 | 33.1124 | 287.7792 | 1000.15 | 25.0058 | 0.010298 | 0.4572 | 367.585 | 395.059
Populus deltoides | LI2A |  | RACiR 800 to 300 | 33.0486 | 286.503 | 1000.12 | 25.0042 | 0.010297 | 0.45722 | 366.134 | 393.532
Populus deltoides | LI2A |  | RACiR 800 to 300 | 32.8713 | 285.3006 | 999.766 | 25.0012 | 0.010294 | 0.4572 | 364.513 | 391.722
Populus deltoides | LI2A |  | RACiR 800 to 300 | 32.8253 | 283.9551 | 1000.12 | 25.0026 | 0.010294 | 0.45712 | 363.048 | 390.201
Populus deltoides | LI2A |  | RACiR 800 to 300 | 32.8022 | 282.3615 | 999.745 | 25.0077 | 0.010295 | 0.45686 | 361.414 | 388.534
Populus deltoides | LI2A |  | RACiR 800 to 300 | 32.5875 | 281.4014 | 1000.12 | 25.0084 | 0.010295 | 0.4568 | 359.968 | 386.861
Populus deltoides | LI2A |  | RACiR 800 to 300 | 32.4401 | 280.1074 | 999.815 | 25.0116 | 0.010294 | 0.4565 | 358.368 | 385.104
Populus deltoides | LI2A |  | RACiR 800 to 300 | 32.2871 | 279.0122 | 1000.09 | 25.0125 | 0.010293 | 0.45638 | 356.929 | 383.502
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 322.6469 | 29.5285 | 25.2254 | 1000.11 | 0.0073922 | 0.3352 | 418.917 | 434.405
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 318.6574 | 28.9017 | 25.2026 | 1000.15 | 0.0073742 | 0.33521 | 412.923 | 427.737
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 313.5283 | 28.6301 | 25.1719 | 1000.18 | 0.0073504 | 0.33529 | 406.814 | 421.308
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 309.3691 | 27.9686 | 25.148 | 999.836 | 0.0073247 | 0.33492 | 400.642 | 414.427
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 305.0127 | 27.403 | 25.1261 | 999.842 | 0.00731 | 0.33516 | 394.401 | 407.575
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 299.3933 | 27.3274 | 25.1049 | 999.914 | 0.0073118 | 0.33646 | 388.083 | 401.14
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 294.778 | 26.953 | 25.0884 | 999.866 | 0.0073233 | 0.33815 | 381.811 | 394.458
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 290.0269 | 26.6775 | 25.0748 | 1000.17 | 0.007343 | 0.34032 | 375.597 | 387.937
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 284.6745 | 26.5893 | 25.0691 | 999.867 | 0.0073645 | 0.34218 | 369.42 | 381.64
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 280.6203 | 26.0214 | 25.0647 | 999.834 | 0.0073801 | 0.34352 | 363.286 | 374.905
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 276.4347 | 25.4719 | 25.057 | 1000.12 | 0.0073911 | 0.34474 | 357.114 | 368.148
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 270.793 | 25.3537 | 25.0547 | 1000.21 | 0.0073975 | 0.3453 | 350.878 | 361.756
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 265.8252 | 24.9778 | 25.0589 | 1000.17 | 0.0074018 | 0.3453 | 344.705 | 355.169
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 261.1807 | 24.4857 | 25.0673 | 999.804 | 0.0074047 | 0.34499 | 338.593 | 348.527
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 256.02 | 24.1035 | 25.0849 | 1000.13 | 0.0074056 | 0.34394 | 332.435 | 341.948
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 251.3631 | 23.6121 | 25.0882 | 1000.11 | 0.0074086 | 0.3439 | 326.244 | 335.227
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 246.3442 | 23.267 | 25.0877 | 1000.09 | 0.0074112 | 0.34407 | 320.067 | 328.666
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 240.6432 | 23.1052 | 25.0991 | 1000.15 | 0.0074124 | 0.34342 | 313.907 | 322.304
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 236.4589 | 22.4658 | 25.1014 | 999.893 | 0.0074116 | 0.34319 | 307.797 | 315.516
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 232.0932 | 21.9153 | 25.0947 | 1000.14 | 0.0074129 | 0.34363 | 301.625 | 308.754
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 226.2 | 21.8132 | 25.1042 | 1000.07 | 0.007412 | 0.34298 | 295.435 | 302.422
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 221.326 | 21.3688 | 25.1164 | 999.79 | 0.0074118 | 0.3422 | 289.301 | 295.804
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 216.4255 | 20.9974 | 25.1188 | 999.841 | 0.0074129 | 0.34211 | 283.215 | 289.308
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 211.0026 | 20.8158 | 25.1227 | 1000.16 | 0.0074144 | 0.34194 | 277.17 | 283.042
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 206.9602 | 20.177 | 25.1192 | 1000.12 | 0.0074129 | 0.34205 | 271.13 | 276.325
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 203.0109 | 19.499 | 25.1138 | 1000.07 | 0.0074119 | 0.34223 | 265.055 | 269.533
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 197.6304 | 19.262 | 25.1179 | 999.715 | 0.0074107 | 0.34184 | 258.927 | 263.128
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 192.7737 | 18.8636 | 25.1206 | 1000.08 | 0.0074089 | 0.34152 | 252.842 | 256.605
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 187.7778 | 18.558 | 25.1139 | 1000.06 | 0.007407 | 0.3417 | 246.798 | 250.216
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 182.8543 | 18.2365 | 25.1091 | 1000.13 | 0.0074023 | 0.34166 | 240.82 | 243.877
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 179.0239 | 17.5204 | 25.1108 | 1000.07 | 0.0073975 | 0.34123 | 234.855 | 237.157
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 174.9619 | 16.8861 | 25.1071 | 1000.14 | 0.0073962 | 0.34129 | 228.814 | 230.443
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 169.9449 | 16.5469 | 25.1087 | 999.789 | 0.0073925 | 0.34088 | 222.745 | 223.995
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 165.7751 | 15.8958 | 25.1136 | 1000.08 | 0.0073869 | 0.34017 | 216.656 | 217.215
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 161.6397 | 15.2475 | 25.1125 | 1000.12 | 0.0073836 | 0.33993 | 210.536 | 210.407
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 156.4227 | 14.9773 | 25.1083 | 999.752 | 0.0073809 | 0.33994 | 204.398 | 203.959
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 151.8938 | 14.4783 | 25.1037 | 1000.02 | 0.0073756 | 0.33981 | 198.301 | 197.323
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 147.0787 | 14.074 | 25.1054 | 1000.04 | 0.0073713 | 0.33938 | 192.233 | 190.811
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 141.879 | 13.8125 | 25.1081 | 1000.1 | 0.0073709 | 0.33911 | 186.176 | 184.454
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 137.9284 | 13.1616 | 25.1009 | 1000.02 | 0.0073673 | 0.33927 | 180.176 | 177.764
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 133.7378 | 12.5601 | 25.1065 | 999.752 | 0.0073634 | 0.33859 | 174.181 | 171.129
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 128.8425 | 12.1975 | 25.1062 | 1000.07 | 0.0073586 | 0.33825 | 168.133 | 164.679
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 124.6802 | 11.6014 | 25.0969 | 1000.01 | 0.0073536 | 0.33842 | 162.077 | 157.988
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 119.6701 | 11.2701 | 25.0965 | 1000.05 | 0.0073473 | 0.33795 | 156.015 | 151.555
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 114.5319 | 11.0063 | 25.0946 | 999.712 | 0.0073408 | 0.33759 | 150.011 | 145.248
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 110.8542 | 10.2696 | 25.0919 | 1000.05 | 0.0073361 | 0.33737 | 144.066 | 138.528
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 106.9708 | 9.5837 | 25.09 | 999.99 | 0.0073316 | 0.33709 | 138.066 | 131.804
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 102.1621 | 9.1946 | 25.0876 | 999.687 | 0.0073267 | 0.33681 | 132.008 | 125.318
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 98.612 | 8.3863 | 25.0825 | 1000.04 | 0.0073204 | 0.33662 | 125.967 | 118.429
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 94.7779 | 7.6654 | 25.0719 | 1000.04 | 0.0073142 | 0.33676 | 119.871 | 111.574
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 89.5337 | 7.3982 | 25.0647 | 1000.07 | 0.0073072 | 0.33668 | 113.713 | 105.109
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 85.0851 | 6.8747 | 25.0603 | 1000.05 | 0.0073015 | 0.33648 | 107.607 | 98.4397
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 80.3216 | 6.4664 | 25.0628 | 999.995 | 0.007297 | 0.33599 | 101.542 | 91.9283
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 75.4366 | 6.1258 | 25.062 | 999.751 | 0.0072933 | 0.33574 | 95.5392 | 85.5466
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 71.7446 | 5.3834 | 25.0567 | 1000.01 | 0.0072854 | 0.33545 | 89.5446 | 78.7705
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 67.8687 | 4.6847 | 25.0558 | 1000.05 | 0.0072809 | 0.3351 | 83.4931 | 71.983
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 63.0145 | 4.3145 | 25.0472 | 1000.03 | 0.0072775 | 0.33531 | 77.4052 | 65.487
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 58.9514 | 3.6935 | 25.0379 | 1000.03 | 0.0072716 | 0.33544 | 71.3749 | 58.7964
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 54.5027 | 3.2048 | 25.0366 | 1000.01 | 0.007266 | 0.33509 | 65.3664 | 52.2618
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 49.8501 | 2.7936 | 25.0364 | 999.98 | 0.0072617 | 0.33474 | 59.3803 | 45.8261
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 46.2623 | 2.0226 | 25.0326 | 999.959 | 0.0072556 | 0.33448 | 53.39 | 39.028
Populus deltoides | LI2A |  | RACiR 500 to 0 2.5 min | 42.6276 | 1.2509 | 25.0298 | 999.962 | 0.0072504 | 0.33427 | 47.3457 | 32.1746
Populus deltoides | LI2A |  | RACiR 500 to 0 | 37.9632 | 337.7224 | 1000.01 | 25.0107 | 0.010254 | 0.45477 | 445.534 | 481.191
Populus deltoides | LI2A |  | RACiR 500 to 0 | 37.7873 | 336.5973 | 1000.09 | 25.0107 | 0.010261 | 0.45504 | 443.555 | 478.851
Populus deltoides | LI2A |  | RACiR 500 to 0 | 37.6088 | 334.9337 | 1000.05 | 25.012 | 0.01026 | 0.45498 | 441.559 | 476.619
Populus deltoides | LI2A |  | RACiR 500 to 0 | 37.5442 | 333.0517 | 1000.01 | 25.0082 | 0.010257 | 0.45484 | 440.1 | 474.894
Populus deltoides | LI2A |  | RACiR 500 to 0 | 37.3728 | 331.7912 | 1000 | 25.0056 | 0.010255 | 0.45492 | 438.035 | 472.642
Populus deltoides | LI2A |  | RACiR 500 to 0 | 37.3385 | 329.4438 | 1000.01 | 25.0064 | 0.01025 | 0.45458 | 435.962 | 470.221
Populus deltoides | LI2A |  | RACiR 500 to 0 | 37.1369 | 327.5581 | 999.964 | 25.0032 | 0.010248 | 0.4544 | 434.435 | 468.62
Populus deltoides | LI2A |  | RACiR 500 to 0 | 37.0866 | 326.1692 | 999.953 | 25.0022 | 0.010241 | 0.45395 | 432.307 | 466.112
Populus deltoides | LI2A |  | RACiR 500 to 0 | 36.8108 | 324.7217 | 999.999 | 24.9982 | 0.010229 | 0.45326 | 430.142 | 463.667
Populus deltoides | LI2A |  | RACiR 500 to 0 | 36.566 | 323.1171 | 1000.08 | 24.9976 | 0.010232 | 0.45344 | 428.565 | 461.903
Populus deltoides | LI2A |  | RACiR 500 to 0 | 36.4577 | 321.7677 | 1000 | 24.9985 | 0.010247 | 0.45412 | 426.339 | 459.485
Populus deltoides | LI2A |  | RACiR 500 to 0 | 36.2137 | 320.0769 | 1000.03 | 25.0009 | 0.010261 | 0.45506 | 424.092 | 457.16
Populus deltoides | LI2A |  | RACiR 500 to 0 | 36.1705 | 318.0158 | 1000.01 | 25.0079 | 0.010247 | 0.45457 | 422.508 | 455.388
Populus deltoides | LI2A |  | RACiR 500 to 0 | 35.9516 | 316.9957 | 1000.04 | 25.0106 | 0.010219 | 0.45317 | 420.292 | 453.106
Populus deltoides | LI2A |  | RACiR 500 to 0 | 35.9652 | 314.8506 | 1000.08 | 25.0122 | 0.010194 | 0.45228 | 418.071 | 450.661
Populus deltoides | LI2A |  | RACiR 500 to 0 | 35.7786 | 313.1061 | 1000.02 | 25.0065 | 0.010193 | 0.45249 | 416.496 | 449.022
Populus deltoides | LI2A |  | RACiR 500 to 0 | 35.7264 | 311.6377 | 999.969 | 25.0001 | 0.010196 | 0.45304 | 414.291 | 446.522
Populus deltoides | LI2A |  | RACiR 500 to 0 | 35.4747 | 309.9529 | 999.991 | 24.9953 | 0.010197 | 0.45323 | 412.066 | 444.037
Populus deltoides | LI2A |  | RACiR 500 to 0 | 35.2416 | 308.1722 | 1000.04 | 24.9954 | 0.010196 | 0.45315 | 410.46 | 442.307
Populus deltoides | LI2A |  | RACiR 500 to 0 | 35.0511 | 306.9579 | 999.996 | 25.0007 | 0.010197 | 0.45305 | 408.205 | 439.796
Populus deltoides | LI2A |  | RACiR 500 to 0 | 34.7249 | 305.4133 | 999.666 | 25.006 | 0.010214 | 0.45341 | 405.945 | 437.479
Populus deltoides | LI2A |  | RACiR 500 to 0 | 34.5512 | 303.5845 | 1000.02 | 25.0082 | 0.010227 | 0.45386 | 404.335 | 435.648
Populus deltoides | LI2A |  | RACiR 500 to 0 | 34.3176 | 302.4336 | 999.987 | 25.0104 | 0.010246 | 0.45451 | 402.061 | 433.371
Populus deltoides | LI2A |  | RACiR 500 to 0 | 34.3396 | 300.1213 | 1000.03 | 25.0099 | 0.010243 | 0.45465 | 399.833 | 430.941
Populus deltoides | LI2A |  | RACiR 500 to 0 | 34.1198 | 298.3599 | 999.98 | 25.0158 | 0.010233 | 0.45449 | 398.237 | 429.274
Populus deltoides | LI2A |  | RACiR 500 to 0 | 34.1026 | 296.8927 | 999.976 | 25.0171 | 0.010216 | 0.45396 | 396.042 | 426.804
Populus deltoides | LI2A |  | RACiR 500 to 0 | 33.8537 | 295.2636 | 999.964 | 25.02 | 0.010204 | 0.45337 | 393.804 | 424.312
Populus deltoides | LI2A |  | RACiR 500 to 0 | 33.6671 | 293.342 | 1000.06 | 25.0164 | 0.010205 | 0.45312 | 392.186 | 422.492
Populus deltoides | LI2A |  | RACiR 500 to 0 | 33.5093 | 291.9393 | 1000.01 | 25.0174 | 0.010206 | 0.45289 | 389.923 | 419.887
Populus deltoides | LI2A |  | RACiR 500 to 0 | 33.331 | 289.8936 | 999.926 | 25.0198 | 0.010211 | 0.45314 | 387.62 | 417.39
Populus deltoides | LI2A |  | RACiR 500 to 0 | 33.3201 | 287.5661 | 1000.05 | 25.0235 | 0.010206 | 0.45277 | 385.987 | 415.507
Populus deltoides | LI2A |  | RACiR 500 to 0 | 33.1881 | 286.3058 | 999.957 | 25.0249 | 0.010204 | 0.45289 | 383.648 | 413.171
Populus deltoides | LI2A |  | RACiR 500 to 0 | 33.2133 | 284.0065 | 999.981 | 25.0297 | 0.010214 | 0.45308 | 381.327 | 410.622
Populus deltoides | LI2A |  | RACiR 500 to 0 | 32.9787 | 282.3306 | 999.955 | 25.0314 | 0.010228 | 0.45373 | 379.679 | 408.95
Populus deltoides | LI2A |  | RACiR 500 to 0 | 32.9343 | 280.8828 | 1000.04 | 25.0306 | 0.010246 | 0.45429 | 377.37 | 406.375
Populus deltoides | LI2A |  | RACiR 500 to 0 | 32.6461 | 279.4665 | 999.976 | 25.0282 | 0.010237 | 0.4539 | 375.052 | 403.851
Populus deltoides | LI2A |  | RACiR 500 to 0 | 32.4255 | 277.9088 | 1000.04 | 25.0262 | 0.01022 | 0.45298 | 373.421 | 402.043
Populus deltoides | LI2A |  | RACiR 500 to 0 | 32.2807 | 276.7262 | 999.998 | 25.025 | 0.010198 | 0.45178 | 371.131 | 399.55
Populus deltoides | LI2A |  | RACiR 500 to 0 | 32.0713 | 275.0279 | 999.976 | 25.0211 | 0.010184 | 0.45104 | 368.851 | 397.232
Populus deltoides | LI2A |  | RACiR 500 to 0 | 31.9508 | 272.9636 | 1000.03 | 25.02 | 0.010186 | 0.4512 | 367.239 | 395.475
Populus deltoides | LI2A |  | RACiR 500 to 0 | 31.7786 | 271.6531 | 1000.06 | 25.0195 | 0.010189 | 0.45108 | 364.967 | 393.213
Populus deltoides | LI2A |  | RACiR 500 to 0 | 31.7993 | 269.2349 | 999.968 | 25.0226 | 0.01019 | 0.45114 | 362.716 | 390.715
Populus deltoides | LI2A |  | RACiR 500 to 0 | 31.6298 | 267.3228 | 1000.05 | 25.0271 | 0.010191 | 0.45133 | 361.102 | 389.044
Populus deltoides | LI2A |  | RACiR 500 to 0 | 31.5878 | 265.8277 | 999.989 | 25.0298 | 0.010206 | 0.45216 | 358.871 | 386.514
Populus deltoides | LI2A |  | RACiR 500 to 0 | 31.2887 | 264.3429 | 1000.03 | 25.0324 | 0.010227 | 0.45319 | 356.618 | 384.026
Populus deltoides | LI2A |  | RACiR 500 to 0 | 31.0449 | 262.6733 | 999.964 | 25.0292 | 0.010239 | 0.45373 | 354.999 | 382.253
Populus deltoides | LI2A |  | RACiR 500 to 0 | 30.8735 | 261.3823 | 999.976 | 25.031 | 0.010243 | 0.45399 | 352.754 | 379.782
Populus deltoides | LI2A |  | RACiR 500 to 0 | 30.68 | 259.6119 | 999.942 | 25.0302 | 0.010228 | 0.45309 | 350.519 | 377.404
Populus deltoides | LI2A |  | RACiR 500 to 0 | 30.5778 | 257.7692 | 1000.02 | 25.0321 | 0.010215 | 0.45234 | 348.914 | 375.612
Populus deltoides | LI2A |  | RACiR 500 to 0 | 30.4227 | 254.5565 | 999.953 | 25.0281 | 0.010201 | 0.45139 | 346.641 | 373.334
Populus deltoides | LI2A |  | RACiR 500 to 0 | 30.5642 | 252.1047 | 1000.02 | 25.0231 | 0.010193 | 0.4508 | 344.402 | 370.911
Populus deltoides | LI2A |  | RACiR 500 to 0 | 30.8284 | 250.0221 | 999.971 | 25.0222 | 0.010192 | 0.45055 | 342.821 | 369.273
Populus deltoides | LI2A |  | RACiR 500 to 0 | 30.6664 | 248.402 | 1000.02 | 25.0204 | 0.010195 | 0.45055 | 340.632 | 366.765
Populus deltoides | LI2A |  | RACiR 500 to 0 | 30.0675 | 247.6467 | 999.987 | 25.0306 | 0.010181 | 0.45023 | 338.423 | 364.291
Populus deltoides | LI2A |  | RACiR 500 to 0 | 29.4139 | 247.5965 | 999.989 | 25.0316 | 0.010166 | 0.44948 | 336.834 | 362.513
Populus deltoides | LI2A |  | RACiR 500 to 0 | 28.8242 | 246.6521 | 999.999 | 25.0329 | 0.010163 | 0.44953 | 334.577 | 360.048
Populus deltoides | LI2A |  | RACiR 500 to 0 | 28.7052 | 244.6008 | 999.988 | 25.0303 | 0.010179 | 0.4503 | 332.344 | 357.7
Populus deltoides | LI2A |  | RACiR 500 to 0 | 28.5559 | 243.3143 | 1000.02 | 25.0293 | 0.010203 | 0.4517 | 328.492 | 353.669
Populus deltoides | LI2A |  | RACiR 500 to 0 | 28.6368 | 241.0106 | 1000.02 | 25.027 | 0.010198 | 0.45173 | 326.293 | 351.595
Populus deltoides | LI2A |  | RACiR 500 to 0 | 28.4968 | 239.2051 | 999.978 | 25.0299 | 0.010195 | 0.45156 | 324.782 | 350.335
Populus deltoides | LI2A |  | RACiR 500 to 0 | 28.4729 | 237.7435 | 999.95 | 25.0297 | 0.010191 | 0.45144 | 322.78 | 348.152
Populus deltoides | LI2A |  | RACiR 500 to 0 | 28.219 | 236.0835 | 999.961 | 25.0312 | 0.010191 | 0.45066 | 320.792 | 345.549
Populus deltoides | LI2A |  | RACiR 500 to 0 | 27.9041 | 234.4634 | 999.939 | 25.0269 | 0.010187 | 0.45039 | 319.315 | 343.407
Populus deltoides | LI2A |  | RACiR 500 to 0 | 27.6948 | 233.1976 | 1000.02 | 25.0271 | 0.010176 | 0.44987 | 317.1 | 340.586
Populus deltoides | LI2A |  | RACiR 500 to 0 | 27.3765 | 231.5125 | 1000.06 | 25.0277 | 0.010169 | 0.44991 | 314.724 | 338.067
Populus deltoides | LI2A |  | RACiR 500 to 0 | 27.2309 | 229.5729 | 999.964 | 25.0276 | 0.010166 | 0.45002 | 313.054 | 336.237
Populus deltoides | LI2A |  | RACiR 500 to 0 | 26.9992 | 228.5353 | 999.953 | 25.0262 | 0.010184 | 0.45111 | 310.719 | 333.964
Populus deltoides | LI2A |  | RACiR 500 to 0 | 26.98 | 226.3784 | 999.947 | 25.028 | 0.010203 | 0.45176 | 308.47 | 331.567
Populus deltoides | LI2A |  | RACiR 500 to 0 | 26.6969 | 224.8535 | 1000.06 | 25.0254 | 0.010217 | 0.45228 | 306.849 | 329.911
Populus deltoides | LI2A |  | RACiR 500 to 0 | 26.6106 | 223.4657 | 999.956 | 25.0258 | 0.010217 | 0.45184 | 304.651 | 327.443
Populus deltoides | LI2A |  | RACiR 500 to 0 | 26.2506 | 222.0541 | 999.97 | 25.0291 | 0.010189 | 0.45055 | 302.468 | 324.92
Populus deltoides | LI2A |  | RACiR 500 to 0 | 26.0189 | 220.4443 | 999.998 | 25.0275 | 0.010166 | 0.44927 | 300.884 | 323.108
Populus deltoides | LI2A |  | RACiR 500 to 0 | 25.8728 | 219.2877 | 999.99 | 25.0251 | 0.010141 | 0.44814 | 298.599 | 320.483
Populus deltoides | LI2A |  | RACiR 500 to 0 | 25.7216 | 217.5404 | 999.985 | 25.0127 | 0.010138 | 0.44815 | 296.284 | 318
Populus deltoides | LI2A |  | RACiR 500 to 0 | 25.6999 | 215.2907 | 999.911 | 25.0125 | 0.01014 | 0.44851 | 294.647 | 316.12
Populus deltoides | LI2A |  | RACiR 500 to 0 | 25.5445 | 213.9735 | 999.978 | 25.0135 | 0.010145 | 0.4488 | 292.358 | 313.794
Populus deltoides | LI2A |  | RACiR 500 to 0 | 25.5252 | 211.7376 | 999.994 | 25.0179 | 0.010143 | 0.44892 | 290.14 | 311.278
Populus deltoides | LI2A |  | RACiR 500 to 0 | 25.2282 | 210.1535 | 999.963 | 25.0183 | 0.010143 | 0.4489 | 288.529 | 309.567
Populus deltoides | LI2A |  | RACiR 500 to 0 | 25.1442 | 208.7182 | 999.947 | 25.0201 | 0.010142 | 0.44867 | 286.308 | 306.97
Populus deltoides | LI2A |  | RACiR 500 to 0 | 24.8158 | 207.1913 | 999.938 | 25.0242 | 0.010161 | 0.44963 | 284.011 | 304.427
Populus deltoides | LI2A |  | RACiR 500 to 0 | 24.5553 | 205.6317 | 999.939 | 25.0265 | 0.010177 | 0.45053 | 282.381 | 302.644
Populus deltoides | LI2A |  | RACiR 500 to 0 | 24.3938 | 204.5228 | 999.981 | 25.0223 | 0.01019 | 0.45178 | 280.086 | 300.184
Populus deltoides | LI2A |  | RACiR 500 to 0 | 24.1694 | 202.9788 | 999.937 | 25.0172 | 0.010183 | 0.45123 | 277.81 | 297.863
Populus deltoides | LI2A |  | RACiR 500 to 0 | 24.0768 | 201.0646 | 999.895 | 25.013 | 0.01017 | 0.45042 | 276.22 | 296.103
Populus deltoides | LI2A |  | RACiR 500 to 0 | 23.8945 | 199.9704 | 999.984 | 25.0158 | 0.01016 | 0.4497 | 273.985 | 293.827
Populus deltoides | LI2A |  | RACiR 500 to 0 | 23.8323 | 198.0603 | 999.937 | 25.0155 | 0.010139 | 0.44881 | 271.809 | 291.336
Populus deltoides | LI2A |  | RACiR 500 to 0 | 23.539 | 196.6362 | 999.941 | 25.0202 | 0.010125 | 0.44814 | 270.237 | 289.664
Populus deltoides | LI2A |  | RACiR 500 to 0 | 23.4513 | 195.307 | 999.928 | 25.0203 | 0.010107 | 0.4473 | 268.04 | 287.115
Populus deltoides | LI2A |  | RACiR 500 to 0 | 23.0961 | 193.9583 | 1000.26 | 25.015 | 0.010106 | 0.44737 | 265.845 | 284.642
Populus deltoides | LI2A |  | RACiR 500 to 0 | 22.8682 | 192.2291 | 999.93 | 25.0115 | 0.01011 | 0.44806 | 264.255 | 282.88
Populus deltoides | LI2A |  | RACiR 500 to 0 | 22.7063 | 191.0253 | 999.872 | 25.0061 | 0.010117 | 0.44896 | 262.052 | 280.438
Populus deltoides | LI2A |  | RACiR 500 to 0 | 22.5216 | 189.1868 | 999.903 | 25.0062 | 0.010118 | 0.44942 | 259.826 | 278.103
Populus deltoides | LI2A |  | RACiR 500 to 0 | 22.3715 | 187.3477 | 1000.24 | 25.0069 | 0.010129 | 0.44973 | 258.262 | 276.348
Populus deltoides | LI2A |  | RACiR 500 to 0 | 22.1484 | 186.2499 | 999.882 | 25.0141 | 0.010151 | 0.45059 | 256.066 | 274.076
Populus deltoides | LI2A |  | RACiR 500 to 0 | 22.0592 | 184.3331 | 1000.27 | 25.0152 | 0.010169 | 0.45088 | 253.924 | 271.631
Populus deltoides | LI2A |  | RACiR 500 to 0 | 21.7763 | 182.8567 | 1000.02 | 25.013 | 0.010172 | 0.4508 | 252.38 | 269.987
Populus deltoides | LI2A |  | RACiR 500 to 0 | 21.6633 | 181.6002 | 999.945 | 25.0125 | 0.010166 | 0.45063 | 250.221 | 267.454
Populus deltoides | LI2A |  | RACiR 500 to 0 | 21.395 | 180.0352 | 999.912 | 25.0165 | 0.010154 | 0.45008 | 248.004 | 264.99
Populus deltoides | LI2A |  | RACiR 500 to 0 | 21.1599 | 178.4214 | 999.868 | 25.0146 | 0.010144 | 0.44992 | 246.424 | 263.233
Populus deltoides | LI2A |  | RACiR 500 to 0 | 21.0116 | 177.2299 | 999.887 | 25.0138 | 0.01013 | 0.4493 | 244.199 | 260.804
Populus deltoides | LI2A |  | RACiR 500 to 0 | 20.7229 | 175.7586 | 1000.23 | 25.0128 | 0.010126 | 0.44917 | 241.993 | 258.428
Populus deltoides | LI2A |  | RACiR 500 to 0 | 20.5575 | 173.9614 | 999.929 | 25.0139 | 0.010125 | 0.44866 | 240.43 | 256.631
Populus deltoides | LI2A |  | RACiR 500 to 0 | 20.331 | 172.9091 | 999.922 | 25.0146 | 0.010127 | 0.44878 | 238.255 | 254.35
Populus deltoides | LI2A |  | RACiR 500 to 0 | 20.2455 | 170.9812 | 1000.15 | 25.009 | 0.010124 | 0.44878 | 236.106 | 251.903
Populus deltoides | LI2A |  | RACiR 500 to 0 | 20.064 | 169.2508 | 999.994 | 25.0077 | 0.010123 | 0.44878 | 234.566 | 250.237
Populus deltoides | LI2A |  | RACiR 500 to 0 | 19.9205 | 168.053 | 999.905 | 25.0062 | 0.010127 | 0.44866 | 232.376 | 247.759
Populus deltoides | LI2A |  | RACiR 500 to 0 | 19.6456 | 166.519 | 1000.15 | 25.004 | 0.010124 | 0.44871 | 230.188 | 245.321
Populus deltoides | LI2A |  | RACiR 500 to 0 | 19.3304 | 165.0962 | 999.869 | 24.9997 | 0.010124 | 0.44879 | 228.626 | 243.599
Populus deltoides | LI2A |  | RACiR 500 to 0 | 19.1303 | 164.0237 | 999.876 | 24.9978 | 0.010124 | 0.4489 | 226.458 | 241.128
Populus deltoides | LI2A |  | RACiR 500 to 0 | 18.7493 | 162.7022 | 999.848 | 25.0012 | 0.010125 | 0.44881 | 224.258 | 238.744
Populus deltoides | LI2A |  | RACiR 500 to 0 | 18.5241 | 161.0621 | 999.896 | 24.999 | 0.010124 | 0.44867 | 222.686 | 236.934
Populus deltoides | LI2A |  | RACiR 500 to 0 | 18.2962 | 160.0132 | 999.835 | 24.9975 | 0.010125 | 0.44909 | 220.475 | 234.618
Populus deltoides | LI2A |  | RACiR 500 to 0 | 18.2618 | 157.8956 | 1000.23 | 24.9968 | 0.010122 | 0.44906 | 218.298 | 232.245
Populus deltoides | LI2A |  | RACiR 500 to 0 | 17.9517 | 156.4316 | 999.997 | 25.0006 | 0.010112 | 0.44869 | 216.787 | 230.577
Populus deltoides | LI2A |  | RACiR 500 to 0 | 17.8468 | 155.153 | 999.844 | 25.0011 | 0.0101 | 0.44832 | 214.633 | 228.127
Populus deltoides | LI2A |  | RACiR 500 to 0 | 17.47 | 153.8728 | 999.902 | 24.9985 | 0.01009 | 0.4482 | 212.475 | 225.635
Populus deltoides | LI2A |  | RACiR 500 to 0 | 17.2156 | 152.2003 | 999.865 | 25.0029 | 0.010089 | 0.4484 | 210.905 | 223.853
Populus deltoides | LI2A |  | RACiR 500 to 0 | 17.0311 | 150.9767 | 1000.03 | 25.006 | 0.01009 | 0.44839 | 208.696 | 221.249
Populus deltoides | LI2A |  | RACiR 500 to 0 | 16.7908 | 149.2662 | 1000.16 | 25.0088 | 0.010096 | 0.44894 | 206.456 | 218.766
Populus deltoides | LI2A |  | RACiR 500 to 0 | 16.6401 | 147.415 | 999.794 | 25.0047 | 0.010097 | 0.44915 | 204.849 | 216.919
Populus deltoides | LI2A |  | RACiR 500 to 0 | 16.4623 | 146.2591 | 1000.21 | 25.0001 | 0.010097 | 0.44925 | 202.596 | 214.614
Populus deltoides | LI2A |  | RACiR 500 to 0 | 16.4432 | 144.0647 | 999.829 | 25.0035 | 0.010105 | 0.44936 | 200.393 | 212.087
Populus deltoides | LI2A |  | RACiR 500 to 0 | 16.1847 | 142.4482 | 999.864 | 25.0097 | 0.010115 | 0.4498 | 198.81 | 210.388
Populus deltoides | LI2A |  | RACiR 500 to 0 | 16.0456 | 141.2396 | 1000.16 | 25.0149 | 0.010127 | 0.4504 | 196.601 | 207.787
Populus deltoides | LI2A |  | RACiR 500 to 0 | 15.6681 | 139.9595 | 1000.17 | 25.0174 | 0.010127 | 0.44991 | 194.365 | 205.28
Populus deltoides | LI2A |  | RACiR 500 to 0 | 15.3925 | 138.4856 | 999.831 | 25.0131 | 0.010122 | 0.44936 | 192.749 | 203.467
Populus deltoides | LI2A |  | RACiR 500 to 0 | 15.217 | 137.3402 | 1000.04 | 25.01 | 0.010119 | 0.44897 | 190.494 | 200.954
Populus deltoides | LI2A |  | RACiR 500 to 0 | 15.0299 | 135.5821 | 1000.18 | 25.0118 | 0.01011 | 0.44884 | 188.269 | 198.559
Populus deltoides | LI2A |  | RACiR 500 to 0 | 14.9428 | 133.6927 | 999.819 | 25.0107 | 0.010101 | 0.44876 | 186.69 | 196.788
Populus deltoides | LI2A |  | RACiR 500 to 0 | 14.751 | 132.6227 | 1000.15 | 25.0132 | 0.010088 | 0.44791 | 184.476 | 194.535
Populus deltoides | LI2A |  | RACiR 500 to 0 | 14.6063 | 130.9241 | 999.815 | 25.0025 | 0.010081 | 0.44718 | 182.302 | 192.086
Populus deltoides | LI2A |  | RACiR 500 to 0 | 14.269 | 129.454 | 999.907 | 25.0039 | 0.010083 | 0.44695 | 180.772 | 190.405
Populus deltoides | LI2A |  | RACiR 500 to 0 | 14.1427 | 128.1841 | 999.843 | 25.0006 | 0.010086 | 0.44706 | 178.6 | 187.84
Populus deltoides | LI2A |  | RACiR 500 to 0 | 13.8249 | 126.6926 | 999.854 | 25.0028 | 0.01009 | 0.44764 | 176.42 | 185.368
Populus deltoides | LI2A |  | RACiR 500 to 0 | 13.5708 | 125.1542 | 1000.09 | 25.002 | 0.010089 | 0.44786 | 174.834 | 183.595
Populus deltoides | LI2A |  | RACiR 500 to 0 | 13.3792 | 124.0753 | 999.948 | 25.0039 | 0.010095 | 0.44799 | 172.604 | 181.162
Populus deltoides | LI2A |  | RACiR 500 to 0 | 13.0931 | 122.5615 | 999.857 | 25.0098 | 0.010117 | 0.44908 | 170.392 | 178.849
Populus deltoides | LI2A |  | RACiR 500 to 0 | 12.9248 | 120.8123 | 1000.18 | 25.0086 | 0.010132 | 0.4495 | 168.837 | 177.093
Populus deltoides | LI2A |  | RACiR 500 to 0 | 12.6996 | 119.7552 | 1000.21 | 25.0137 | 0.010144 | 0.45057 | 166.692 | 174.787
Populus deltoides | LI2A |  | RACiR 500 to 0 | 12.6609 | 117.682 | 999.796 | 25.0143 | 0.010114 | 0.44883 | 164.558 | 172.296
Populus deltoides | LI2A |  | RACiR 500 to 0 | 12.3973 | 116.1566 | 999.864 | 25.02 | 0.010095 | 0.44816 | 163.021 | 170.618
Populus deltoides | LI2A |  | RACiR 500 to 0 | 12.2912 | 114.9188 | 1000.14 | 25.0159 | 0.010075 | 0.44717 | 160.848 | 168.107
Populus deltoides | LI2A |  | RACiR 500 to 0 | 11.9393 | 113.6313 | 1000.14 | 25.0187 | 0.010076 | 0.44749 | 158.676 | 165.665
Populus deltoides | LI2A |  | RACiR 500 to 0 | 11.6259 | 112.2323 | 1000.14 | 25.0188 | 0.010079 | 0.44758 | 157.134 | 163.92
Populus deltoides | LI2A |  | RACiR 500 to 0 | 11.3775 | 111.3054 | 999.926 | 25.0183 | 0.010083 | 0.44752 | 154.945 | 161.429
Populus deltoides | LI2A |  | RACiR 500 to 0 | 10.9614 | 110.1248 | 1000.14 | 25.0129 | 0.010086 | 0.44786 | 152.752 | 159.051
Populus deltoides | LI2A |  | RACiR 500 to 0 | 10.7687 | 108.3183 | 999.812 | 25.0153 | 0.010086 | 0.44749 | 151.188 | 157.25
Populus deltoides | LI2A |  | RACiR 500 to 0 | 10.5911 | 107.0819 | 999.835 | 25.0176 | 0.010087 | 0.44751 | 148.979 | 154.985
Populus deltoides | LI2A |  | RACiR 500 to 0 | 10.5709 | 104.8835 | 1000.13 | 25.0267 | 0.010085 | 0.44703 | 146.856 | 152.583
Populus deltoides | LI2A |  | RACiR 500 to 0 | 10.3849 | 103.221 | 1000.15 | 25.0186 | 0.010085 | 0.44735 | 145.328 | 150.937
Populus deltoides | LI2A |  | RACiR 500 to 0 | 10.2673 | 102.0151 | 999.784 | 25.0174 | 0.010086 | 0.44719 | 143.226 | 148.467
Populus deltoides | LI2A |  | RACiR 500 to 0 | 9.949 | 100.7029 | 1000.13 | 25.0086 | 0.010097 | 0.4476 | 141.063 | 145.976
Populus deltoides | LI2A |  | RACiR 500 to 0 | 9.6067 | 99.3367 | 1000.08 | 25.0164 | 0.010114 | 0.44839 | 139.508 | 144.163
Populus deltoides | LI2A |  | RACiR 500 to 0 | 9.356 | 98.3834 | 1000.09 | 25.0183 | 0.010126 | 0.44907 | 137.324 | 141.549
Populus deltoides | LI2A |  | RACiR 500 to 0 | 8.9528 | 97.1327 | 1000.08 | 25.0243 | 0.010113 | 0.44803 | 135.101 | 139.114
Populus deltoides | LI2A |  | RACiR 500 to 0 | 8.7428 | 95.4126 | 1000.11 | 25.0267 | 0.010091 | 0.4467 | 133.509 | 137.33
Populus deltoides | LI2A |  | RACiR 500 to 0 | 8.4977 | 94.4083 | 1000.08 | 25.0239 | 0.010073 | 0.44514 | 131.303 | 135.083
Populus deltoides | LI2A |  | RACiR 500 to 0 | 8.4286 | 92.3749 | 1000.13 | 25.0166 | 0.010071 | 0.44579 | 129.141 | 132.719
Populus deltoides | LI2A |  | RACiR 500 to 0 | 8.0963 | 90.9915 | 1000.12 | 25.0074 | 0.010075 | 0.44619 | 127.619 | 131.068
Populus deltoides | LI2A |  | RACiR 500 to 0 | 7.9695 | 89.7494 | 1000.1 | 25.0065 | 0.010079 | 0.44712 | 125.504 | 128.619
Populus deltoides | LI2A |  | RACiR 500 to 0 | 7.6483 | 88.3368 | 999.76 | 25.0073 | 0.010082 | 0.44683 | 123.349 | 126.106
Populus deltoides | LI2A |  | RACiR 500 to 0 | 7.3851 | 86.7655 | 1000.03 | 25.0125 | 0.010085 | 0.44686 | 121.805 | 124.3
Populus deltoides | LI2A |  | RACiR 500 to 0 | 7.196 | 85.6448 | 1000.06 | 25.014 | 0.010085 | 0.44654 | 119.63 | 121.706
Populus deltoides | LI2A |  | RACiR 500 to 0 | 6.9218 | 84.0664 | 1000.07 | 25.0143 | 0.01008 | 0.44615 | 117.413 | 119.261
Populus deltoides | LI2A |  | RACiR 500 to 0 | 6.821 | 82.0848 | 1000.05 | 25.0115 | 0.010067 | 0.44576 | 115.846 | 117.436
Populus deltoides | LI2A |  | RACiR 500 to 0 | 6.6288 | 80.962 | 999.768 | 25.0138 | 0.010056 | 0.44588 | 113.601 | 115.103
Populus deltoides | LI2A |  | RACiR 500 to 0 | 6.5139 | 79.0787 | 1000.11 | 25.0106 | 0.010052 | 0.44651 | 111.402 | 112.555
Populus deltoides | LI2A |  | RACiR 500 to 0 | 6.1917 | 77.7124 | 999.864 | 25.0084 | 0.010061 | 0.44714 | 109.818 | 110.833
Populus deltoides | LI2A |  | RACiR 500 to 0 | 6.0442 | 76.5416 | 1000.1 | 25.0077 | 0.010088 | 0.44842 | 107.597 | 108.277
Populus deltoides | LI2A |  | RACiR 500 to 0 | 5.6802 | 75.2214 | 1000.06 | 25.0102 | 0.010108 | 0.44876 | 105.386 | 105.788
Populus deltoides | LI2A |  | RACiR 500 to 0 | 5.3787 | 73.7305 | 999.73 | 25.0111 | 0.010118 | 0.44891 | 103.81 | 104.012
Populus deltoides | LI2A |  | RACiR 500 to 0 | 5.1892 | 72.5812 | 999.834 | 25.0098 | 0.010108 | 0.44806 | 101.609 | 101.519
Populus deltoides | LI2A |  | RACiR 500 to 0 | 4.9383 | 70.9667 | 1000.06 | 25.0095 | 0.010094 | 0.44746 | 99.3756 | 99.1661
Populus deltoides | LI2A |  | RACiR 500 to 0 | 4.8016 | 69.0948 | 999.783 | 25.0107 | 0.010084 | 0.4468 | 97.8158 | 97.4013
Populus deltoides | LI2A |  | RACiR 500 to 0 | 4.5876 | 68.0613 | 1000.09 | 25.0113 | 0.010075 | 0.44672 | 95.6307 | 95.0836
Populus deltoides | LI2A |  | RACiR 500 to 0 | 4.5049 | 66.1017 | 1000.06 | 25.0134 | 0.010073 | 0.44681 | 93.5001 | 92.6145
Populus deltoides | LI2A |  | RACiR 500 to 0 | 4.1982 | 64.7018 | 1000.1 | 25.0152 | 0.010075 | 0.44708 | 91.9595 | 90.9148
Populus deltoides | LI2A |  | RACiR 500 to 0 | 4.0673 | 63.514 | 1000.06 | 25.0116 | 0.010077 | 0.447 | 89.7888 | 88.3641
Populus deltoides | LI2A |  | RACiR 500 to 0 | 3.7221 | 62.2027 | 999.836 | 25.0073 | 0.010078 | 0.44701 | 87.5815 | 85.8387
Populus deltoides | LI2A |  | RACiR 500 to 0 | 3.4569 | 60.6803 | 1000.1 | 25.004 | 0.010067 | 0.44664 | 85.9867 | 84.0413
Populus deltoides | LI2A |  | RACiR 500 to 0 | 3.2753 | 59.5435 | 1000.07 | 25.0074 | 0.010044 | 0.44565 | 83.7918 | 81.5768
Populus deltoides | LI2A |  | RACiR 500 to 0 | 2.9911 | 58.0366 | 1000.06 | 25.0095 | 0.010026 | 0.44505 | 81.5811 | 79.2112
Populus deltoides | LI2A |  | RACiR 500 to 0 | 2.7712 | 56.4031 | 999.721 | 25.0065 | 0.010018 | 0.44482 | 80.0424 | 77.4464
Populus deltoides | LI2A |  | RACiR 500 to 0 | 2.5084 | 55.4821 | 1000.09 | 25.0023 | 0.010023 | 0.44519 | 77.8412 | 75.146
Populus deltoides | LI2A |  | RACiR 500 to 0 | 2.3631 | 53.6951 | 1000.08 | 25.0011 | 0.010036 | 0.44581 | 75.6991 | 72.6818
Populus deltoides | LI2A |  | RACiR 500 to 0 | 2.068 | 52.2411 | 999.759 | 25.004 | 0.010057 | 0.44718 | 74.1576 | 70.9995
Populus deltoides | LI2A |  | RACiR 500 to 0 | 1.9137 | 51.0891 | 1000.08 | 25.0063 | 0.010081 | 0.44861 | 72.0078 | 68.4902
Populus deltoides | LI2A |  | RACiR 500 to 0 | 1.5958 | 49.6986 | 999.739 | 25.0124 | 0.010093 | 0.44917 | 69.8429 | 66.0444
Populus deltoides | LI2A |  | RACiR 500 to 0 | 1.301 | 48.2341 | 1000.03 | 25.0162 | 0.010089 | 0.44855 | 68.283 | 64.2909
Populus deltoides | LI2A |  | RACiR 300 to 800 | 26.0738 | 237.6076 | 1000.05 | 24.9384 | 0.010272 | 0.45884 | 300.456 | 338.654
Populus deltoides | LI2A |  | RACiR 300 to 800 | 26.2457 | 238.6732 | 1000.06 | 24.9326 | 0.010274 | 0.45938 | 301.851 | 340.226
Populus deltoides | LI2A |  | RACiR 300 to 800 | 26.4905 | 239.7276 | 999.96 | 24.9286 | 0.010274 | 0.45965 | 303.43 | 342.063
Populus deltoides | LI2A |  | RACiR 300 to 800 | 26.5624 | 240.98 | 999.982 | 24.9301 | 0.010276 | 0.4597 | 304.863 | 343.577
Populus deltoides | LI2A |  | RACiR 300 to 800 | 26.5903 | 242.4595 | 1000.04 | 24.9338 | 0.010276 | 0.45943 | 306.475 | 345.226
Populus deltoides | LI2A |  | RACiR 300 to 800 | 26.7455 | 243.478 | 999.722 | 24.9368 | 0.010275 | 0.4592 | 307.89 | 346.806
Populus deltoides | LI2A |  | RACiR 300 to 800 | 26.8369 | 244.7942 | 1000.03 | 24.9439 | 0.010282 | 0.459 | 309.47 | 348.493
Populus deltoides | LI2A |  | RACiR 300 to 800 | 26.8977 | 246.0998 | 1000.04 | 24.9476 | 0.010295 | 0.4593 | 310.905 | 350
Populus deltoides | LI2A |  | RACiR 300 to 800 | 27.1885 | 247.0454 | 1000.08 | 24.95 | 0.01031 | 0.45969 | 312.465 | 351.865
Populus deltoides | LI2A |  | RACiR 300 to 800 | 27.3456 | 248.1666 | 999.996 | 24.9492 | 0.010322 | 0.46027 | 313.882 | 353.454
Populus deltoides | LI2A |  | RACiR 300 to 800 | 27.4951 | 249.4199 | 1000.01 | 24.949 | 0.01033 | 0.46046 | 315.468 | 355.199
Populus deltoides | LI2A |  | RACiR 300 to 800 | 27.7744 | 250.1184 | 1000.09 | 24.9494 | 0.01032 | 0.45975 | 316.894 | 356.909
Populus deltoides | LI2A |  | RACiR 300 to 800 | 27.8904 | 251.3389 | 1000.06 | 24.9482 | 0.010306 | 0.45909 | 318.486 | 358.626
Populus deltoides | LI2A |  | RACiR 300 to 800 | 27.9761 | 252.525 | 1000.05 | 24.9469 | 0.010296 | 0.45864 | 319.949 | 360.178
Populus deltoides | LI2A |  | RACiR 300 to 800 | 28.2092 | 253.5356 | 1000.02 | 24.9442 | 0.010284 | 0.4583 | 321.539 | 362.01
Populus deltoides | LI2A |  | RACiR 300 to 800 | 28.2564 | 254.9029 | 999.978 | 24.9399 | 0.010281 | 0.45861 | 322.994 | 363.523
Populus deltoides | LI2A |  | RACiR 300 to 800 | 28.3074 | 256.4237 | 999.999 | 24.9375 | 0.010281 | 0.45884 | 324.627 | 365.217
Populus deltoides | LI2A |  | RACiR 300 to 800 | 28.5413 | 257.3289 | 999.979 | 24.9393 | 0.010283 | 0.45887 | 326.065 | 366.895
Populus deltoides | LI2A |  | RACiR 300 to 800 | 28.6511 | 258.6609 | 999.967 | 24.9434 | 0.010284 | 0.45869 | 327.697 | 368.654
Populus deltoides | LI2A |  | RACiR 300 to 800 | 28.7825 | 259.8357 | 1000.02 | 24.9438 | 0.010287 | 0.45883 | 329.169 | 370.266
Populus deltoides | LI2A |  | RACiR 300 to 800 | 29.1155 | 260.6795 | 1000.05 | 24.9456 | 0.01029 | 0.45889 | 330.758 | 372.195
Populus deltoides | LI2A |  | RACiR 300 to 800 | 29.2017 | 261.9894 | 1000.01 | 24.943 | 0.010292 | 0.45926 | 332.232 | 373.77
Populus deltoides | LI2A |  | RACiR 300 to 800 | 29.2441 | 263.5008 | 999.742 | 24.9444 | 0.010292 | 0.45924 | 333.874 | 375.464
Populus deltoides | LI2A |  | RACiR 300 to 800 | 29.4062 | 264.5314 | 999.989 | 24.949 | 0.010292 | 0.45894 | 335.332 | 377.096
Populus deltoides | LI2A |  | RACiR 300 to 800 | 29.4504 | 266.0268 | 999.998 | 24.9505 | 0.010293 | 0.45893 | 336.96 | 378.78
Populus deltoides | LI2A |  | RACiR 300 to 800 | 29.5005 | 267.3202 | 1000.1 | 24.953 | 0.010293 | 0.45877 | 338.419 | 380.293
Populus deltoides | LI2A |  | RACiR 300 to 800 | 29.7773 | 268.2915 | 1000.06 | 24.9529 | 0.010293 | 0.45885 | 340.01 | 382.17
Populus deltoides | LI2A |  | RACiR 300 to 800 | 29.8486 | 269.6158 | 999.977 | 24.9521 | 0.010295 | 0.45907 | 341.488 | 383.735
Populus deltoides | LI2A |  | RACiR 300 to 800 | 29.9584 | 271.0041 | 1000.04 | 24.9527 | 0.010297 | 0.4592 | 343.13 | 385.495
Populus deltoides | LI2A |  | RACiR 300 to 800 | 30.2113 | 271.904 | 999.695 | 24.9538 | 0.0103 | 0.45929 | 344.595 | 387.224
Populus deltoides | LI2A |  | RACiR 300 to 800 | 30.2808 | 273.3671 | 1000.05 | 24.9557 | 0.010299 | 0.45918 | 346.262 | 388.977
Populus deltoides | LI2A |  | RACiR 300 to 800 | 30.3747 | 274.6233 | 1000.05 | 24.9569 | 0.010303 | 0.45935 | 347.729 | 390.541
Populus deltoides | LI2A |  | RACiR 300 to 800 | 30.5679 | 275.7751 | 1000.01 | 24.9598 | 0.010303 | 0.45915 | 349.364 | 392.376
Populus deltoides | LI2A |  | RACiR 300 to 800 | 30.5696 | 277.185 | 999.983 | 24.9631 | 0.010303 | 0.45892 | 350.846 | 393.877
Populus deltoides | LI2A |  | RACiR 300 to 800 | 30.5763 | 278.7235 | 1000.02 | 24.9692 | 0.010306 | 0.45869 | 352.475 | 395.522
Populus deltoides | LI2A |  | RACiR 300 to 800 | 30.7173 | 279.799 | 1000.06 | 24.9736 | 0.010306 | 0.45839 | 353.935 | 397.133
Populus deltoides | LI2A |  | RACiR 300 to 800 | 30.7611 | 281.2465 | 1000.08 | 24.9759 | 0.010305 | 0.45823 | 355.537 | 398.794
Populus deltoides | LI2A |  | RACiR 300 to 800 | 30.8365 | 282.477 | 999.729 | 24.9771 | 0.010307 | 0.45831 | 356.951 | 400.289
Populus deltoides | LI2A |  | RACiR 300 to 800 | 31.101 | 283.4774 | 1000.08 | 24.9765 | 0.010308 | 0.45838 | 358.546 | 402.158
Populus deltoides | LI2A |  | RACiR 300 to 800 | 31.2162 | 284.6491 | 1000.04 | 24.9756 | 0.010307 | 0.4585 | 359.978 | 403.719
Populus deltoides | LI2A |  | RACiR 300 to 800 | 31.3348 | 286.0036 | 1000.02 | 24.9738 | 0.010309 | 0.45878 | 361.579 | 405.452
Populus deltoides | LI2A |  | RACiR 300 to 800 | 31.6024 | 286.8737 | 1000.05 | 24.9721 | 0.010309 | 0.45893 | 363.034 | 407.184
Populus deltoides | LI2A |  | RACiR 300 to 800 | 31.7163 | 288.2459 | 1000.05 | 24.9692 | 0.01031 | 0.45922 | 364.64 | 408.915
Populus deltoides | LI2A |  | RACiR 300 to 800 | 31.791 | 289.52 | 999.887 | 24.9695 | 0.010311 | 0.4593 | 366.095 | 410.448
Populus deltoides | LI2A |  | RACiR 300 to 800 | 31.9921 | 290.6704 | 1000.03 | 24.9716 | 0.010311 | 0.45916 | 367.738 | 412.304
Populus deltoides | LI2A |  | RACiR 300 to 800 | 32.024 | 292.0101 | 1000.07 | 24.9757 | 0.010313 | 0.45898 | 369.211 | 413.817
Populus deltoides | LI2A |  | RACiR 300 to 800 | 32.0115 | 293.6079 | 1000.04 | 24.9802 | 0.010312 | 0.45863 | 370.876 | 415.481
Populus deltoides | LI2A |  | RACiR 300 to 800 | 32.2381 | 294.5277 | 1000.02 | 24.9808 | 0.010311 | 0.45854 | 372.332 | 417.176
Populus deltoides | LI2A |  | RACiR 300 to 800 | 32.3456 | 295.9056 | 1000.05 | 24.9801 | 0.010312 | 0.45867 | 373.956 | 418.922
Populus deltoides | LI2A |  | RACiR 300 to 800 | 32.4525 | 297.1294 | 999.672 | 24.9809 | 0.010314 | 0.45873 | 375.434 | 420.512
Populus deltoides | LI2A |  | RACiR 300 to 800 | 32.7545 | 298.0983 | 1000.08 | 24.9805 | 0.010317 | 0.45897 | 377.052 | 422.443
Populus deltoides | LI2A |  | RACiR 300 to 800 | 32.8026 | 299.4544 | 1000.03 | 24.9825 | 0.010319 | 0.459 | 378.54 | 423.992
Populus deltoides | LI2A |  | RACiR 300 to 800 | 32.8284 | 301.0201 | 1000.03 | 24.985 | 0.010323 | 0.45909 | 380.186 | 425.675
Populus deltoides | LI2A |  | RACiR 300 to 800 | 33.0106 | 302.1623 | 1000.04 | 24.9842 | 0.010336 | 0.45981 | 381.639 | 427.32
Populus deltoides | LI2A |  | RACiR 300 to 800 | 33.0495 | 303.8485 | 1000.09 | 24.9802 | 0.010353 | 0.46095 | 383.263 | 429.006
Populus deltoides | LI2A |  | RACiR 300 to 800 | 33.1163 | 305.2833 | 999.848 | 24.9794 | 0.01037 | 0.46177 | 384.748 | 430.572
Populus deltoides | LI2A |  | RACiR 300 to 800 | 33.3651 | 306.4393 | 1000.03 | 24.9766 | 0.010384 | 0.46241 | 386.373 | 432.464
Populus deltoides | LI2A |  | RACiR 300 to 800 | 33.4054 | 307.8431 | 1000.04 | 24.9734 | 0.010387 | 0.46269 | 387.851 | 433.998
Populus deltoides | LI2A |  | RACiR 300 to 800 | 33.5345 | 309.0524 | 1000.04 | 24.9764 | 0.01038 | 0.46198 | 389.489 | 435.772
Populus deltoides | LI2A |  | RACiR 300 to 800 | 33.7677 | 309.8706 | 1000.05 | 24.9778 | 0.010371 | 0.46125 | 390.958 | 437.477
Populus deltoides | LI2A |  | RACiR 300 to 800 | 33.8657 | 311.1553 | 1000.03 | 24.9792 | 0.010362 | 0.46063 | 392.591 | 439.22
Populus deltoides | LI2A |  | RACiR 300 to 800 | 33.989 | 312.271 | 1000.07 | 24.9832 | 0.01036 | 0.46021 | 394.076 | 440.83
Populus deltoides | LI2A |  | RACiR 300 to 800 | 34.1585 | 313.5302 | 999.815 | 24.982 | 0.010358 | 0.46018 | 395.74 | 442.678
Populus deltoides | LI2A |  | RACiR 300 to 800 | 34.1313 | 315.1024 | 1000.02 | 24.9793 | 0.010359 | 0.46046 | 397.24 | 444.162
Populus deltoides | LI2A |  | RACiR 300 to 800 | 34.1006 | 316.8695 | 1000.06 | 24.9777 | 0.010363 | 0.46091 | 398.902 | 445.803
Populus deltoides | LI2A |  | RACiR 300 to 800 | 34.2255 | 318.0906 | 1000.06 | 24.9751 | 0.010363 | 0.4612 | 400.373 | 447.409
Populus deltoides | LI2A |  | RACiR 300 to 800 | 34.2433 | 319.653 | 1000.08 | 24.9735 | 0.010361 | 0.46125 | 401.998 | 449.066
Populus deltoides | LI2A |  | RACiR 300 to 800 | 34.3146 | 320.9083 | 999.868 | 24.9753 | 0.010363 | 0.46126 | 403.439 | 450.583
Populus deltoides | LI2A |  | RACiR 300 to 800 | 34.6057 | 321.7745 | 1000.07 | 24.9781 | 0.010361 | 0.46086 | 405.034 | 452.478
Populus deltoides | LI2A |  | RACiR 300 to 800 | 34.7315 | 322.8536 | 1000.08 | 24.9806 | 0.010358 | 0.46047 | 406.482 | 454.065
Populus deltoides | LI2A |  | RACiR 300 to 800 | 34.8624 | 324.1162 | 1000.08 | 24.9849 | 0.01036 | 0.46026 | 408.103 | 455.82
Populus deltoides | LI2A |  | RACiR 300 to 800 | 35.1051 | 324.9323 | 1000.09 | 24.9882 | 0.010358 | 0.45981 | 409.551 | 457.525
Populus deltoides | LI2A |  | RACiR 300 to 800 | 35.1337 | 326.4353 | 999.664 | 24.9878 | 0.010352 | 0.45958 | 411.189 | 459.202
Populus deltoides | LI2A |  | RACiR 300 to 800 | 35.1463 | 327.8706 | 1000.13 | 24.9866 | 0.010348 | 0.45955 | 412.689 | 460.721
Populus deltoides | LI2A |  | RACiR 300 to 800 | 35.3506 | 328.9846 | 1000.07 | 24.9844 | 0.010341 | 0.4594 | 414.299 | 462.544
Populus deltoides | LI2A |  | RACiR 300 to 800 | 35.3464 | 330.5194 | 1000.02 | 24.9804 | 0.010339 | 0.45982 | 415.78 | 464.033
Populus deltoides | LI2A |  | RACiR 300 to 800 | 35.436 | 331.9241 | 1000.04 | 24.9821 | 0.010339 | 0.4598 | 417.417 | 465.769
Populus deltoides | LI2A |  | RACiR 300 to 800 | 35.6356 | 332.9346 | 999.759 | 24.986 | 0.010343 | 0.45983 | 418.887 | 467.448
Populus deltoides | LI2A |  | RACiR 300 to 800 | 35.7482 | 334.2698 | 1000.08 | 24.9924 | 0.010346 | 0.45966 | 420.536 | 469.227
Populus deltoides | LI2A |  | RACiR 300 to 800 | 35.8724 | 335.3843 | 1000.08 | 24.9968 | 0.010346 | 0.45931 | 422.014 | 470.836
Populus deltoides | LI2A |  | RACiR 300 to 800 | 36.1517 | 336.4034 | 1000.06 | 24.9983 | 0.010349 | 0.45943 | 423.651 | 472.76
Populus deltoides | LI2A |  | RACiR 300 to 800 | 36.1827 | 337.7769 | 1000.09 | 24.9999 | 0.010348 | 0.45928 | 425.153 | 474.31
Populus deltoides | LI2A |  | RACiR 300 to 800 | 36.2198 | 339.3616 | 1000.09 | 25.0002 | 0.010351 | 0.45951 | 426.813 | 476.014
Populus deltoides | LI2A |  | RACiR 300 to 800 | 36.3936 | 341.317 | 1000.05 | 24.9993 | 0.010354 | 0.45984 | 429.136 | 478.524
Populus deltoides | LI2A |  | RACiR 300 to 800 | 36.3936 | 341.324 | 1000.04 | 24.9993 | 0.010354 | 0.45987 | 429.136 | 478.524
Populus deltoides | LI2A |  | RACiR 300 to 800 | 36.2494 | 343.8848 | 1000.06 | 24.9998 | 0.010355 | 0.45996 | 431.423 | 480.69
Populus deltoides | LI2A |  | RACiR 300 to 800 | 36.3818 | 345.2228 | 999.796 | 25 | 0.010358 | 0.46012 | 433.052 | 482.461
Populus deltoides | LI2A |  | RACiR 300 to 800 | 36.5176 | 346.2656 | 1000.08 | 25.003 | 0.010356 | 0.4598 | 434.479 | 484.04
Populus deltoides | LI2A |  | RACiR 300 to 800 | 36.733 | 347.3706 | 1000.06 | 25.0062 | 0.010362 | 0.45989 | 436.069 | 485.837
Populus deltoides | LI2A |  | RACiR 300 to 800 | 37.1237 | 347.9475 | 1000.04 | 25.0078 | 0.010361 | 0.45982 | 437.532 | 487.708
Populus deltoides | LI2A |  | RACiR 300 to 800 | 37.3765 | 349.0216 | 1000.1 | 25.0083 | 0.010359 | 0.45974 | 439.202 | 489.648
Populus deltoides | LI2A |  | RACiR 300 to 800 | 37.3886 | 350.4949 | 1000.1 | 25.0106 | 0.01036 | 0.45966 | 440.751 | 491.215
Populus deltoides | LI2A |  | RACiR 300 to 800 | 37.4405 | 352.0932 | 1000.02 | 25.0096 | 0.010361 | 0.4598 | 442.474 | 493
Populus deltoides | LI2A |  | RACiR 300 to 800 | 37.4172 | 353.5985 | 1000.06 | 25.0097 | 0.01036 | 0.45976 | 443.968 | 494.486
Populus deltoides | LI2A |  | RACiR 300 to 800 | 37.3207 | 355.4662 | 999.722 | 25.0101 | 0.010361 | 0.45985 | 445.651 | 496.086
Populus deltoides | LI2A |  | RACiR 300 to 800 | 37.4425 | 356.6712 | 1000.08 | 25.0074 | 0.010361 | 0.46007 | 447.106 | 497.679
Populus deltoides | LI2A |  | RACiR 300 to 800 | 37.5102 | 358.0762 | 1000.04 | 25.0119 | 0.010364 | 0.45992 | 448.724 | 499.374
Populus deltoides | LI2A |  | RACiR 300 to 800 | 37.591 | 359.3233 | 999.732 | 25.016 | 0.010368 | 0.45991 | 450.184 | 500.922
Populus deltoides | LI2A |  | RACiR 300 to 800 | 37.8227 | 360.3931 | 1000.07 | 25.0186 | 0.010369 | 0.45977 | 451.817 | 502.795
Populus deltoides | LI2A |  | RACiR 300 to 800 | 37.9546 | 361.6552 | 1000.06 | 25.0186 | 0.010382 | 0.46039 | 453.285 | 504.414
Populus deltoides | LI2A |  | RACiR 300 to 800 | 38.0295 | 363.2994 | 1000.03 | 25.0168 | 0.010401 | 0.46132 | 454.96 | 506.184
Populus deltoides | LI2A |  | RACiR 300 to 800 | 38.1958 | 364.5661 | 1000.06 | 25.0111 | 0.010415 | 0.46217 | 456.462 | 507.868
Populus deltoides | LI2A |  | RACiR 300 to 800 | 38.2487 | 366.1945 | 1000.08 | 25.0073 | 0.01043 | 0.46295 | 458.1 | 509.577
Populus deltoides | LI2A |  | RACiR 300 to 800 | 38.228 | 367.6588 | 1000.07 | 25.0084 | 0.010432 | 0.46269 | 459.605 | 511.071
Populus deltoides | LI2A |  | RACiR 300 to 800 | 38.375 | 368.8998 | 1000.07 | 25.0034 | 0.010425 | 0.46245 | 461.238 | 512.856
Populus deltoides | LI2A |  | RACiR 300 to 800 | 38.348 | 370.3514 | 999.738 | 25.0021 | 0.010417 | 0.46198 | 462.753 | 514.356
Populus deltoides | LI2A |  | RACiR 300 to 800 | 38.3943 | 371.7863 | 1000.07 | 24.9991 | 0.01041 | 0.46161 | 464.389 | 516.043
Populus deltoides | LI2A |  | RACiR 300 to 800 | 38.6373 | 372.6873 | 1000.05 | 24.9939 | 0.010405 | 0.46166 | 465.83 | 517.738
Populus deltoides | LI2A |  | RACiR 300 to 800 | 38.7387 | 374.0631 | 1000.07 | 24.9911 | 0.0104 | 0.46145 | 467.496 | 519.518
Populus deltoides | LI2A |  | RACiR 300 to 800 | 38.8781 | 375.1556 | 1000.05 | 24.9925 | 0.0104 | 0.46126 | 468.959 | 521.121
Populus deltoides | LI2A |  | RACiR 300 to 800 | 39.1197 | 376.2782 | 999.994 | 24.9895 | 0.010398 | 0.46134 | 470.619 | 523.033
Populus deltoides | LI2A |  | RACiR 300 to 800 | 39.138 | 377.6643 | 1000.07 | 24.9878 | 0.010393 | 0.46112 | 472.116 | 524.56
Populus deltoides | LI2A |  | RACiR 300 to 800 | 39.1844 | 379.1801 | 1000.05 | 24.986 | 0.010392 | 0.46115 | 473.762 | 526.258
Populus deltoides | LI2A |  | RACiR 300 to 800 | 39.3238 | 380.3964 | 1000.03 | 24.9799 | 0.010389 | 0.46143 | 475.25 | 527.894
Populus deltoides | LI2A |  | RACiR 300 to 800 | 39.2995 | 382.052 | 999.701 | 24.9794 | 0.010387 | 0.46133 | 476.908 | 529.544
Populus deltoides | LI2A |  | RACiR 300 to 800 | 39.3808 | 383.2793 | 999.891 | 24.9827 | 0.010387 | 0.46114 | 478.382 | 531.105
Populus deltoides | LI2A |  | RACiR 300 to 800 | 39.5821 | 384.3806 | 1000.02 | 24.9861 | 0.010385 | 0.46076 | 480.029 | 532.971
Populus deltoides | LI2A |  | RACiR 300 to 800 | 39.6402 | 385.6811 | 1000.12 | 24.9869 | 0.010384 | 0.46065 | 481.509 | 534.516
Populus deltoides | LI2A |  | RACiR 300 to 800 | 39.7352 | 387.1174 | 1000.01 | 24.9857 | 0.010385 | 0.46079 | 483.158 | 536.268
Populus deltoides | LI2A |  | RACiR 300 to 800 | 39.946 | 388.2101 | 1000.1 | 24.9803 | 0.010385 | 0.46117 | 484.657 | 537.985
Populus deltoides | LI2A |  | RACiR 300 to 800 | 39.9755 | 389.855 | 1000.07 | 24.9754 | 0.010384 | 0.46146 | 486.341 | 539.716
Populus deltoides | LI2A |  | RACiR 300 to 800 | 40.0312 | 391.2003 | 1000.11 | 24.9764 | 0.010385 | 0.46141 | 487.85 | 541.284
Populus deltoides | LI2A |  | RACiR 300 to 800 | 40.2509 | 392.309 | 999.78 | 24.9765 | 0.010383 | 0.46131 | 489.485 | 543.152
Populus deltoides | LI2A |  | RACiR 300 to 800 | 40.1953 | 393.8889 | 1000.06 | 24.9789 | 0.010382 | 0.4611 | 491.024 | 544.648
Populus deltoides | LI2A |  | RACiR 300 to 800 | 40.1752 | 395.5135 | 1000.06 | 24.9781 | 0.010378 | 0.46093 | 492.674 | 546.287
Populus deltoides | LI2A |  | RACiR 300 to 800 | 40.2948 | 396.7382 | 1000.08 | 24.9763 | 0.010378 | 0.46111 | 494.151 | 547.9
Populus deltoides | LI2A |  | RACiR 300 to 800 | 40.2845 | 398.4202 | 999.774 | 24.9742 | 0.010378 | 0.46122 | 495.825 | 549.57
Populus deltoides | LI2A |  | RACiR 300 to 800 | 40.336 | 399.8169 | 999.918 | 24.9696 | 0.010377 | 0.46152 | 497.302 | 551.103
Populus deltoides | LI2A |  | RACiR 300 to 800 | 40.5799 | 400.9431 | 1000.1 | 24.9675 | 0.010378 | 0.46176 | 498.936 | 552.997
Populus deltoides | LI2A |  | RACiR 300 to 800 | 40.6399 | 402.3019 | 1000.06 | 24.967 | 0.010375 | 0.46167 | 500.475 | 554.605
Populus deltoides | LI2A |  | RACiR 300 to 800 | 40.7041 | 403.7328 | 1000.06 | 24.9699 | 0.010374 | 0.46135 | 502.146 | 556.346
Populus deltoides | LI2A |  | RACiR 300 to 800 | 40.9368 | 404.6162 | 1000.08 | 24.9744 | 0.010375 | 0.46105 | 503.625 | 558.066
Populus deltoides | LI2A |  | RACiR 300 to 800 | 40.9518 | 406.1167 | 1000.04 | 24.9806 | 0.010373 | 0.46048 | 505.316 | 559.792
Populus deltoides | LI2A |  | RACiR 300 to 800 | 40.976 | 407.5181 | 1000.11 | 24.98 | 0.010374 | 0.46054 | 506.79 | 561.295
Populus deltoides | LI2A |  | RACiR 300 to 800 | 41.1098 | 408.9528 | 999.813 | 24.9769 | 0.010377 | 0.46097 | 508.46 | 563.109
Populus deltoides | LI2A |  | RACiR 300 to 800 | 41.0786 | 410.545 | 1000.08 | 24.9728 | 0.010374 | 0.46119 | 509.969 | 564.605
Populus deltoides | LI2A |  | RACiR 300 to 800 | 41.0259 | 412.3655 | 1000.07 | 24.967 | 0.010373 | 0.46154 | 511.638 | 566.226
Populus deltoides | LI2A |  | RACiR 300 to 800 | 41.248 | 413.3408 | 1000.02 | 24.9673 | 0.010374 | 0.46162 | 513.106 | 567.927
Populus deltoides | LI2A |  | RACiR 300 to 800 | 41.3465 | 414.7841 | 999.754 | 24.9662 | 0.010374 | 0.46168 | 514.783 | 569.716
Populus deltoides | LI2A |  | RACiR 300 to 800 | 41.4405 | 416.0812 | 999.88 | 24.9659 | 0.010375 | 0.46183 | 516.285 | 571.317
Populus deltoides | LI2A |  | RACiR 300 to 800 | 41.7347 | 417.068 | 1000.07 | 24.9694 | 0.010379 | 0.46181 | 517.946 | 573.281
Populus deltoides | LI2A |  | RACiR 300 to 800 | 41.7395 | 418.4938 | 1000.1 | 24.9724 | 0.010377 | 0.46152 | 519.478 | 574.833
Populus deltoides | LI2A |  | RACiR 300 to 800 | 41.6922 | 420.3178 | 1000.07 | 24.9705 | 0.010379 | 0.46176 | 521.188 | 576.508
Populus deltoides | LI2A |  | RACiR 300 to 800 | 41.8199 | 421.489 | 1000.02 | 24.9713 | 0.010378 | 0.46168 | 522.683 | 578.146
Populus deltoides | LI2A |  | RACiR 300 to 800 | 41.7857 | 423.2795 | 1000.1 | 24.9684 | 0.010378 | 0.46194 | 524.382 | 579.822
Populus deltoides | LI2A |  | RACiR 300 to 800 | 41.7602 | 424.9039 | 1000.14 | 24.9652 | 0.01038 | 0.46233 | 525.903 | 581.324
Populus deltoides | LI2A |  | RACiR 300 to 800 | 41.9978 | 425.9545 | 999.99 | 24.9678 | 0.01038 | 0.46212 | 527.544 | 583.216
Populus deltoides | LI2A |  | RACiR 300 to 800 | 42.01 | 427.4531 | 1000.08 | 24.9676 | 0.01038 | 0.46219 | 529.087 | 584.783
Populus deltoides | LI2A |  | RACiR 300 to 800 | 42.0369 | 429.0554 | 1000.06 | 24.9663 | 0.010379 | 0.46227 | 530.767 | 586.496
Populus deltoides | LI2A |  | RACiR 300 to 800 | 42.2471 | 430.1094 | 1000.09 | 24.9652 | 0.01038 | 0.4624 | 532.276 | 588.227
Populus deltoides | LI2A |  | RACiR 300 to 800 | 42.2729 | 431.8972 | 1000.06 | 24.962 | 0.010394 | 0.46334 | 533.967 | 589.967
Populus deltoides | LI2A |  | RACiR 300 to 800 | 42.2822 | 433.4363 | 1000.13 | 24.9647 | 0.01041 | 0.46376 | 535.482 | 591.509
Populus deltoides | LI2A |  | RACiR 300 to 800 | 42.4569 | 434.7955 | 1000.03 | 24.9677 | 0.010428 | 0.46428 | 537.159 | 593.38
Populus deltoides | LI2A |  | RACiR 300 to 800 | 42.3819 | 436.4571 | 1000.09 | 24.9709 | 0.01044 | 0.46431 | 538.699 | 594.864
Populus deltoides | LI2A |  | RACiR 300 to 800 | 42.3001 | 438.1851 | 999.72 | 24.9731 | 0.010441 | 0.46389 | 540.383 | 596.474
Populus deltoides | LI2A |  | RACiR 300 to 800 | 42.3982 | 439.3676 | 1000.07 | 24.9703 | 0.010437 | 0.4636 | 541.867 | 598.059
Populus deltoides | LI2A |  | RACiR 300 to 800 | 42.3735 | 440.9568 | 1000 | 24.9677 | 0.010427 | 0.46314 | 543.531 | 599.714
Populus deltoides | LI2A |  | RACiR 300 to 800 | 42.4429 | 442.1877 | 999.737 | 24.9673 | 0.010422 | 0.46283 | 545.005 | 601.262
Populus deltoides | LI2A |  | RACiR 300 to 800 | 42.6637 | 443.2489 | 1000.09 | 24.969 | 0.010418 | 0.46241 | 546.663 | 603.145
Populus deltoides | LI2A |  | RACiR 300 to 800 | 42.7604 | 444.4627 | 1000.12 | 24.9669 | 0.010412 | 0.46226 | 548.144 | 604.74
Populus deltoides | LI2A |  | RACiR 300 to 800 | 42.86 | 445.9149 | 1000.07 | 24.965 | 0.010411 | 0.46239 | 549.818 | 606.517
Populus deltoides | LI2A |  | RACiR 300 to 800 | 43.0756 | 446.925 | 1000.06 | 24.9632 | 0.01041 | 0.4624 | 551.318 | 608.241
Populus deltoides | LI2A |  | RACiR 300 to 800 | 43.1238 | 448.486 | 1000.06 | 24.9589 | 0.010406 | 0.46246 | 553.003 | 609.989
Populus deltoides | LI2A |  | RACiR 300 to 800 | 43.1045 | 449.9506 | 1000.09 | 24.9625 | 0.010405 | 0.4621 | 554.539 | 611.511
Populus deltoides | LI2A |  | RACiR 300 to 800 | 43.245 | 451.2664 | 1000.04 | 24.9653 | 0.010406 | 0.46193 | 556.233 | 613.358
Populus deltoides | LI2A |  | RACiR 300 to 800 | 43.1631 | 452.973 | 1000.09 | 24.9635 | 0.010405 | 0.46204 | 557.772 | 614.831
Populus deltoides | LI2A |  | RACiR 300 to 800 | 43.1142 | 454.7443 | 999.663 | 24.9644 | 0.010407 | 0.46208 | 559.471 | 616.487
Populus deltoides | LI2A |  | RACiR 300 to 800 | 43.292 | 455.873 | 1000.08 | 24.9623 | 0.010407 | 0.46229 | 560.966 | 618.174
Populus deltoides | LI2A |  | RACiR 300 to 800 | 43.3224 | 457.4492 | 1000.07 | 24.96 | 0.010402 | 0.4622 | 562.66 | 619.906
Populus deltoides | LI2A |  | RACiR 300 to 800 | 43.3802 | 458.765 | 999.99 | 24.9615 | 0.010399 | 0.46194 | 564.188 | 621.499
Populus deltoides | LI2A |  | RACiR 300 to 800 | 43.6387 | 459.8664 | 999.769 | 24.9602 | 0.010397 | 0.46203 | 565.856 | 623.438
Populus deltoides | LI2A |  | RACiR 300 to 800 | 43.6625 | 461.2423 | 1000.12 | 24.9634 | 0.010395 | 0.46177 | 567.374 | 624.991
Populus deltoides | LI2A |  | RACiR 300 to 800 | 43.6853 | 462.9002 | 1000.12 | 24.9646 | 0.010397 | 0.46194 | 569.082 | 626.733
Populus deltoides | LI2A |  | RACiR 300 to 800 | 43.7963 | 464.1707 | 999.716 | 24.9679 | 0.0104 | 0.46198 | 570.619 | 628.391
Populus deltoides | LI2A |  | RACiR 300 to 800 | 43.7545 | 465.8848 | 1000.09 | 24.9716 | 0.010399 | 0.46171 | 572.342 | 630.087
Populus deltoides | LI2A |  | RACiR 300 to 800 | 43.7489 | 467.4138 | 1000.11 | 24.9731 | 0.0104 | 0.46172 | 573.894 | 631.641
Populus deltoides | LI2A |  | RACiR 300 to 800 | 43.9388 | 468.6309 | 1000.01 | 24.9737 | 0.010399 | 0.46161 | 575.579 | 633.528
Populus deltoides | LI2A |  | RACiR 300 to 800 | 43.9208 | 470.1172 | 1000.09 | 24.9755 | 0.010398 | 0.46138 | 577.112 | 635.058
Populus deltoides | LI2A |  | RACiR 300 to 800 | 43.9562 | 471.6924 | 999.69 | 24.9772 | 0.010399 | 0.46135 | 578.808 | 636.798
Populus deltoides | LI2A |  | RACiR 300 to 800 | 44.128 | 472.8209 | 1000.14 | 24.9781 | 0.0104 | 0.4614 | 580.328 | 638.502
Populus deltoides | LI2A |  | RACiR 300 to 800 | 44.1423 | 474.3388 | 1000.06 | 24.9856 | 0.0104 | 0.4609 | 582.03 | 640.23
Populus deltoides | LI2A |  | RACiR 300 to 800 | 44.1231 | 475.9247 | 999.722 | 24.9847 | 0.0104 | 0.46101 | 583.586 | 641.776
Populus deltoides | LI2A |  | RACiR 300 to 800 | 44.279 | 477.3147 | 1000.07 | 24.9814 | 0.010402 | 0.46139 | 585.26 | 643.62
Populus deltoides | LI2A |  | RACiR 300 to 800 | 44.2251 | 478.9884 | 1000.07 | 24.9793 | 0.010401 | 0.46154 | 586.818 | 645.134
Populus deltoides | LI2A |  | RACiR 300 to 800 | 44.1517 | 480.8425 | 1000.07 | 24.9759 | 0.010401 | 0.46176 | 588.503 | 646.751
Populus deltoides | LI2A |  | RACiR 300 to 800 | 44.2554 | 482.0774 | 1000.05 | 24.9777 | 0.010403 | 0.46176 | 589.995 | 648.356
Populus deltoides | LI2A |  | RACiR 300 to 800 | 44.1857 | 483.804 | 1000.11 | 24.9819 | 0.010402 | 0.46145 | 591.681 | 649.992
Populus deltoides | LI2A |  | RACiR 300 to 800 | 44.2462 | 485.1195 | 1000.03 | 24.9833 | 0.010403 | 0.46141 | 593.169 | 651.545
Populus deltoides | LI2A |  | RACiR 300 to 800 | 44.4554 | 486.2526 | 1000.06 | 24.9874 | 0.010406 | 0.46127 | 594.823 | 653.419
Populus deltoides | LI2A |  | RACiR 300 to 800 | 44.5309 | 487.6454 | 999.713 | 24.9853 | 0.010407 | 0.46156 | 596.346 | 655.035
Populus deltoides | LI2A |  | RACiR 300 to 800 | 44.5645 | 489.2768 | 1000.07 | 24.984 | 0.010408 | 0.46172 | 598.051 | 656.783
Populus deltoides | LI2A |  | RACiR 300 to 800 | 44.7447 | 490.3405 | 1000.06 | 24.9865 | 0.01041 | 0.4617 | 599.541 | 658.464
Populus deltoides | LI2A |  | RACiR 300 to 800 | 44.7056 | 492.174 | 1000.09 | 24.982 | 0.010409 | 0.46202 | 601.251 | 660.154
Populus deltoides | LI2A |  | RACiR 300 to 800 | 44.6729 | 493.7159 | 1000.05 | 24.9834 | 0.01041 | 0.46194 | 602.778 | 661.653
Populus deltoides | LI2A |  | RACiR 300 to 800 | 44.8504 | 494.9118 | 1000.04 | 24.9867 | 0.010411 | 0.46178 | 604.429 | 663.494
Populus deltoides | LI2A |  | RACiR 300 to 800 | 44.7971 | 496.5426 | 999.68 | 24.9859 | 0.01041 | 0.46182 | 605.97 | 664.997
Populus deltoides | LI2A |  | RACiR 300 to 800 | 44.7613 | 498.2912 | 999.994 | 24.9867 | 0.010412 | 0.46188 | 607.668 | 666.663
Populus deltoides | LI2A |  | RACiR 300 to 800 | 44.9214 | 499.4216 | 1000.14 | 24.9889 | 0.010411 | 0.46175 | 609.204 | 668.373
Populus deltoides | LI2A |  | RACiR 300 to 800 | 44.9553 | 501.0714 | 1000.03 | 24.9862 | 0.010411 | 0.46197 | 610.913 | 670.131
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.0561 | 502.3887 | 999.876 | 24.9858 | 0.010413 | 0.4622 | 612.428 | 671.754
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.2776 | 503.6057 | 1000.04 | 24.9846 | 0.010413 | 0.4623 | 614.135 | 673.692
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.2676 | 505.1315 | 1000.06 | 24.9846 | 0.010414 | 0.46232 | 615.668 | 675.228
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.1759 | 506.9699 | 1000.07 | 24.9879 | 0.010414 | 0.4621 | 617.402 | 676.883
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.2502 | 508.3121 | 1000.08 | 24.9867 | 0.010412 | 0.46207 | 618.941 | 678.5
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.1919 | 510.0624 | 1000.09 | 24.9875 | 0.010411 | 0.46197 | 620.628 | 680.151
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.1999 | 511.5668 | 1000.06 | 24.9869 | 0.010411 | 0.46203 | 622.168 | 681.707
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.4399 | 512.7195 | 999.706 | 24.9875 | 0.010414 | 0.46221 | 623.834 | 683.627
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.4199 | 514.2528 | 1000.1 | 24.9931 | 0.010423 | 0.46225 | 625.359 | 685.15
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.4069 | 516.1034 | 1000.07 | 24.9982 | 0.010443 | 0.46283 | 627.107 | 686.903
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.6051 | 517.221 | 1000.09 | 25.0014 | 0.010459 | 0.46324 | 628.605 | 688.618
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.5765 | 519.041 | 1000.06 | 25.0033 | 0.010474 | 0.46359 | 630.336 | 690.345
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.5687 | 520.6386 | 1000.04 | 25.001 | 0.01048 | 0.46385 | 631.899 | 691.909
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.7476 | 521.8278 | 1000.07 | 24.9957 | 0.010465 | 0.46341 | 633.595 | 693.789
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.6615 | 523.4945 | 1000.05 | 24.9904 | 0.010451 | 0.46308 | 635.173 | 695.286
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.5431 | 525.3332 | 999.722 | 24.987 | 0.010435 | 0.46245 | 636.928 | 696.921
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.6519 | 526.5141 | 1000.08 | 24.9862 | 0.010424 | 0.46207 | 638.453 | 698.562
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.5651 | 528.3357 | 1000.09 | 24.9876 | 0.010421 | 0.46187 | 640.171 | 700.208
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.5761 | 529.8351 | 999.776 | 24.9867 | 0.010422 | 0.46208 | 641.68 | 701.734
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.7671 | 531.1415 | 1000.13 | 24.985 | 0.010425 | 0.46243 | 643.351 | 703.607
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.7936 | 532.5276 | 1000.09 | 24.9877 | 0.010425 | 0.46232 | 644.855 | 705.153
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.8571 | 534.0794 | 1000.1 | 24.986 | 0.010426 | 0.46245 | 646.548 | 706.913
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.9699 | 535.3798 | 1000.02 | 24.9861 | 0.010426 | 0.46256 | 648.099 | 708.592
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.9554 | 536.9691 | 1000.05 | 24.9898 | 0.010423 | 0.46208 | 649.808 | 710.303
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.937 | 538.5635 | 1000.11 | 24.9877 | 0.010425 | 0.46234 | 651.336 | 711.817
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.9991 | 540.0657 | 1000.07 | 24.9904 | 0.010424 | 0.46213 | 653.06 | 713.624
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.9707 | 541.662 | 1000.1 | 24.9901 | 0.010426 | 0.46223 | 654.607 | 715.15
Populus deltoides | LI2A |  | RACiR 300 to 800 | 45.9621 | 543.3786 | 1000.03 | 24.9887 | 0.010428 | 0.4625 | 656.28 | 716.823
Populus deltoides | LI2A |  | RACiR 300 to 800 | 46.1173 | 544.5428 | 1000.02 | 24.9898 | 0.010428 | 0.46242 | 657.83 | 718.541
Populus deltoides | LI2A |  | RACiR 300 to 800 | 46.2048 | 546.0008 | 1000.1 | 24.9923 | 0.010429 | 0.46235 | 659.533 | 720.343
Populus deltoides | LI2A |  | RACiR 300 to 800 | 46.1976 | 547.5549 | 1000.07 | 24.9915 | 0.010428 | 0.46236 | 661.102 | 721.917
Populus deltoides | LI2A |  | RACiR 300 to 800 | 46.3881 | 548.7999 | 999.989 | 24.9939 | 0.010428 | 0.46221 | 662.83 | 723.852
Populus deltoides | LI2A |  | RACiR 0 to 500 | 8.7953 | 140.514 | 999.937 | 24.9762 | 0.010013 | 0.44749 | 163.571 | 183.852
Populus deltoides | LI2A |  | RACiR 0 to 500 | 8.9134 | 141.5968 | 999.987 | 24.9765 | 0.010017 | 0.44778 | 164.931 | 185.34
Populus deltoides | LI2A |  | RACiR 0 to 500 | 9.1975 | 142.4749 | 1000.02 | 24.9749 | 0.01002 | 0.44816 | 166.451 | 187.157
Populus deltoides | LI2A |  | RACiR 0 to 500 | 9.3714 | 143.4427 | 999.952 | 24.9736 | 0.010022 | 0.44841 | 167.82 | 188.713
Populus deltoides | LI2A |  | RACiR 0 to 500 | 9.5131 | 144.6228 | 999.989 | 24.9787 | 0.010024 | 0.44818 | 169.36 | 190.407
Populus deltoides | LI2A |  | RACiR 0 to 500 | 9.7718 | 145.4038 | 999.984 | 24.9801 | 0.010025 | 0.44821 | 170.741 | 192.057
Populus deltoides | LI2A |  | RACiR 0 to 500 | 9.89 | 146.6597 | 999.979 | 24.9833 | 0.010026 | 0.4481 | 172.299 | 193.748
Populus deltoides | LI2A |  | RACiR 0 to 500 | 9.9688 | 147.8551 | 999.977 | 24.9886 | 0.010028 | 0.44784 | 173.715 | 195.253
Populus deltoides | LI2A |  | RACiR 0 to 500 | 10.1844 | 148.9003 | 1000 | 24.9899 | 0.010027 | 0.44771 | 175.278 | 197.044
Populus deltoides | LI2A |  | RACiR 0 to 500 | 10.2596 | 150.1232 | 999.959 | 24.9864 | 0.010029 | 0.44808 | 176.677 | 198.532
Populus deltoides | LI2A |  | RACiR 0 to 500 | 10.331 | 151.5638 | 999.926 | 24.9794 | 0.010044 | 0.44936 | 178.24 | 200.18
Populus deltoides | LI2A |  | RACiR 0 to 500 | 10.5836 | 152.4063 | 1000.01 | 24.9764 | 0.010056 | 0.45009 | 179.63 | 201.837
Populus deltoides | LI2A |  | RACiR 0 to 500 | 10.7596 | 153.5458 | 1000.01 | 24.9789 | 0.01007 | 0.45049 | 181.171 | 203.571
Populus deltoides | LI2A |  | RACiR 0 to 500 | 10.9015 | 154.6442 | 1000.01 | 24.9803 | 0.010082 | 0.4508 | 182.598 | 205.151
Populus deltoides | LI2A |  | RACiR 0 to 500 | 11.2331 | 155.4203 | 999.961 | 24.9819 | 0.010079 | 0.45036 | 184.163 | 207.06
Populus deltoides | LI2A |  | RACiR 0 to 500 | 11.3185 | 156.6206 | 999.974 | 24.9798 | 0.010073 | 0.45005 | 185.6 | 208.596
Populus deltoides | LI2A |  | RACiR 0 to 500 | 11.3706 | 158.0861 | 999.982 | 24.9743 | 0.01007 | 0.45019 | 187.205 | 210.264
Populus deltoides | LI2A |  | RACiR 0 to 500 | 11.5735 | 159.0027 | 999.958 | 24.973 | 0.010065 | 0.44996 | 188.612 | 211.884
Populus deltoides | LI2A |  | RACiR 0 to 500 | 11.621 | 160.412 | 1000.21 | 24.9764 | 0.010059 | 0.44937 | 190.197 | 213.531
Populus deltoides | LI2A |  | RACiR 0 to 500 | 11.7358 | 161.4753 | 1000.15 | 24.9812 | 0.010051 | 0.44852 | 191.597 | 215.054
Populus deltoides | LI2A |  | RACiR 0 to 500 | 12.0612 | 162.2093 | 1000 | 24.9844 | 0.010041 | 0.44781 | 193.127 | 216.918
Populus deltoides | LI2A |  | RACiR 0 to 500 | 12.1889 | 163.3042 | 1000.04 | 24.9851 | 0.010034 | 0.44748 | 194.555 | 218.487
Populus deltoides | LI2A |  | RACiR 0 to 500 | 12.3365 | 164.5239 | 999.945 | 24.9829 | 0.010028 | 0.44747 | 196.135 | 220.224
Populus deltoides | LI2A |  | RACiR 0 to 500 | 12.6266 | 165.272 | 999.963 | 24.9812 | 0.010028 | 0.44774 | 197.536 | 221.929
Populus deltoides | LI2A |  | RACiR 0 to 500 | 12.7301 | 166.6405 | 1000.04 | 24.9811 | 0.010032 | 0.44802 | 199.15 | 223.66
Populus deltoides | LI2A |  | RACiR 0 to 500 | 12.8102 | 167.8991 | 999.914 | 24.9823 | 0.010034 | 0.44815 | 200.609 | 225.211
Populus deltoides | LI2A |  | RACiR 0 to 500 | 13.0453 | 168.9411 | 1000.01 | 24.9824 | 0.010035 | 0.44822 | 202.2 | 227.049
Populus deltoides | LI2A |  | RACiR 0 to 500 | 13.1077 | 170.2265 | 1000.01 | 24.98 | 0.010036 | 0.44854 | 203.632 | 228.558
Populus deltoides | LI2A |  | RACiR 0 to 500 | 13.1555 | 171.7016 | 999.944 | 24.9787 | 0.010039 | 0.4489 | 205.222 | 230.209
Populus deltoides | LI2A |  | RACiR 0 to 500 | 13.3567 | 172.6279 | 999.948 | 24.9796 | 0.01004 | 0.44898 | 206.617 | 231.816
Populus deltoides | LI2A |  | RACiR 0 to 500 | 13.479 | 173.8794 | 999.953 | 24.981 | 0.010042 | 0.44902 | 208.17 | 233.507
Populus deltoides | LI2A |  | RACiR 0 to 500 | 13.5802 | 175.0506 | 1000 | 24.9848 | 0.010045 | 0.44896 | 209.602 | 235.051
Populus deltoides | LI2A |  | RACiR 0 to 500 | 13.8776 | 175.896 | 999.944 | 24.9882 | 0.010045 | 0.44871 | 211.156 | 236.917
Populus deltoides | LI2A |  | RACiR 0 to 500 | 14.0437 | 176.9186 | 999.984 | 24.9891 | 0.010047 | 0.44882 | 212.568 | 238.507
Populus deltoides | LI2A |  | RACiR 0 to 500 | 14.1926 | 178.1515 | 999.992 | 24.9887 | 0.010048 | 0.44888 | 214.16 | 240.259
Populus deltoides | LI2A |  | RACiR 0 to 500 | 14.4668 | 178.9604 | 999.956 | 24.9838 | 0.010048 | 0.44931 | 215.57 | 241.953
Populus deltoides | LI2A |  | RACiR 0 to 500 | 14.6135 | 180.2132 | 999.927 | 24.9792 | 0.010048 | 0.44968 | 217.15 | 243.695
Populus deltoides | LI2A |  | RACiR 0 to 500 | 14.6877 | 181.4705 | 999.962 | 24.9767 | 0.010047 | 0.4498 | 218.592 | 245.222
Populus deltoides | LI2A |  | RACiR 0 to 500 | 14.8925 | 182.5513 | 1000.23 | 24.9772 | 0.010047 | 0.4498 | 220.158 | 247.005
Populus deltoides | LI2A |  | RACiR 0 to 500 | 14.9551 | 183.7917 | 1000 | 24.9802 | 0.010046 | 0.44961 | 221.583 | 248.507
Populus deltoides | LI2A |  | RACiR 0 to 500 | 15.0406 | 185.1277 | 999.986 | 24.9846 | 0.010047 | 0.44932 | 223.167 | 250.186
Populus deltoides | LI2A |  | RACiR 0 to 500 | 15.4309 | 186.417 | 999.96 | 24.9876 | 0.01005 | 0.44929 | 225.367 | 252.793
Populus deltoides | LI2A |  | RACiR 0 to 500 | 15.4309 | 186.4174 | 999.951 | 24.9876 | 0.01005 | 0.4493 | 225.367 | 252.793
Populus deltoides | LI2A |  | RACiR 0 to 500 | 15.5508 | 188.3054 | 999.935 | 24.9928 | 0.01005 | 0.44896 | 227.598 | 255.165
Populus deltoides | LI2A |  | RACiR 0 to 500 | 15.8169 | 189.2747 | 999.977 | 24.9934 | 0.01006 | 0.4494 | 229.154 | 257.001
Populus deltoides | LI2A |  | RACiR 0 to 500 | 15.9466 | 190.4354 | 1000.02 | 24.9891 | 0.010068 | 0.45003 | 230.579 | 258.574
Populus deltoides | LI2A |  | RACiR 0 to 500 | 16.1265 | 191.6268 | 999.954 | 24.9857 | 0.010082 | 0.45082 | 232.137 | 260.32
Populus deltoides | LI2A |  | RACiR 0 to 500 | 16.4439 | 192.3452 | 1000.01 | 24.984 | 0.010088 | 0.45113 | 233.558 | 262.078
Populus deltoides | LI2A |  | RACiR 0 to 500 | 16.6453 | 193.4562 | 1000.04 | 24.9862 | 0.010091 | 0.45095 | 235.164 | 263.899
Populus deltoides | LI2A |  | RACiR 0 to 500 | 16.726 | 194.6764 | 999.987 | 24.9865 | 0.010088 | 0.45064 | 236.62 | 265.448
Populus deltoides | LI2A |  | RACiR 0 to 500 | 16.9465 | 195.7565 | 1000.25 | 24.9862 | 0.010088 | 0.45061 | 238.222 | 267.28
Populus deltoides | LI2A |  | RACiR 0 to 500 | 17.0288 | 196.9777 | 999.932 | 24.9837 | 0.010082 | 0.45044 | 239.669 | 268.825
Populus deltoides | LI2A |  | RACiR 0 to 500 | 17.1423 | 198.2526 | 999.926 | 24.9791 | 0.010069 | 0.45007 | 241.257 | 270.534
Populus deltoides | LI2A |  | RACiR 0 to 500 | 17.4209 | 198.9766 | 1000.02 | 24.98 | 0.010056 | 0.44944 | 242.679 | 272.242
Populus deltoides | LI2A |  | RACiR 0 to 500 | 17.5571 | 200.1917 | 999.99 | 24.977 | 0.01004 | 0.44892 | 244.272 | 273.982
Populus deltoides | LI2A |  | RACiR 0 to 500 | 17.6815 | 201.2948 | 1000.01 | 24.9767 | 0.01003 | 0.44857 | 245.71 | 275.551
Populus deltoides | LI2A |  | RACiR 0 to 500 | 17.9537 | 202.2456 | 1000.04 | 24.979 | 0.010027 | 0.44847 | 247.305 | 277.433
Populus deltoides | LI2A |  | RACiR 0 to 500 | 18.029 | 203.5365 | 999.957 | 24.977 | 0.010032 | 0.44903 | 248.74 | 278.958
Populus deltoides | LI2A |  | RACiR 0 to 500 | 18.0468 | 205.1067 | 999.983 | 24.9775 | 0.010036 | 0.44943 | 250.347 | 280.595
Populus deltoides | LI2A |  | RACiR 0 to 500 | 18.2375 | 206.0629 | 999.978 | 24.9784 | 0.010039 | 0.44961 | 251.737 | 282.189
Populus deltoides | LI2A |  | RACiR 0 to 500 | 18.3556 | 207.3357 | 1000.01 | 24.9783 | 0.010042 | 0.44982 | 253.284 | 283.866
Populus deltoides | LI2A |  | RACiR 0 to 500 | 18.4251 | 208.5894 | 999.932 | 24.9813 | 0.010045 | 0.44988 | 254.717 | 285.379
Populus deltoides | LI2A |  | RACiR 0 to 500 | 18.7404 | 209.3804 | 999.934 | 24.9862 | 0.010046 | 0.44965 | 256.26 | 287.256
Populus deltoides | LI2A |  | RACiR 0 to 500 | 18.8885 | 210.4399 | 999.962 | 24.988 | 0.010049 | 0.44972 | 257.67 | 288.824
Populus deltoides | LI2A |  | RACiR 0 to 500 | 19.0613 | 211.6042 | 999.968 | 24.9905 | 0.010051 | 0.4498 | 259.242 | 290.582
Populus deltoides | LI2A |  | RACiR 0 to 500 | 19.3395 | 212.3605 | 999.97 | 24.9905 | 0.01005 | 0.44977 | 260.644 | 292.274
Populus deltoides | LI2A |  | RACiR 0 to 500 | 19.4815 | 213.619 | 1000.06 | 24.9875 | 0.01005 | 0.45001 | 262.224 | 294.008
Populus deltoides | LI2A |  | RACiR 0 to 500 | 19.6018 | 214.7215 | 1000.13 | 24.9909 | 0.010052 | 0.44988 | 263.637 | 295.551
Populus deltoides | LI2A |  | RACiR 0 to 500 | 19.8525 | 215.6829 | 1000 | 24.9934 | 0.010053 | 0.44977 | 265.196 | 297.375
Populus deltoides | LI2A |  | RACiR 0 to 500 | 19.9095 | 216.9909 | 1000.01 | 24.9923 | 0.010055 | 0.45004 | 266.633 | 298.885
Populus deltoides | LI2A |  | RACiR 0 to 500 | 20.0264 | 218.2772 | 1000.02 | 24.9923 | 0.010058 | 0.45026 | 268.188 | 300.568
Populus deltoides | LI2A |  | RACiR 0 to 500 | 20.3156 | 219.0321 | 999.987 | 24.989 | 0.010058 | 0.45058 | 269.573 | 302.256
Populus deltoides | LI2A |  | RACiR 0 to 500 | 20.4885 | 220.2623 | 1000.02 | 24.9829 | 0.010061 | 0.45121 | 271.149 | 304.014
Populus deltoides | LI2A |  | RACiR 0 to 500 | 20.6701 | 221.2513 | 999.981 | 24.9838 | 0.010061 | 0.45121 | 272.567 | 305.626
Populus deltoides | LI2A |  | RACiR 0 to 500 | 20.9978 | 222.0609 | 999.975 | 24.9876 | 0.010064 | 0.45112 | 274.141 | 307.539
Populus deltoides | LI2A |  | RACiR 0 to 500 | 21.0607 | 223.3544 | 999.955 | 24.9876 | 0.010062 | 0.45111 | 275.606 | 309.084
Populus deltoides | LI2A |  | RACiR 0 to 500 | 21.1441 | 224.7544 | 999.973 | 24.9895 | 0.010063 | 0.45107 | 277.23 | 310.804
Populus deltoides | LI2A |  | RACiR 0 to 500 | 21.3444 | 225.6961 | 1000.01 | 24.9914 | 0.010065 | 0.45098 | 278.653 | 312.439
Populus deltoides | LI2A |  | RACiR 0 to 500 | 21.4278 | 227.1258 | 999.952 | 24.9902 | 0.010074 | 0.45149 | 280.246 | 314.134
Populus deltoides | LI2A |  | RACiR 0 to 500 | 21.5449 | 228.2909 | 999.942 | 24.9906 | 0.010083 | 0.4518 | 281.668 | 315.685
Populus deltoides | LI2A |  | RACiR 0 to 500 | 21.8532 | 229.1744 | 999.904 | 24.9918 | 0.010094 | 0.45215 | 283.226 | 317.57
Populus deltoides | LI2A |  | RACiR 0 to 500 | 21.9711 | 230.372 | 1000.29 | 24.9887 | 0.010103 | 0.45265 | 284.658 | 319.134
Populus deltoides | LI2A |  | RACiR 0 to 500 | 22.1179 | 231.6464 | 1000.02 | 24.9842 | 0.010102 | 0.45282 | 286.269 | 320.904
Populus deltoides | LI2A |  | RACiR 0 to 500 | 22.4289 | 232.3107 | 1000.01 | 24.9847 | 0.0101 | 0.45254 | 287.676 | 322.628
Populus deltoides | LI2A |  | RACiR 0 to 500 | 22.5337 | 233.5888 | 999.975 | 24.9887 | 0.010094 | 0.45183 | 289.302 | 324.376
Populus deltoides | LI2A |  | RACiR 0 to 500 | 22.699 | 234.5636 | 999.958 | 24.9912 | 0.010091 | 0.45141 | 290.719 | 325.965
Populus deltoides | LI2A |  | RACiR 0 to 500 | 22.9086 | 235.6151 | 999.968 | 24.9941 | 0.010087 | 0.45099 | 292.315 | 327.788
Populus deltoides | LI2A |  | RACiR 0 to 500 | 22.9669 | 236.8991 | 1000.01 | 24.9924 | 0.010086 | 0.45099 | 293.758 | 329.3
Populus deltoides | LI2A |  | RACiR 0 to 500 | 23.0033 | 238.422 | 1000.08 | 24.9866 | 0.010086 | 0.45145 | 295.339 | 330.928
Populus deltoides | LI2A |  | RACiR 0 to 500 | 23.1663 | 239.4727 | 999.967 | 24.9844 | 0.010087 | 0.45163 | 296.756 | 332.518
Populus deltoides | LI2A |  | RACiR 0 to 500 | 23.2906 | 240.753 | 1000.04 | 24.9811 | 0.010086 | 0.45181 | 298.319 | 334.219
Populus deltoides | LI2A |  | RACiR 0 to 500 | 23.3946 | 241.8787 | 1000.01 | 24.9832 | 0.010086 | 0.4517 | 299.718 | 335.734
Populus deltoides | LI2A |  | RACiR 0 to 500 | 23.7243 | 242.6518 | 999.944 | 24.9853 | 0.010087 | 0.45166 | 301.253 | 337.61
Populus deltoides | LI2A |  | RACiR 0 to 500 | 23.8783 | 243.74 | 999.948 | 24.9814 | 0.010087 | 0.452 | 302.665 | 339.192
Populus deltoides | LI2A |  | RACiR 0 to 500 | 24.0506 | 244.9435 | 1000.04 | 24.9789 | 0.010089 | 0.45238 | 304.232 | 340.943
Populus deltoides | LI2A |  | RACiR 0 to 500 | 24.3026 | 245.8283 | 999.971 | 24.9777 | 0.010092 | 0.45261 | 305.671 | 342.643
Populus deltoides | LI2A |  | RACiR 0 to 500 | 24.4147 | 247.147 | 999.964 | 24.9798 | 0.010092 | 0.4525 | 307.285 | 344.386
Populus deltoides | LI2A |  | RACiR 0 to 500 | 24.5219 | 248.2668 | 1000.01 | 24.9856 | 0.010093 | 0.45212 | 308.722 | 345.938
Populus deltoides | LI2A |  | RACiR 0 to 500 | 24.7067 | 249.3898 | 1000.28 | 24.9896 | 0.010091 | 0.45173 | 310.335 | 347.755
Populus deltoides | LI2A |  | RACiR 0 to 500 | 24.7567 | 250.6674 | 999.955 | 24.9923 | 0.010092 | 0.45156 | 311.777 | 349.258
Populus deltoides | LI2A |  | RACiR 0 to 500 | 24.8582 | 252.0328 | 999.942 | 24.9888 | 0.010093 | 0.45194 | 313.349 | 350.946
Populus deltoides | LI2A |  | RACiR 0 to 500 | 25.1107 | 252.8764 | 999.958 | 24.9861 | 0.010091 | 0.45204 | 314.762 | 352.624
Populus deltoides | LI2A |  | RACiR 0 to 500 | 25.2845 | 254.048 | 999.938 | 24.9861 | 0.010091 | 0.45204 | 316.348 | 354.394
Populus deltoides | LI2A |  | RACiR 0 to 500 | 25.4849 | 255.0234 | 999.968 | 24.9844 | 0.010092 | 0.45223 | 317.769 | 356.024
Populus deltoides | LI2A |  | RACiR 0 to 500 | 25.7844 | 255.9367 | 999.96 | 24.9855 | 0.010092 | 0.45219 | 319.378 | 357.945
Populus deltoides | LI2A |  | RACiR 0 to 500 | 25.8571 | 257.1946 | 999.991 | 24.9881 | 0.010092 | 0.45205 | 320.846 | 359.5
Populus deltoides | LI2A |  | RACiR 0 to 500 | 25.8812 | 258.7893 | 1000.02 | 24.986 | 0.010093 | 0.45229 | 322.498 | 361.192
Populus deltoides | LI2A |  | RACiR 0 to 500 | 26.0676 | 259.7881 | 1000.28 | 24.9864 | 0.010094 | 0.4523 | 323.934 | 362.826
Populus deltoides | LI2A |  | RACiR 0 to 500 | 26.1325 | 261.2197 | 999.95 | 24.987 | 0.010093 | 0.45222 | 325.554 | 364.528
Populus deltoides | LI2A |  | RACiR 0 to 500 | 26.1992 | 262.5149 | 999.966 | 24.9842 | 0.010092 | 0.45239 | 327.004 | 366.053
Populus deltoides | LI2A |  | RACiR 0 to 500 | 26.4802 | 263.4356 | 1000 | 24.9835 | 0.01009 | 0.45231 | 328.584 | 367.925
Populus deltoides | LI2A |  | RACiR 0 to 500 | 26.577 | 264.6298 | 999.961 | 24.9833 | 0.010088 | 0.45226 | 330.027 | 369.478
Populus deltoides | LI2B | 2 | Standard | 278.4835 | 27.0805 | 24.9939 | 1000.14 | 0.0072148 | 0.32046 | 370.292 | 400.043
Populus deltoides | LI2B | 2 | Standard | 244.4165 | 24.0332 | 24.9963 | 1000.14 | 0.0074526 | 0.33028 | 323.59 | 350.035
Populus deltoides | LI2B | 2 | Standard | 211.5934 | 20.4977 | 25.0023 | 1000.01 | 0.0076502 | 0.33967 | 277.446 | 300.063
Populus deltoides | LI2B | 2 | Standard | 178.5569 | 16.8397 | 25.0024 | 1000.05 | 0.0077592 | 0.34844 | 231.451 | 250.086
Populus deltoides | LI2B | 2 | Standard | 145.4928 | 12.8458 | 24.9927 | 1000.06 | 0.0079393 | 0.35191 | 185.732 | 200.052
Populus deltoides | LI2B | 2 | Standard | 113.3922 | 8.6249 | 25.0263 | 999.973 | 0.008074 | 0.35906 | 140.265 | 150.022
Populus deltoides | LI2B | 2 | Standard | 81.3549 | 4.2827 | 25.0262 | 999.971 | 0.0081478 | 0.36572 | 95.03 | 100.087
Populus deltoides | LI2B | 2 | Standard | 49.6228 | -0.30417 | 24.9962 | 1000.02 | 0.0083048 | 0.369 | 49.9186 | 50.029
Populus deltoides | LI2B | 2 | Standard | 291.1748 | 26.8522 | 25.0308 | 1000.03 | 0.0085129 | 0.37622 | 370.029 | 400.032
Populus deltoides | LI2B | 2 | Standard | 372.0541 | 31.4578 | 25.0666 | 1000.05 | 0.0085094 | 0.37811 | 464.666 | 500.079
Populus deltoides | LI2B | 2 | Standard | 455.451 | 35.2061 | 25.0125 | 1000.06 | 0.0085367 | 0.37803 | 560.049 | 600.034
Populus deltoides | LI2B | 2 | Standard | 543.0168 | 37.5795 | 25.0324 | 999.963 | 0.008469 | 0.37467 | 656.878 | 700.022
Populus deltoides | LI2B | 2 | Standard | 634.0068 | 39.0593 | 25.0365 | 1000.13 | 0.0084069 | 0.37198 | 754.703 | 800.108
Populus deltoides | LI2B | 2 | Standard | 725.6945 | 40.082 | 25.0245 | 1000.12 | 0.0082844 | 0.36641 | 852.931 | 900.082
Populus deltoides | LI2B | 2 | Standard | 818.4343 | 40.7167 | 25.0311 | 999.942 | 0.008173 | 0.36022 | 951.548 | 1000.04
Populus deltoides | LI2B | 2 | Standard | 912.1088 | 41.0916 | 25.0287 | 1000.13 | 0.0080207 | 0.35367 | 1050.55 | 1100.07
Populus deltoides | LI2B | 2 | Standard | 1007.1887 | 41.2079 | 25.0227 | 999.799 | 0.0078992 | 0.34884 | 1149.74 | 1200.03
Populus deltoides | LI2B | 2 | Standard | 1101.5835 | 41.5677 | 25.0099 | 1000.08 | 0.0077871 | 0.34459 | 1248.77 | 1300.06
Populus deltoides | LI2B | 2 | Standard | 1196.7264 | 41.623 | 25.0241 | 1000.12 | 0.0076834 | 0.33971 | 1348.03 | 1400.01
Populus deltoides | LI2B | 2 | Standard | 1290.989 | 41.9681 | 25.0193 | 1000.12 | 0.0075912 | 0.33541 | 1447.1 | 1500.05
Populus deltoides | LI2B | 2 | Standard | 1578.9723 | 42.1742 | 25.0269 | 1000.12 | 0.0074714 | 0.32916 | 1744.82 | 1800.03
Populus deltoides | LI2B | 2 | Standard | 1768.768 | 42.8092 | 25.0237 | 1000.07 | 0.0073554 | 0.32403 | 1943.01 | 2000.11
Populus deltoides | LI2B | 1 | Standard | 282.3292 | 25.8321 | 25.0103 | 1000.13 | 0.0071284 | 0.31574 | 371.525 | 400.005
Populus deltoides | LI2B | 1 | Standard | 246.4736 | 22.9805 | 25.0143 | 1000.15 | 0.0072147 | 0.32012 | 324.698 | 350.023
Populus deltoides | LI2B | 1 | Standard | 212.0531 | 19.6881 | 24.9878 | 1000.12 | 0.0072857 | 0.32413 | 278.305 | 300.02
Populus deltoides | LI2B | 1 | Standard | 178.6762 | 16.1406 | 24.9961 | 1000.1 | 0.0074144 | 0.33024 | 232.164 | 250.026
Populus deltoides | LI2B | 1 | Standard | 145.6383 | 12.3423 | 25.005 | 999.683 | 0.0075161 | 0.33434 | 186.285 | 200.027
Populus deltoides | LI2B | 1 | Standard | 113.6161 | 8.2428 | 25.0092 | 999.958 | 0.0076511 | 0.34043 | 140.687 | 150.006
Populus deltoides | LI2B | 1 | Standard | 81.3824 | 4.1041 | 25.0054 | 1000.05 | 0.0077721 | 0.34723 | 95.1779 | 100.022
Populus deltoides | LI2B | 1 | Standard | 49.7155 | -0.3164 | 25.008 | 999.97 | 0.0079081 | 0.35175 | 49.9362 | 50.0147
Populus deltoides | LI2B | 1 | Standard | 294.3987 | 25.0824 | 25.0038 | 1000.09 | 0.0080665 | 0.358 | 371.969 | 400.051
Populus deltoides | LI2B | 1 | Standard | 378.9285 | 28.9343 | 25.0214 | 1000.08 | 0.008227 | 0.36686 | 467.286 | 500.063
Populus deltoides | LI2B | 1 | Standard | 464.6365 | 32.3871 | 24.9879 | 1000.21 | 0.0084014 | 0.37348 | 562.912 | 600.027
Populus deltoides | LI2B | 1 | Standard | 554.1508 | 34.3941 | 25.023 | 1000.21 | 0.0083586 | 0.3724 | 660.135 | 700.046
Populus deltoides | LI2B | 1 | Standard | 647.9509 | 35.4068 | 25.0198 | 999.945 | 0.0083999 | 0.3746 | 758.235 | 800.009
Populus deltoides | LI2B | 1 | Standard | 742.8715 | 36.0028 | 25.0114 | 999.762 | 0.008443 | 0.37506 | 856.869 | 900.105
Populus deltoides | LI2B | 1 | Standard | 838.3779 | 36.146 | 25.0362 | 999.83 | 0.0083358 | 0.37105 | 955.948 | 1000.06
Populus deltoides | LI2B | 1 | Standard | 934.2568 | 36.2423 | 25.0153 | 1000.2 | 0.0082923 | 0.36853 | 1054.98 | 1099.97
Populus deltoides | LI2B | 1 | Standard | 1031.2329 | 36.171 | 25.0356 | 1000.14 | 0.0082134 | 0.36643 | 1154.44 | 1200.09
Populus deltoides | LI2B | 1 | Standard | 1127.5563 | 35.9673 | 25.0093 | 999.84 | 0.0081289 | 0.3618 | 1253.72 | 1299.88
Populus deltoides | LI2B | 1 | Standard | 1220.8629 | 36.5451 | 25.0134 | 999.83 | 0.0079961 | 0.35535 | 1352.66 | 1400.02
Populus deltoides | LI2B | 1 | Standard | 1317.2861 | 36.5619 | 25.004 | 1000.11 | 0.0079658 | 0.3532 | 1452.03 | 1500.16
Populus deltoides | LI2B | 1 | Standard | 1605.2956 | 36.6618 | 25.0383 | 999.774 | 0.0077883 | 0.34422 | 1749.76 | 1800.05
Populus deltoides | LI2B | 1 | Standard | 1797.1164 | 36.9352 | 25.0199 | 1000.17 | 0.0076624 | 0.33975 | 1948.06 | 1999.92
Populus deltoides | LI2B | 3 | Standard | 288.3611 | 24.1739 | 25.0378 | 1000.09 | 0.0070778 | 0.31239 | 373.24 | 400.055
Populus deltoides | LI2B | 3 | Standard | 246.9861 | 22.2722 | 25.0029 | 1000.08 | 0.0069954 | 0.30915 | 325.506 | 350.055
Populus deltoides | LI2B | 3 | Standard | 210.7334 | 19.2004 | 25.0126 | 999.805 | 0.0069373 | 0.3067 | 278.873 | 300.007
Populus deltoides | LI2B | 3 | Standard | 176.7769 | 15.747 | 25.0346 | 1000.1 | 0.0069499 | 0.30715 | 232.677 | 250.04
Populus deltoides | LI2B | 3 | Standard | 144.1007 | 11.9964 | 25.015 | 1000.08 | 0.0069861 | 0.30849 | 186.734 | 200.035
Populus deltoides | LI2B | 3 | Standard | 112.023 | 8.1038 | 25.0267 | 1000.06 | 0.007025 | 0.3109 | 140.947 | 150.041
Populus deltoides | LI2B | 3 | Standard | 80.7605 | 3.9747 | 25.0111 | 999.774 | 0.0070849 | 0.31405 | 95.4041 | 100.055
Populus deltoides | LI2B | 3 | Standard | 49.7565 | -0.29615 | 25.0056 | 1000.09 | 0.0071977 | 0.31867 | 49.9532 | 50.0166
Populus deltoides | LI2B | 3 | Standard | 282.0335 | 26.5192 | 24.9963 | 1000.11 | 0.0074112 | 0.32611 | 370.77 | 400.038
Populus deltoides | LI2B | 3 | Standard | 359.8972 | 31.4154 | 25.0398 | 1000.18 | 0.0074138 | 0.32755 | 465.144 | 500.006
Populus deltoides | LI2B | 3 | Standard | 442.2699 | 35.3755 | 25.0141 | 999.866 | 0.007467 | 0.33094 | 560.478 | 600.04
Populus deltoides | LI2B | 3 | Standard | 529.1757 | 37.9564 | 25.0249 | 999.941 | 0.0075022 | 0.33145 | 657.118 | 700.005
Populus deltoides | LI2B | 3 | Standard | 622.1516 | 39.3899 | 25.0201 | 999.821 | 0.0075701 | 0.3359 | 754.934 | 800.038
Populus deltoides | LI2B | 3 | Standard | 714.3802 | 40.4687 | 25.0168 | 999.906 | 0.0075415 | 0.33428 | 853.126 | 900.028
Populus deltoides | LI2B | 3 | Standard | 809.4161 | 40.9845 | 25.0088 | 999.857 | 0.0075961 | 0.33462 | 951.888 | 1000.1
Populus deltoides | LI2B | 3 | Standard | 903.3389 | 41.4947 | 25.0241 | 999.89 | 0.0075039 | 0.33134 | 1050.7 | 1100.08
Populus deltoides | LI2B | 3 | Standard | 998.9817 | 41.6177 | 25.015 | 999.871 | 0.0074303 | 0.32866 | 1149.9 | 1200.06
Populus deltoides | LI2B | 3 | Standard | 1092.7958 | 42.1467 | 24.9998 | 1000.03 | 0.0073873 | 0.32581 | 1248.7 | 1300.07
Populus deltoides | LI2B | 3 | Standard | 1187.7042 | 42.3883 | 25.024 | 1000.02 | 0.0073243 | 0.32279 | 1347.79 | 1400.05
Populus deltoides | LI2B | 3 | Standard | 1283.8446 | 42.4639 | 25.0012 | 1000.14 | 0.0073091 | 0.32155 | 1446.94 | 1499.98
Populus deltoides | LI2B | 3 | Standard | 1572.9973 | 42.7634 | 24.9969 | 999.831 | 0.0072494 | 0.31918 | 1744.65 | 1800.06
Populus deltoides | LI2B |  | RACiR 800 to 300 | 573.8031 | 33.5888 | 25 | 999.894 | 0.008401 | 0.3701 | 678.776 | 707.664
Populus deltoides | LI2B |  | RACiR 800 to 300 | 571.7425 | 33.5411 | 24.991 | 1000.25 | 0.0084008 | 0.37064 | 676.381 | 705.211
Populus deltoides | LI2B |  | RACiR 800 to 300 | 569.1827 | 33.6306 | 24.9923 | 999.933 | 0.0084048 | 0.37079 | 673.972 | 702.877
Populus deltoides | LI2B |  | RACiR 800 to 300 | 567.9357 | 33.4763 | 24.9954 | 999.976 | 0.0084078 | 0.37076 | 672.29 | 701.037
Populus deltoides | LI2B |  | RACiR 800 to 300 | 565.6858 | 33.4697 | 25.0033 | 999.923 | 0.0084295 | 0.37123 | 669.873 | 698.615
Populus deltoides | LI2B |  | RACiR 800 to 300 | 563.5537 | 33.4251 | 25.0066 | 1000.22 | 0.0084483 | 0.37172 | 667.463 | 696.162
Populus deltoides | LI2B |  | RACiR 800 to 300 | 561.9907 | 33.3937 | 25.0064 | 999.928 | 0.0084536 | 0.37179 | 665.769 | 694.431
Populus deltoides | LI2B |  | RACiR 800 to 300 | 559.6181 | 33.3623 | 25.0023 | 1000.24 | 0.0084422 | 0.37129 | 663.376 | 691.982
Populus deltoides | LI2B |  | RACiR 800 to 300 | 557.6907 | 33.1482 | 24.9981 | 999.861 | 0.0084259 | 0.37061 | 660.985 | 689.355
Populus deltoides | LI2B |  | RACiR 800 to 300 | 555.6299 | 33.2204 | 25.0022 | 999.946 | 0.0084195 | 0.37 | 659.237 | 687.666
Populus deltoides | LI2B |  | RACiR 800 to 300 | 552.9424 | 33.219 | 25.0054 | 999.944 | 0.0083996 | 0.36884 | 656.78 | 685.185
Populus deltoides | LI2B |  | RACiR 800 to 300 | 549.8853 | 33.4388 | 25.0006 | 999.875 | 0.0083811 | 0.36834 | 654.368 | 682.965
Populus deltoides | LI2B |  | RACiR 800 to 300 | 548.3016 | 33.3857 | 24.9978 | 999.937 | 0.0083709 | 0.36815 | 652.641 | 681.171
Populus deltoides | LI2B |  | RACiR 800 to 300 | 546.1787 | 33.368 | 24.9908 | 999.923 | 0.0083727 | 0.3688 | 650.238 | 678.736
Populus deltoides | LI2B |  | RACiR 800 to 300 | 543.9781 | 33.3031 | 24.9998 | 999.953 | 0.0083794 | 0.36876 | 647.83 | 676.255
Populus deltoides | LI2B |  | RACiR 800 to 300 | 542.3287 | 33.3059 | 25.0011 | 999.957 | 0.0083818 | 0.36888 | 646.119 | 674.542
Populus deltoides | LI2B |  | RACiR 800 to 300 | 540.1442 | 33.2791 | 25.0017 | 999.915 | 0.0083861 | 0.36916 | 643.738 | 672.123
Populus deltoides | LI2B |  | RACiR 800 to 300 | 538.3896 | 33.0925 | 24.9973 | 1000.23 | 0.0083885 | 0.36959 | 641.315 | 669.505
Populus deltoides | LI2B |  | RACiR 800 to 300 | 536.3519 | 33.1961 | 25.0005 | 999.92 | 0.0083897 | 0.36947 | 639.552 | 667.835
Populus deltoides | LI2B |  | RACiR 800 to 300 | 533.9683 | 33.1441 | 25.0089 | 999.897 | 0.0083902 | 0.36903 | 637.095 | 665.318
Populus deltoides | LI2B |  | RACiR 800 to 300 | 531.2523 | 33.2575 | 25.0132 | 1000.24 | 0.0083935 | 0.36896 | 634.652 | 662.971
Populus deltoides | LI2B |  | RACiR 800 to 300 | 529.9541 | 33.1221 | 25.0127 | 999.906 | 0.0083944 | 0.36907 | 632.924 | 661.101
Populus deltoides | LI2B |  | RACiR 800 to 300 | 527.6551 | 33.1076 | 25.0103 | 999.939 | 0.0083946 | 0.36922 | 630.491 | 658.64
Populus deltoides | LI2B |  | RACiR 800 to 300 | 525.4813 | 33.0616 | 25.008 | 1000.26 | 0.0083955 | 0.36944 | 628.081 | 656.174
Populus deltoides | LI2B |  | RACiR 800 to 300 | 523.8898 | 33.0129 | 25.0085 | 999.929 | 0.0083939 | 0.36932 | 626.349 | 654.386
Populus deltoides | LI2B |  | RACiR 800 to 300 | 521.6422 | 32.9823 | 25.0089 | 1000.23 | 0.0083971 | 0.36948 | 623.927 | 651.92
Populus deltoides | LI2B |  | RACiR 800 to 300 | 519.4941 | 32.8875 | 25.0091 | 999.997 | 0.0083953 | 0.36936 | 621.501 | 649.388
Populus deltoides | LI2B |  | RACiR 800 to 300 | 517.3022 | 33.0728 | 25.0077 | 999.859 | 0.0083979 | 0.36956 | 619.712 | 647.771
Populus deltoides | LI2B |  | RACiR 800 to 300 | 515.1329 | 33.0034 | 25.0046 | 999.898 | 0.0083959 | 0.36967 | 617.271 | 645.251
Populus deltoides | LI2B |  | RACiR 800 to 300 | 512.5438 | 33.0718 | 25.0067 | 999.913 | 0.0083963 | 0.36955 | 614.844 | 642.876
Populus deltoides | LI2B |  | RACiR 800 to 300 | 511.2107 | 32.9152 | 25.0088 | 999.932 | 0.008394 | 0.36926 | 613.128 | 640.994
Populus deltoides | LI2B |  | RACiR 800 to 300 | 508.7131 | 32.9265 | 25.0123 | 999.906 | 0.008393 | 0.36894 | 610.691 | 638.555
Populus deltoides | LI2B |  | RACiR 800 to 300 | 506.5882 | 32.8732 | 25.0058 | 999.931 | 0.0083937 | 0.36939 | 608.247 | 636.045
Populus deltoides | LI2B |  | RACiR 800 to 300 | 504.9019 | 32.8696 | 25.0067 | 999.955 | 0.0083944 | 0.3694 | 606.511 | 634.299
Populus deltoides | LI2B |  | RACiR 800 to 300 | 502.3959 | 32.9085 | 25.009 | 999.896 | 0.0083958 | 0.36931 | 604.08 | 631.891
Populus deltoides | LI2B |  | RACiR 800 to 300 | 500.009 | 32.8435 | 25.0186 | 1000.21 | 0.0083935 | 0.36859 | 601.658 | 629.395
Populus deltoides | LI2B |  | RACiR 800 to 300 | 497.8009 | 33.0173 | 25.0189 | 999.939 | 0.0083929 | 0.36852 | 599.895 | 627.793
Populus deltoides | LI2B |  | RACiR 800 to 300 | 495.582 | 32.9751 | 25.0183 | 999.923 | 0.0083913 | 0.36849 | 597.515 | 625.361
Populus deltoides | LI2B |  | RACiR 800 to 300 | 493.1666 | 33.0347 | 25.0146 | 1000.27 | 0.0083929 | 0.36877 | 595.133 | 623.021
Populus deltoides | LI2B |  | RACiR 800 to 300 | 492.0169 | 32.8515 | 25.0121 | 999.925 | 0.0083908 | 0.36878 | 593.446 | 621.146
Populus deltoides | LI2B |  | RACiR 800 to 300 | 489.7527 | 32.7916 | 25.0155 | 999.915 | 0.0083897 | 0.36851 | 591.04 | 618.664
Populus deltoides | LI2B |  | RACiR 800 to 300 | 487.6958 | 32.6828 | 25.0174 | 1000.27 | 0.0083905 | 0.36843 | 588.66 | 616.165
Populus deltoides | LI2B |  | RACiR 800 to 300 | 485.892 | 32.6989 | 25.0215 | 999.905 | 0.0083893 | 0.36811 | 586.945 | 614.457
Populus deltoides | LI2B |  | RACiR 800 to 300 | 483.2065 | 32.8241 | 25.0201 | 1000.25 | 0.0083898 | 0.36821 | 584.516 | 612.139
Populus deltoides | LI2B |  | RACiR 800 to 300 | 480.8836 | 32.7968 | 25.0192 | 999.942 | 0.0083852 | 0.36799 | 582.118 | 609.7
Populus deltoides | LI2B |  | RACiR 800 to 300 | 479.0447 | 32.9079 | 25.0113 | 999.915 | 0.0083852 | 0.36843 | 580.42 | 608.103
Populus deltoides | LI2B |  | RACiR 800 to 300 | 477.0464 | 32.8149 | 25.0046 | 999.93 | 0.0083816 | 0.36862 | 578.062 | 605.641
Populus deltoides | LI2B |  | RACiR 800 to 300 | 474.7818 | 32.8396 | 24.9984 | 1000.24 | 0.0083826 | 0.36903 | 575.701 | 603.288
Populus deltoides | LI2B |  | RACiR 800 to 300 | 473.4607 | 32.6667 | 25.0069 | 999.93 | 0.0083815 | 0.3685 | 574.021 | 601.429
Populus deltoides | LI2B |  | RACiR 800 to 300 | 471.1043 | 32.628 | 25.0152 | 999.967 | 0.0083829 | 0.36806 | 571.627 | 598.983
Populus deltoides | LI2B |  | RACiR 800 to 300 | 468.8633 | 32.5611 | 25.0198 | 1000.26 | 0.0083849 | 0.36788 | 569.204 | 596.482
Populus deltoides | LI2B |  | RACiR 800 to 300 | 467.2943 | 32.5671 | 25.0112 | 999.904 | 0.0083851 | 0.36838 | 567.477 | 594.753
Populus deltoides | LI2B |  | RACiR 800 to 300 | 465.0789 | 32.5643 | 25.0024 | 999.985 | 0.0083853 | 0.3689 | 565.061 | 592.321
Populus deltoides | LI2B |  | RACiR 800 to 300 | 462.9432 | 32.4515 | 25.0043 | 999.873 | 0.008383 | 0.36861 | 562.644 | 589.778
Populus deltoides | LI2B |  | RACiR 800 to 300 | 461.0518 | 32.5135 | 25.0067 | 1000.3 | 0.0083828 | 0.36846 | 560.919 | 588.105
Populus deltoides | LI2B |  | RACiR 800 to 300 | 459.1097 | 32.3539 | 25.0074 | 999.958 | 0.0083828 | 0.36843 | 558.503 | 585.518
Populus deltoides | LI2B |  | RACiR 800 to 300 | 456.6055 | 32.4194 | 25.0061 | 999.945 | 0.0083837 | 0.36857 | 556.083 | 583.148
Populus deltoides | LI2B |  | RACiR 800 to 300 | 455.2392 | 32.3121 | 25.002 | 999.964 | 0.0083827 | 0.36875 | 554.34 | 581.291
Populus deltoides | LI2B |  | RACiR 800 to 300 | 452.8156 | 32.3409 | 25.0075 | 1000.32 | 0.0083948 | 0.36894 | 551.906 | 578.881
Populus deltoides | LI2B |  | RACiR 800 to 300 | 450.5959 | 32.3637 | 25.0098 | 1000.02 | 0.0084201 | 0.36986 | 549.482 | 576.48
Populus deltoides | LI2B |  | RACiR 800 to 300 | 448.913 | 32.3705 | 25.0148 | 999.956 | 0.0084326 | 0.36995 | 547.773 | 574.777
Populus deltoides | LI2B |  | RACiR 800 to 300 | 446.4807 | 32.3822 | 25.0144 | 999.899 | 0.0084325 | 0.36964 | 545.398 | 572.398
Populus deltoides | LI2B |  | RACiR 800 to 300 | 444.4551 | 32.1873 | 25.011 | 999.955 | 0.0084147 | 0.36873 | 543.015 | 569.797
Populus deltoides | LI2B |  | RACiR 800 to 300 | 442.3505 | 32.2647 | 25.0121 | 999.982 | 0.0084016 | 0.36792 | 541.276 | 568.121
Populus deltoides | LI2B |  | RACiR 800 to 300 | 440.1754 | 32.1318 | 25.0086 | 999.936 | 0.008379 | 0.36708 | 538.882 | 565.573
Populus deltoides | LI2B |  | RACiR 800 to 300 | 437.4539 | 32.2304 | 25.0026 | 1000.31 | 0.0083589 | 0.36657 | 536.482 | 563.245
Populus deltoides | LI2B |  | RACiR 800 to 300 | 436.1313 | 32.0928 | 24.9985 | 999.899 | 0.0083472 | 0.36642 | 534.774 | 561.386
Populus deltoides | LI2B |  | RACiR 800 to 300 | 433.7557 | 32.1255 | 24.9986 | 999.95 | 0.0083482 | 0.3667 | 532.361 | 558.991
Populus deltoides | LI2B |  | RACiR 800 to 300 | 431.5705 | 32.0731 | 25.0033 | 999.949 | 0.0083522 | 0.36678 | 529.964 | 556.533
Populus deltoides | LI2B |  | RACiR 800 to 300 | 430.1375 | 31.9868 | 25.0061 | 999.922 | 0.0083541 | 0.36678 | 528.263 | 554.74
Populus deltoides | LI2B |  | RACiR 800 to 300 | 428.0869 | 31.9278 | 24.9978 | 999.954 | 0.0083568 | 0.36742 | 525.831 | 552.236
Populus deltoides | LI2B |  | RACiR 800 to 300 | 426.1442 | 31.7833 | 24.9984 | 1000.02 | 0.0083571 | 0.36746 | 523.436 | 549.688
Populus deltoides | LI2B |  | RACiR 800 to 300 | 424.0679 | 31.9455 | 24.9964 | 999.928 | 0.00836 | 0.36778 | 521.676 | 548.079
Populus deltoides | LI2B |  | RACiR 800 to 300 | 421.6997 | 31.898 | 25.0077 | 999.955 | 0.0083605 | 0.36717 | 519.288 | 545.632
Populus deltoides | LI2B |  | RACiR 800 to 300 | 419.2556 | 31.9647 | 25.0039 | 1000.27 | 0.0083622 | 0.36747 | 516.893 | 543.288
Populus deltoides | LI2B |  | RACiR 800 to 300 | 418.1806 | 31.7537 | 25.0025 | 999.929 | 0.0083618 | 0.36751 | 515.201 | 541.379
Populus deltoides | LI2B |  | RACiR 800 to 300 | 415.9385 | 31.7186 | 25.0013 | 1000 | 0.0083584 | 0.3674 | 512.837 | 538.969
Populus deltoides | LI2B |  | RACiR 800 to 300 | 413.8145 | 31.5995 | 25.0088 | 999.955 | 0.008359 | 0.36702 | 510.437 | 536.436
Populus deltoides | LI2B |  | RACiR 800 to 300 | 412.0954 | 31.5905 | 25.0129 | 999.889 | 0.0083563 | 0.36666 | 508.747 | 534.73
Populus deltoides | LI2B |  | RACiR 800 to 300 | 409.6032 | 31.6182 | 25.0162 | 1000.01 | 0.0083566 | 0.36644 | 506.332 | 532.326
Populus deltoides | LI2B |  | RACiR 800 to 300 | 407.3276 | 31.5787 | 25.0162 | 999.931 | 0.0083562 | 0.36638 | 503.911 | 529.855
Populus deltoides | LI2B |  | RACiR 800 to 300 | 405.3469 | 31.7086 | 25.0148 | 999.991 | 0.0083585 | 0.36654 | 502.203 | 528.265
Populus deltoides | LI2B |  | RACiR 800 to 300 | 403.257 | 31.6333 | 25.0144 | 1000.32 | 0.0083581 | 0.36653 | 499.859 | 525.837
Populus deltoides | LI2B |  | RACiR 800 to 300 | 401.0309 | 31.6237 | 25.0128 | 999.928 | 0.0083577 | 0.36658 | 497.542 | 523.495
Populus deltoides | LI2B |  | RACiR 800 to 300 | 399.8629 | 31.4469 | 25.0139 | 999.94 | 0.0083571 | 0.36649 | 495.882 | 521.652
Populus deltoides | LI2B |  | RACiR 800 to 300 | 397.7727 | 31.3748 | 25.0099 | 999.994 | 0.0083558 | 0.36663 | 493.506 | 519.19
Populus deltoides | LI2B |  | RACiR 800 to 300 | 395.7692 | 31.2508 | 25.0116 | 999.959 | 0.0083553 | 0.36651 | 491.144 | 516.693
Populus deltoides | LI2B |  | RACiR 800 to 300 | 394.1344 | 31.2402 | 25.0098 | 1000.3 | 0.0083529 | 0.36646 | 489.452 | 514.983
Populus deltoides | LI2B |  | RACiR 800 to 300 | 391.8514 | 31.2429 | 25.0076 | 999.95 | 0.0083531 | 0.36657 | 487.095 | 512.612
Populus deltoides | LI2B |  | RACiR 800 to 300 | 389.8216 | 31.1223 | 25.0077 | 999.959 | 0.0083504 | 0.3664 | 484.727 | 510.114
Populus deltoides | LI2B |  | RACiR 800 to 300 | 387.8202 | 31.2237 | 25.0109 | 999.98 | 0.0083507 | 0.36622 | 483.007 | 508.484
Populus deltoides | LI2B |  | RACiR 800 to 300 | 385.8189 | 31.084 | 25.0149 | 999.974 | 0.008348 | 0.36581 | 480.679 | 506.006
Populus deltoides | LI2B |  | RACiR 800 to 300 | 383.4591 | 31.1192 | 25.0123 | 999.989 | 0.0083482 | 0.36594 | 478.328 | 503.674
Populus deltoides | LI2B |  | RACiR 800 to 300 | 382.4999 | 30.8981 | 25.0074 | 999.967 | 0.0083448 | 0.36605 | 476.703 | 501.823
Populus deltoides | LI2B |  | RACiR 800 to 300 | 380.371 | 30.836 | 25.0077 | 999.956 | 0.0083464 | 0.36611 | 474.34 | 499.382
Populus deltoides | LI2B |  | RACiR 800 to 300 | 378.3119 | 30.7241 | 25.0106 | 999.974 | 0.0083453 | 0.36589 | 471.98 | 496.899
Populus deltoides | LI2B |  | RACiR 800 to 300 | 376.6375 | 30.7232 | 25.0119 | 1000 | 0.0083455 | 0.36583 | 470.28 | 495.189
Populus deltoides | LI2B |  | RACiR 800 to 300 | 374.5163 | 30.6452 | 25.0124 | 999.993 | 0.0083433 | 0.36568 | 467.931 | 492.75
Populus deltoides | LI2B |  | RACiR 800 to 300 | 372.823 | 30.4081 | 25.015 | 1000.01 | 0.0083421 | 0.36546 | 465.597 | 490.169
Populus deltoides | LI2B |  | RACiR 800 to 300 | 371.0087 | 30.4512 | 25.0156 | 1000.07 | 0.0083408 | 0.36531 | 463.897 | 488.5
Populus deltoides | LI2B |  | RACiR 800 to 300 | 369.2816 | 30.2522 | 25.0138 | 999.993 | 0.0083412 | 0.36545 | 461.544 | 485.938
Populus deltoides | LI2B |  | RACiR 800 to 300 | 366.7807 | 30.3166 | 25.0115 | 1000.02 | 0.008341 | 0.36554 | 459.139 | 483.581
Populus deltoides | LI2B |  | RACiR 800 to 300 | 365.6153 | 30.1438 | 25.01 | 999.963 | 0.0083402 | 0.36561 | 457.452 | 481.715
Populus deltoides | LI2B |  | RACiR 800 to 300 | 363.182 | 30.1663 | 25.0083 | 1000.01 | 0.0083389 | 0.36558 | 455.029 | 479.301
Populus deltoides | LI2B |  | RACiR 800 to 300 | 360.811 | 30.1582 | 25.0062 | 999.979 | 0.0083319 | 0.3653 | 452.646 | 476.893
Populus deltoides | LI2B |  | RACiR 800 to 300 | 359.0999 | 30.1365 | 25.008 | 999.992 | 0.0083224 | 0.36472 | 450.972 | 475.188
Populus deltoides | LI2B |  | RACiR 800 to 300 | 356.8586 | 30.0939 | 25.0073 | 999.944 | 0.0083127 | 0.36436 | 448.641 | 472.795
Populus deltoides | LI2B |  | RACiR 800 to 300 | 355.0003 | 29.9037 | 25.0108 | 999.974 | 0.0083077 | 0.36401 | 446.294 | 470.247
Populus deltoides | LI2B |  | RACiR 800 to 300 | 353.3283 | 29.9386 | 25.0075 | 999.967 | 0.0083083 | 0.36425 | 444.621 | 468.598
Populus deltoides | LI2B |  | RACiR 800 to 300 | 351.4558 | 29.781 | 25.0103 | 1000.04 | 0.0083104 | 0.36424 | 442.272 | 466.083
Populus deltoides | LI2B |  | RACiR 800 to 300 | 349.2135 | 29.78 | 25.0092 | 999.959 | 0.0083134 | 0.36446 | 439.924 | 463.719
Populus deltoides | LI2B |  | RACiR 800 to 300 | 348.1072 | 29.5301 | 25.0191 | 999.945 | 0.0083124 | 0.36386 | 438.248 | 461.788
Populus deltoides | LI2B |  | RACiR 800 to 300 | 346.146 | 29.3983 | 25.0204 | 999.994 | 0.0083136 | 0.36384 | 435.881 | 459.276
Populus deltoides | LI2B |  | RACiR 800 to 300 | 344.2385 | 29.2306 | 25.0203 | 999.949 | 0.008311 | 0.36367 | 433.503 | 456.718
Populus deltoides | LI2B |  | RACiR 800 to 300 | 342.7379 | 29.2141 | 25.011 | 1000 | 0.0083107 | 0.36418 | 431.795 | 454.987
Populus deltoides | LI2B |  | RACiR 800 to 300 | 340.4493 | 29.2043 | 25.0078 | 999.965 | 0.0083099 | 0.36431 | 429.393 | 452.56
Populus deltoides | LI2B |  | RACiR 800 to 300 | 338.4451 | 29.1068 | 25.0029 | 999.986 | 0.0083076 | 0.36449 | 427.025 | 450.083
Populus deltoides | LI2B |  | RACiR 800 to 300 | 336.4254 | 29.2058 | 25.0086 | 999.943 | 0.0083069 | 0.36411 | 425.327 | 448.473
Populus deltoides | LI2B |  | RACiR 800 to 300 | 334.3625 | 29.1203 | 25.0093 | 999.988 | 0.0083066 | 0.36405 | 422.993 | 446.043
Populus deltoides | LI2B |  | RACiR 800 to 300 | 332.1079 | 29.1253 | 25.0065 | 999.985 | 0.0083073 | 0.36419 | 420.667 | 443.707
Populus deltoides | LI2B |  | RACiR 800 to 300 | 331.0775 | 28.9116 | 25.0058 | 999.948 | 0.0083056 | 0.36416 | 419.025 | 441.847
Populus deltoides | LI2B |  | RACiR 800 to 300 | 329.1088 | 28.8013 | 25.005 | 1000.03 | 0.0083047 | 0.36415 | 416.705 | 439.403
Populus deltoides | LI2B |  | RACiR 800 to 300 | 327.0356 | 28.6956 | 25.0108 | 1000.04 | 0.0083057 | 0.3639 | 414.351 | 436.932
Populus deltoides | LI2B |  | RACiR 800 to 300 | 325.5747 | 28.6298 | 25.0093 | 999.95 | 0.0083031 | 0.36383 | 412.689 | 435.197
Populus deltoides | LI2B |  | RACiR 800 to 300 | 323.4019 | 28.5835 | 25.0091 | 1000.02 | 0.0083027 | 0.36377 | 410.351 | 432.798
Populus deltoides | LI2B |  | RACiR 800 to 300 | 321.5496 | 28.4078 | 25.0114 | 999.981 | 0.0083003 | 0.36346 | 408.039 | 430.3
Populus deltoides | LI2B |  | RACiR 800 to 300 | 319.6953 | 28.4948 | 25.0105 | 999.953 | 0.0083013 | 0.36357 | 406.359 | 428.697
Populus deltoides | LI2B |  | RACiR 800 to 300 | 317.9233 | 28.3361 | 25.012 | 999.971 | 0.0083065 | 0.36371 | 404.077 | 426.248
Populus deltoides | LI2B |  | RACiR 800 to 300 | 316.0772 | 28.2853 | 25.0088 | 999.992 | 0.0083295 | 0.36494 | 401.776 | 423.891
Populus deltoides | LI2B |  | RACiR 800 to 300 | 315.0691 | 28.0732 | 25.0136 | 1000.02 | 0.0083402 | 0.36506 | 400.14 | 422.042
Populus deltoides | LI2B |  | RACiR 800 to 300 | 313.004 | 28.0057 | 25.0116 | 1000.01 | 0.0083444 | 0.36521 | 397.809 | 419.631
Populus deltoides | LI2B |  | RACiR 800 to 300 | 310.8002 | 27.9141 | 25.0086 | 999.961 | 0.0083247 | 0.36428 | 395.499 | 417.21
Populus deltoides | LI2B |  | RACiR 800 to 300 | 309.152 | 27.8541 | 25.0088 | 999.95 | 0.0083047 | 0.36329 | 393.859 | 415.495
Populus deltoides | LI2B |  | RACiR 800 to 300 | 306.6897 | 27.806 | 25.0127 | 999.937 | 0.0082744 | 0.36171 | 391.549 | 413.113
Populus deltoides | LI2B |  | RACiR 800 to 300 | 304.8267 | 27.5869 | 25.0172 | 999.998 | 0.0082517 | 0.36064 | 389.262 | 410.589
Populus deltoides | LI2B |  | RACiR 800 to 300 | 303.1657 | 27.6435 | 25.0047 | 1000.04 | 0.0082433 | 0.36117 | 387.588 | 408.958
Populus deltoides | LI2B |  | RACiR 800 to 300 | 301.4803 | 27.5149 | 24.9985 | 999.967 | 0.0082488 | 0.36206 | 385.3 | 406.533
Populus deltoides | LI2B |  | RACiR 800 to 300 | 299.441 | 27.5074 | 24.9952 | 999.996 | 0.0082572 | 0.36285 | 383.015 | 404.23
Populus deltoides | LI2B |  | RACiR 800 to 300 | 298.3459 | 27.3256 | 25.0001 | 1000 | 0.0082614 | 0.36286 | 381.389 | 402.417
Populus deltoides | LI2B |  | RACiR 800 to 300 | 296.5009 | 27.1652 | 25.0099 | 1000 | 0.0082663 | 0.3626 | 379.12 | 399.978
Populus deltoides | LI2B |  | RACiR 800 to 300 | 294.7685 | 27.0171 | 25.0037 | 1000.03 | 0.0082674 | 0.36306 | 376.831 | 397.53
Populus deltoides | LI2B |  | RACiR 800 to 300 | 293.4119 | 26.9128 | 25.0075 | 1000.07 | 0.0082654 | 0.36274 | 375.223 | 395.81
Populus deltoides | LI2B |  | RACiR 800 to 300 | 291.494 | 26.8081 | 25.0035 | 999.979 | 0.0082671 | 0.36311 | 372.887 | 393.358
Populus deltoides | LI2B |  | RACiR 800 to 300 | 290.0533 | 26.4859 | 25.0099 | 1000.01 | 0.0082678 | 0.36283 | 370.579 | 390.718
Populus deltoides | LI2B |  | RACiR 800 to 300 | 288.4236 | 26.4841 | 25.0099 | 999.984 | 0.0082697 | 0.36292 | 368.888 | 389.016
Populus deltoides | LI2B |  | RACiR 800 to 300 | 286.502 | 26.3608 | 25.009 | 1000.01 | 0.0082687 | 0.36289 | 366.584 | 386.579
Populus deltoides | LI2B |  | RACiR 800 to 300 | 284.0311 | 26.45 | 25.0068 | 1000.03 | 0.0082705 | 0.3631 | 364.259 | 384.329
Populus deltoides | LI2B |  | RACiR 800 to 300 | 282.764 | 26.334 | 25.0047 | 999.984 | 0.0082702 | 0.36321 | 362.616 | 382.563
Populus deltoides | LI2B |  | RACiR 800 to 300 | 280.9727 | 26.1766 | 25.0011 | 1000.01 | 0.008268 | 0.36331 | 360.32 | 380.098
Populus deltoides | LI2B |  | RACiR 800 to 300 | 279.2075 | 25.9806 | 25.0042 | 1000.02 | 0.0082658 | 0.36302 | 358.03 | 377.601
Populus deltoides | LI2B |  | RACiR 800 to 300 | 277.9266 | 25.8513 | 25.0052 | 1000.01 | 0.008263 | 0.36281 | 356.403 | 375.837
Populus deltoides | LI2B |  | RACiR 800 to 300 | 275.8732 | 25.7551 | 25.0107 | 1000.03 | 0.0082627 | 0.36247 | 354.108 | 373.431
Populus deltoides | LI2B |  | RACiR 800 to 300 | 274.1715 | 25.5356 | 25.0116 | 999.97 | 0.0082611 | 0.36225 | 351.8 | 370.892
Populus deltoides | LI2B |  | RACiR 800 to 300 | 272.4794 | 25.5692 | 25.0085 | 1000.02 | 0.0082598 | 0.36233 | 350.144 | 369.26
Populus deltoides | LI2B |  | RACiR 800 to 300 | 270.8976 | 25.3569 | 25.0022 | 1000 | 0.0082584 | 0.36259 | 347.877 | 366.77
Populus deltoides | LI2B |  | RACiR 800 to 300 | 268.7974 | 25.3005 | 25.0019 | 1000.03 | 0.0082592 | 0.36266 | 345.556 | 364.379
Populus deltoides | LI2B |  | RACiR 800 to 300 | 267.7011 | 25.0997 | 25.0047 | 1000.01 | 0.0082583 | 0.36247 | 343.915 | 362.531
Populus deltoides | LI2B |  | RACiR 800 to 300 | 265.5921 | 25.0247 | 25.0055 | 999.993 | 0.0082559 | 0.3623 | 341.582 | 360.111
Populus deltoides | LI2B |  | RACiR 800 to 300 | 263.8641 | 24.8325 | 25.0053 | 1000 | 0.0082549 | 0.36223 | 339.291 | 357.616
Populus deltoides | LI2B |  | RACiR 800 to 300 | 262.6464 | 24.6904 | 25.004 | 1000.01 | 0.0082532 | 0.36222 | 337.65 | 355.826
Populus deltoides | LI2B |  | RACiR 800 to 300 | 260.7661 | 24.5214 | 25.0107 | 999.991 | 0.0082546 | 0.36188 | 335.325 | 353.32
Populus deltoides | LI2B |  | RACiR 800 to 300 | 259.0705 | 24.3349 | 25.0038 | 999.97 | 0.0082522 | 0.36216 | 333.01 | 350.808
Populus deltoides | LI2B |  | RACiR 800 to 300 | 257.1544 | 24.4243 | 25.0073 | 999.969 | 0.0082525 | 0.36194 | 331.346 | 349.224
Populus deltoides | LI2B |  | RACiR 800 to 300 | 255.31 | 24.2918 | 25.0017 | 1000 | 0.0082501 | 0.36209 | 329.054 | 346.786
Populus deltoides | LI2B |  | RACiR 800 to 300 | 253.0659 | 24.2643 | 25.0071 | 999.97 | 0.0082509 | 0.36173 | 326.757 | 344.449
Populus deltoides | LI2B |  | RACiR 800 to 300 | 252.2309 | 24.0036 | 25.0019 | 1000.01 | 0.008248 | 0.36184 | 325.149 | 342.574
Populus deltoides | LI2B |  | RACiR 800 to 300 | 250.6665 | 23.7813 | 24.9976 | 999.986 | 0.0082474 | 0.36205 | 322.884 | 340.075
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 345.5852 | 36.1263 | 25.3593 | 1000.07 | 0.0082505 | 0.40703 | 442.327 | 473.391
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 343.7551 | 35.7074 | 25.3396 | 1000.11 | 0.0082287 | 0.40673 | 439.47 | 470.085
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 341.227 | 35.5146 | 25.3164 | 1000.16 | 0.0082014 | 0.40614 | 436.523 | 466.917
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 338.8597 | 35.2483 | 25.2889 | 1000.09 | 0.0081713 | 0.40568 | 433.526 | 463.626
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 336.3801 | 35.0533 | 25.2624 | 1000.05 | 0.0081435 | 0.40537 | 430.555 | 460.428
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 333.8163 | 34.8628 | 25.2312 | 1000.06 | 0.0081065 | 0.40467 | 427.593 | 457.241
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 331.6428 | 34.4886 | 25.1977 | 1000.15 | 0.0080687 | 0.40394 | 424.577 | 453.819
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 329.1583 | 34.2005 | 25.1716 | 999.93 | 0.008035 | 0.40301 | 421.503 | 450.424
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 326.1068 | 34.1025 | 25.144 | 1000.09 | 0.0079947 | 0.40167 | 418.417 | 447.204
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 323.5114 | 33.8094 | 25.1127 | 1000.09 | 0.0079501 | 0.40027 | 415.315 | 443.774
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 320.921 | 33.5044 | 25.0775 | 999.732 | 0.0079007 | 0.39873 | 412.212 | 440.328
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 317.5824 | 33.4646 | 25.0384 | 1000.03 | 0.0078462 | 0.39699 | 409.06 | 437.094
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 314.6915 | 33.2381 | 25.004 | 1000.04 | 0.007791 | 0.3949 | 405.969 | 433.738
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 311.7399 | 33.0676 | 24.9746 | 1000.03 | 0.0077443 | 0.3931 | 402.899 | 430.461
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 309.087 | 32.9581 | 24.9295 | 1000.09 | 0.0077069 | 0.39323 | 399.847 | 427.268
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 306.9052 | 32.6444 | 24.8865 | 1000.12 | 0.0076701 | 0.39323 | 396.784 | 423.862
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 304.4159 | 32.4318 | 24.8573 | 1000.06 | 0.0076438 | 0.3932 | 393.683 | 420.522
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 302.4989 | 32.3401 | 24.8215 | 1000.07 | 0.0076608 | 0.39741 | 390.557 | 417.292
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 301.5626 | 32.085 | 24.7968 | 999.995 | 0.0077222 | 0.40455 | 387.45 | 413.937
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 300.6958 | 31.8431 | 24.7773 | 999.978 | 0.0077992 | 0.41269 | 384.329 | 410.588
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 299.4254 | 31.6641 | 24.7587 | 1000.05 | 0.0078672 | 0.42014 | 381.163 | 407.248
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 297.7201 | 31.4132 | 24.7536 | 999.753 | 0.0079123 | 0.42458 | 378.003 | 403.837
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 295.4336 | 31.2426 | 24.7556 | 1000.06 | 0.0079393 | 0.42674 | 374.882 | 400.538
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 292.7628 | 31.1444 | 24.7629 | 1000.05 | 0.0079549 | 0.42743 | 371.811 | 397.356
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 290.3894 | 30.8358 | 24.77 | 999.738 | 0.0079589 | 0.42716 | 368.718 | 393.94
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 288.013 | 30.5208 | 24.7722 | 1000.05 | 0.0079589 | 0.42682 | 365.614 | 390.501
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 285.0268 | 30.4513 | 24.7717 | 999.961 | 0.0079534 | 0.42629 | 362.501 | 387.298
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 282.3076 | 30.2201 | 24.7791 | 1000.03 | 0.0079466 | 0.42497 | 359.418 | 383.966
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 278.7522 | 30.0172 | 24.7848 | 1000.01 | 0.0079379 | 0.42359 | 355.559 | 379.886
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 276.1135 | 30.1543 | 24.7894 | 1000 | 0.0079365 | 0.42293 | 353.318 | 377.761
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 273.9066 | 29.7664 | 24.799 | 999.922 | 0.0079327 | 0.42175 | 350.356 | 374.394
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 272.0787 | 29.2055 | 24.8062 | 1000.03 | 0.0079278 | 0.42059 | 347.355 | 370.812
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 269.407 | 28.9889 | 24.8078 | 999.971 | 0.0079229 | 0.41996 | 344.223 | 367.443
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 266.6155 | 28.8418 | 24.8062 | 999.974 | 0.0079187 | 0.41962 | 341.082 | 364.137
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 263.8058 | 28.7359 | 24.8075 | 999.958 | 0.007919 | 0.41939 | 338.006 | 360.938
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 260.966 | 28.6371 | 24.8152 | 1000.03 | 0.0079201 | 0.41882 | 334.977 | 357.793
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 258.8665 | 28.2305 | 24.8246 | 999.995 | 0.0079212 | 0.41806 | 331.992 | 354.385
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 256.5632 | 27.9109 | 24.8293 | 999.962 | 0.0079223 | 0.41767 | 328.941 | 350.996
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 253.8224 | 27.7803 | 24.8318 | 999.964 | 0.007925 | 0.41758 | 325.85 | 347.758
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 251.6008 | 27.4123 | 24.839 | 999.974 | 0.0079252 | 0.41696 | 322.803 | 344.329
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 249.2998 | 27.0956 | 24.8407 | 999.97 | 0.0079272 | 0.41687 | 319.707 | 340.898
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 246.4701 | 27.0222 | 24.8345 | 1000.04 | 0.0079269 | 0.41728 | 316.576 | 337.674
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 244.0564 | 26.7654 | 24.8334 | 1000.01 | 0.0079265 | 0.41729 | 313.492 | 334.317
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 241.4638 | 26.5705 | 24.8417 | 1000.02 | 0.0079313 | 0.41691 | 310.445 | 331.059
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 238.5317 | 26.4948 | 24.8522 | 1000.05 | 0.0079331 | 0.41615 | 307.401 | 327.921
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 236.4229 | 26.1444 | 24.8556 | 1000.04 | 0.007938 | 0.41627 | 304.386 | 324.54
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 234.3082 | 25.7837 | 24.8611 | 1000.03 | 0.0079437 | 0.4162 | 301.37 | 321.147
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 231.6096 | 25.6278 | 24.8716 | 999.98 | 0.007949 | 0.41574 | 298.32 | 317.926
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 229.27 | 25.354 | 24.8763 | 1000.02 | 0.0079541 | 0.41572 | 295.277 | 314.594
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 226.7759 | 25.1449 | 24.8792 | 999.999 | 0.0079576 | 0.41567 | 292.237 | 311.328
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 223.9402 | 25.0423 | 24.89 | 999.987 | 0.0079622 | 0.41511 | 289.188 | 308.161
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 221.8152 | 24.6952 | 24.8959 | 999.99 | 0.0079675 | 0.41503 | 286.194 | 304.805
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 219.677 | 24.3753 | 24.8965 | 999.938 | 0.0079725 | 0.41532 | 283.194 | 301.469
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 217.0994 | 24.1876 | 24.9039 | 999.908 | 0.0079766 | 0.41502 | 280.158 | 298.229
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 215.0455 | 23.7726 | 24.9141 | 1000.01 | 0.007984 | 0.41475 | 277.101 | 294.741
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 212.7677 | 23.4184 | 24.928 | 999.991 | 0.0079894 | 0.41401 | 274.03 | 291.3
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 210.0107 | 23.2901 | 24.936 | 999.96 | 0.0079953 | 0.41379 | 270.942 | 288.067
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 207.6653 | 23.0223 | 24.9379 | 999.941 | 0.008 | 0.41397 | 267.874 | 284.716
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 205.0276 | 22.8351 | 24.9502 | 1000.01 | 0.0080042 | 0.41329 | 264.83 | 281.467
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 202.2696 | 22.732 | 24.9567 | 999.933 | 0.0080088 | 0.41313 | 261.792 | 278.31
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 200.2124 | 22.3606 | 24.9627 | 1000.28 | 0.0080152 | 0.41307 | 258.798 | 274.93
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 198.115 | 21.9613 | 24.9776 | 999.939 | 0.0080205 | 0.41226 | 255.801 | 271.518
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 195.4902 | 21.8052 | 24.9818 | 999.939 | 0.0080257 | 0.4123 | 252.739 | 268.281
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 193.2563 | 21.5495 | 24.9732 | 999.992 | 0.0080286 | 0.41317 | 249.716 | 264.988
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 191.0198 | 21.2392 | 24.9791 | 999.946 | 0.008033 | 0.41304 | 246.698 | 261.644
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 188.5689 | 20.9689 | 24.9964 | 999.977 | 0.0080401 | 0.41217 | 243.659 | 258.317
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 186.3022 | 20.6372 | 25.0051 | 999.962 | 0.008043 | 0.41168 | 240.601 | 254.912
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 183.8571 | 20.3871 | 25.0121 | 999.96 | 0.0080455 | 0.41123 | 237.556 | 251.599
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 181.3257 | 20.2164 | 25.0145 | 999.945 | 0.0080486 | 0.41114 | 234.568 | 248.423
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 179.1903 | 19.8214 | 25.0179 | 1000.19 | 0.0080372 | 0.40989 | 231.576 | 245.018
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 176.887 | 19.4334 | 25.0215 | 999.909 | 0.0080133 | 0.40758 | 228.552 | 241.583
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 174.1835 | 19.2466 | 25.0217 | 1000.21 | 0.0079944 | 0.40598 | 225.527 | 238.351
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 171.8679 | 18.9796 | 25.0123 | 999.944 | 0.0079876 | 0.40596 | 222.5 | 235.036
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 169.5134 | 18.757 | 25.0054 | 1000.14 | 0.0079878 | 0.40634 | 219.498 | 231.791
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 167.0947 | 18.5876 | 24.9997 | 999.914 | 0.0079901 | 0.40694 | 216.537 | 228.646
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 165.3195 | 18.1737 | 24.99 | 1000.26 | 0.0079948 | 0.40807 | 213.567 | 225.248
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 163.0938 | 17.8687 | 24.9946 | 999.923 | 0.007999 | 0.40796 | 210.557 | 221.916
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 160.6856 | 17.6755 | 24.9906 | 999.871 | 0.008002 | 0.40843 | 207.57 | 218.718
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 158.7377 | 17.3089 | 24.9815 | 999.888 | 0.0080042 | 0.40928 | 204.581 | 215.348
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 156.4814 | 17.058 | 24.9791 | 1000.07 | 0.0080094 | 0.40982 | 201.604 | 212.103
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 154.1224 | 16.8376 | 24.98 | 999.884 | 0.008011 | 0.40984 | 198.654 | 208.915
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 152.1172 | 16.4544 | 24.9818 | 999.909 | 0.0080095 | 0.40952 | 195.698 | 205.56
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 150.1842 | 16.0572 | 24.9807 | 999.897 | 0.0080105 | 0.40959 | 192.741 | 202.189
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 147.7902 | 15.8324 | 24.978 | 999.903 | 0.008011 | 0.40977 | 189.727 | 198.933
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 145.8344 | 15.4282 | 24.9742 | 1000.21 | 0.008012 | 0.41008 | 186.705 | 195.49
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 143.7886 | 15.0418 | 24.9771 | 1000.07 | 0.0080133 | 0.40985 | 183.691 | 192.073
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 141.196 | 14.8531 | 24.9825 | 1000.2 | 0.0080131 | 0.40928 | 180.637 | 188.812
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 139.0754 | 14.5091 | 24.9805 | 999.898 | 0.0080121 | 0.40932 | 177.622 | 185.436
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 136.6878 | 14.2838 | 24.9794 | 999.867 | 0.008014 | 0.40947 | 174.618 | 182.189
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 134.1918 | 14.0939 | 24.9822 | 999.93 | 0.0080149 | 0.40922 | 171.627 | 178.991
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 132.2436 | 13.6794 | 24.9865 | 999.861 | 0.0080137 | 0.40874 | 168.661 | 175.594
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 130.082 | 13.3646 | 24.9876 | 1000.08 | 0.0080146 | 0.40861 | 165.69 | 172.291
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 127.6704 | 13.1353 | 24.9879 | 1000.21 | 0.0080135 | 0.40839 | 162.683 | 169.037
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 125.7317 | 12.7368 | 24.9826 | 1000.14 | 0.0080095 | 0.40848 | 159.711 | 165.65
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 123.6276 | 12.4044 | 24.9793 | 999.869 | 0.0080101 | 0.40871 | 156.725 | 162.314
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 121.3071 | 12.1201 | 24.9856 | 1000.16 | 0.0080101 | 0.40806 | 153.708 | 158.995
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 119.1358 | 11.778 | 24.9876 | 999.83 | 0.0080087 | 0.40776 | 150.67 | 155.598
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 116.8024 | 11.5067 | 24.9899 | 1000.04 | 0.0080082 | 0.40747 | 147.64 | 152.279
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 114.2516 | 11.3413 | 24.9854 | 1000.17 | 0.0080059 | 0.40755 | 144.622 | 149.078
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 112.2683 | 10.948 | 24.9885 | 999.864 | 0.0080067 | 0.40729 | 141.644 | 145.69
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 110.3167 | 10.5548 | 24.9867 | 999.874 | 0.0080085 | 0.40749 | 138.665 | 142.301
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 107.8569 | 10.3534 | 24.9847 | 1000.16 | 0.0080073 | 0.40745 | 135.66 | 139.077
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 105.6393 | 10.035 | 24.9897 | 1000.17 | 0.0080033 | 0.40669 | 132.659 | 135.74
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 103.3671 | 9.7699 | 24.9876 | 999.969 | 0.0080054 | 0.40695 | 129.667 | 132.466
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 101.0662 | 9.5231 | 24.986 | 1000.17 | 0.0080056 | 0.40702 | 126.703 | 129.238
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 99.0619 | 9.13 | 24.9926 | 999.825 | 0.0080036 | 0.40627 | 123.729 | 125.854
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 97.038 | 8.7638 | 24.9944 | 999.8 | 0.0080039 | 0.40603 | 120.769 | 122.511
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 94.8557 | 8.4786 | 24.9881 | 1000.16 | 0.0080057 | 0.40658 | 117.803 | 119.243
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 92.8172 | 8.1124 | 24.9909 | 1000.16 | 0.0080063 | 0.40636 | 114.826 | 115.883
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 90.6158 | 7.7956 | 24.9996 | 999.976 | 0.0080033 | 0.40539 | 111.844 | 112.567
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 88.4427 | 7.4837 | 25.0005 | 1000.15 | 0.0080033 | 0.40523 | 108.859 | 109.253
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 86.4894 | 7.0825 | 25.0005 | 999.839 | 0.0080049 | 0.40529 | 105.864 | 105.84
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 84.1949 | 6.8077 | 25.0057 | 999.838 | 0.0080047 | 0.4048 | 102.862 | 102.546
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 81.7369 | 6.6183 | 25.0063 | 999.777 | 0.0080039 | 0.40459 | 99.8916 | 99.3689
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 79.8664 | 6.2006 | 25.0029 | 999.832 | 0.0080027 | 0.40469 | 96.9361 | 95.9789
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 77.9849 | 5.7785 | 24.9982 | 999.979 | 0.0080003 | 0.40482 | 93.958 | 92.5617
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 75.6189 | 5.5288 | 25.0044 | 999.84 | 0.0079976 | 0.40405 | 90.951 | 89.2875
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 73.5217 | 5.1753 | 25.0062 | 999.814 | 0.0079936 | 0.40345 | 87.949 | 85.9147
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 71.2881 | 4.8888 | 25.005 | 999.848 | 0.007993 | 0.40343 | 84.9541 | 82.6159
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 68.8965 | 4.681 | 25.0031 | 1000.21 | 0.0079956 | 0.4037 | 81.986 | 79.4234
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 66.9283 | 4.2971 | 25.0019 | 1000.1 | 0.0079924 | 0.40348 | 79.024 | 76.0601
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 65.1302 | 3.8486 | 25.0056 | 999.962 | 0.0079937 | 0.40322 | 76.0746 | 72.6459
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 62.8795 | 3.5762 | 25.0099 | 1000.15 | 0.0079951 | 0.40288 | 73.1056 | 69.3881
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 60.8067 | 3.2302 | 25.0125 | 999.797 | 0.0079931 | 0.40244 | 70.1335 | 66.0528
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 58.8514 | 2.8389 | 25.0128 | 999.806 | 0.0079917 | 0.40229 | 67.1599 | 62.6708
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 56.7563 | 2.5031 | 25.0064 | 999.858 | 0.0079897 | 0.4025 | 64.1754 | 59.3336
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 54.8125 | 2.0949 | 25.0102 | 999.83 | 0.0079875 | 0.402 | 61.176 | 55.9089
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 52.6688 | 1.7663 | 25.0195 | 999.783 | 0.0079884 | 0.4013 | 58.1735 | 52.5608
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 50.1861 | 1.5783 | 25.0186 | 1000.15 | 0.0079872 | 0.40113 | 55.1696 | 49.3518
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 48.1998 | 1.1887 | 25.0189 | 1000.14 | 0.0079862 | 0.40097 | 52.1639 | 45.9389
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 46.3418 | 0.75046 | 25.0273 | 999.8 | 0.0079882 | 0.40045 | 49.1685 | 42.4885
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 44.0884 | 0.47084 | 25.0311 | 999.827 | 0.0079863 | 0.39991 | 46.167 | 39.191
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 41.9905 | 0.12296 | 25.025 | 999.8 | 0.0079835 | 0.40009 | 43.1473 | 35.806
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 39.7228 | -0.15194 | 25.0208 | 999.95 | 0.0079822 | 0.40027 | 40.1395 | 32.5055
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 37.354 | -0.3717 | 25.026 | 999.869 | 0.0079805 | 0.39966 | 37.168 | 29.2973
Populus deltoides | LI2B |  | RACiR 500 to 0 no VPD control | 35.5539 | -0.81929 | 25.0307 | 1000.18 | 0.0079816 | 0.39933 | 34.1993 | 25.8637
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 325.4052 | 29.513 | 24.8291 | 1000.19 | 0.0074045 | 0.35935 | 415.164 | 430.618
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 322.2239 | 28.8546 | 24.8033 | 999.781 | 0.0074377 | 0.36363 | 409.053 | 423.821
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 317.8273 | 28.7417 | 24.7812 | 1000.06 | 0.0074982 | 0.36985 | 402.844 | 417.483
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 313.9075 | 28.4038 | 24.7688 | 1000.1 | 0.0075632 | 0.37579 | 396.641 | 410.929
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 310.264 | 27.9334 | 24.7521 | 1000.13 | 0.00762 | 0.38142 | 390.499 | 404.297
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 305.7496 | 27.6176 | 24.7414 | 1000.17 | 0.0076596 | 0.38534 | 384.278 | 397.733
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 301.2834 | 27.0785 | 24.7433 | 999.842 | 0.0076796 | 0.38687 | 378.02 | 390.904
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 295.9476 | 26.8132 | 24.7499 | 1000.14 | 0.0076939 | 0.3875 | 371.772 | 384.353
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 290.0551 | 26.7381 | 24.7543 | 1000.04 | 0.0076978 | 0.38749 | 365.571 | 378.036
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 284.8585 | 26.3777 | 24.7553 | 1000.12 | 0.0076927 | 0.38703 | 359.415 | 371.478
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 279.8875 | 25.9435 | 24.7595 | 1000.12 | 0.007689 | 0.38639 | 353.331 | 364.917
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 274.6056 | 25.595 | 24.7671 | 999.817 | 0.0076853 | 0.3855 | 347.198 | 358.393
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 269.568 | 25.1455 | 24.7725 | 1000.09 | 0.0076801 | 0.38465 | 341.037 | 351.74
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 264.1593 | 24.8439 | 24.7792 | 1000.12 | 0.0076763 | 0.38381 | 334.88 | 345.238
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 258.6013 | 24.6048 | 24.7877 | 1000.07 | 0.0076724 | 0.38289 | 328.746 | 338.822
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 254.1896 | 23.934 | 24.7944 | 999.822 | 0.0076664 | 0.38194 | 322.643 | 332.007
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 249.5301 | 23.372 | 24.7988 | 1000.18 | 0.0076641 | 0.38142 | 316.493 | 325.253
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 243.6493 | 23.2413 | 24.8055 | 999.786 | 0.0076625 | 0.38081 | 310.254 | 318.841
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 238.318 | 22.8671 | 24.8252 | 1000.08 | 0.0076618 | 0.37936 | 304.073 | 312.245
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 233.346 | 22.4141 | 24.8388 | 999.761 | 0.0076634 | 0.37848 | 297.946 | 305.624
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 228.2526 | 22.0703 | 24.8374 | 1000.04 | 0.0076647 | 0.37866 | 291.802 | 299.095
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 223.6953 | 21.4887 | 24.8405 | 1000.13 | 0.0076647 | 0.37844 | 285.635 | 292.306
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 218.3924 | 21.1686 | 24.8485 | 1000.1 | 0.0076662 | 0.37797 | 279.448 | 285.758
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 212.7372 | 21.0236 | 24.8528 | 1000.07 | 0.0076672 | 0.37774 | 273.326 | 279.45
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 207.9595 | 20.5662 | 24.8607 | 1000.09 | 0.007671 | 0.37749 | 267.27 | 272.898
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 203.26 | 20.0857 | 24.8693 | 1000.04 | 0.0076765 | 0.37725 | 261.228 | 266.336
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 198.367 | 19.6775 | 24.8751 | 1000.07 | 0.0076809 | 0.37718 | 255.155 | 259.815
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 193.7768 | 19.1478 | 24.8816 | 1000.04 | 0.0076814 | 0.37681 | 249.106 | 253.197
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 188.8749 | 18.726 | 24.8934 | 999.994 | 0.0076835 | 0.37616 | 243.069 | 246.698
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 183.6218 | 18.4378 | 24.9053 | 1000.02 | 0.0076884 | 0.37565 | 237.008 | 240.309
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 179.6612 | 17.7093 | 24.9074 | 1000.04 | 0.0076933 | 0.37589 | 230.977 | 233.511
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 175.4813 | 17.0422 | 24.9089 | 1000.05 | 0.0076965 | 0.37603 | 224.9 | 226.727
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 170.0823 | 16.7726 | 24.925 | 1000.07 | 0.0077015 | 0.37529 | 218.765 | 220.282
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 165.4113 | 16.2604 | 24.9354 | 1000.05 | 0.0077043 | 0.37478 | 212.684 | 213.649
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 160.7635 | 15.7631 | 24.944 | 999.998 | 0.007708 | 0.37441 | 206.647 | 207.075
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 155.7085 | 15.4065 | 24.9492 | 999.936 | 0.0077109 | 0.37427 | 200.543 | 200.574
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 151.5745 | 14.7254 | 24.9499 | 1000.04 | 0.0077151 | 0.37454 | 194.454 | 193.764
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 146.8211 | 14.2482 | 24.9654 | 999.983 | 0.0077192 | 0.3738 | 188.399 | 187.192
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 141.4549 | 14.0211 | 24.9733 | 1000.01 | 0.0077244 | 0.37364 | 182.327 | 180.853
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 136.8096 | 13.5604 | 24.9727 | 1000.02 | 0.0077256 | 0.37378 | 176.325 | 174.35
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 132.3677 | 13.0318 | 24.9749 | 999.995 | 0.0077295 | 0.37392 | 170.348 | 167.805
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 127.7811 | 12.5134 | 24.9848 | 999.973 | 0.0077336 | 0.37351 | 164.307 | 161.206
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 123.618 | 11.8558 | 24.987 | 999.975 | 0.0077365 | 0.37354 | 158.276 | 154.477
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 118.8292 | 11.4267 | 24.993 | 999.993 | 0.0077392 | 0.37333 | 152.246 | 147.978
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 113.6896 | 11.115 | 25.0076 | 1000 | 0.0077418 | 0.37252 | 146.228 | 141.608
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 109.9168 | 10.322 | 25.0138 | 1000.01 | 0.0077438 | 0.37224 | 140.248 | 134.795
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 105.8982 | 9.6038 | 25.019 | 999.993 | 0.0077474 | 0.37216 | 134.203 | 127.992
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 100.7785 | 9.2894 | 25.0157 | 999.918 | 0.0077459 | 0.37211 | 128.125 | 121.559
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 96.2497 | 8.7363 | 25.0175 | 999.985 | 0.0077277 | 0.37061 | 122.099 | 114.936
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 91.834 | 8.1347 | 25.0256 | 999.983 | 0.0077039 | 0.36828 | 116.097 | 108.29
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 87.0533 | 7.6787 | 25.028 | 999.998 | 0.0076845 | 0.36658 | 110.066 | 101.761
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 83.1198 | 6.9438 | 25.0203 | 999.988 | 0.0076769 | 0.36646 | 104.028 | 94.9463
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 78.5719 | 6.4459 | 25.0055 | 999.953 | 0.0076779 | 0.36746 | 97.9582 | 88.3386
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 73.5757 | 6.099 | 25.0068 | 999.948 | 0.0076786 | 0.36743 | 91.9039 | 81.8969
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 69.2416 | 5.5235 | 25.0089 | 1000.03 | 0.0076807 | 0.36742 | 85.8963 | 75.2729
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 65.0019 | 4.9344 | 24.9994 | 999.987 | 0.0076818 | 0.36805 | 79.9212 | 68.6699
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 60.7362 | 4.3474 | 24.9867 | 999.959 | 0.0076858 | 0.36916 | 73.9146 | 62.0375
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 56.8715 | 3.595 | 24.9865 | 999.959 | 0.0076883 | 0.36931 | 67.9046 | 55.2356
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 52.4496 | 3.043 | 24.99 | 999.98 | 0.0076879 | 0.36905 | 61.8858 | 48.6238
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 47.5901 | 2.6515 | 24.9865 | 999.96 | 0.0076878 | 0.36927 | 55.8472 | 42.1536
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 44.0176 | 1.7826 | 24.9927 | 1000 | 0.0076897 | 0.36895 | 49.8272 | 35.2243
Populus deltoides | LI2B |  | RACiR 500 to 0 2.5 min | 40.3061 | 0.95162 | 24.9935 | 999.971 | 0.0076922 | 0.36904 | 43.7609 | 28.2864
Populus deltoides | LI2B |  | RACiR 500 to 0 | 331.802 | 31.7285 | 24.9951 | 1000.07 | 0.0066118 | 0.28964 | 450.271 | 476.109
Populus deltoides | LI2B |  | RACiR 500 to 0 | 331.222 | 31.3723 | 24.9825 | 999.648 | 0.0066106 | 0.29019 | 448.208 | 473.686
Populus deltoides | LI2B |  | RACiR 500 to 0 | 329.8628 | 31.1736 | 24.982 | 1000.11 | 0.0066118 | 0.29027 | 446.095 | 471.359
Populus deltoides | LI2B |  | RACiR 500 to 0 | 329.3061 | 30.8996 | 24.9821 | 1000.05 | 0.0066114 | 0.29024 | 444.581 | 469.567
Populus deltoides | LI2B |  | RACiR 500 to 0 | 327.7578 | 30.6974 | 24.9886 | 1000.1 | 0.0066131 | 0.29001 | 442.386 | 467.155
Populus deltoides | LI2B |  | RACiR 500 to 0 | 326.3974 | 30.4609 | 24.9875 | 1000.01 | 0.0066115 | 0.29 | 440.173 | 464.701
Populus deltoides | LI2B |  | RACiR 500 to 0 | 325.1836 | 30.3455 | 24.9899 | 999.701 | 0.0066107 | 0.28988 | 438.574 | 462.983
Populus deltoides | LI2B |  | RACiR 500 to 0 | 323.5866 | 30.1908 | 24.9887 | 1000.09 | 0.0066131 | 0.29008 | 436.326 | 460.568
Populus deltoides | LI2B |  | RACiR 500 to 0 | 322.3988 | 29.8902 | 24.9895 | 999.984 | 0.0066136 | 0.2901 | 434.055 | 457.989
Populus deltoides | LI2B |  | RACiR 500 to 0 | 320.9019 | 29.868 | 24.9892 | 1000.01 | 0.0066146 | 0.29018 | 432.419 | 456.324
Populus deltoides | LI2B |  | RACiR 500 to 0 | 319.6919 | 29.5492 | 24.9923 | 999.718 | 0.006614 | 0.29003 | 430.126 | 453.705
Populus deltoides | LI2B |  | RACiR 500 to 0 | 317.903 | 29.5022 | 24.9893 | 1000.08 | 0.006631 | 0.29097 | 427.794 | 451.319
Populus deltoides | LI2B |  | RACiR 500 to 0 | 317.255 | 29.3102 | 24.9862 | 1000.09 | 0.0066483 | 0.29188 | 426.139 | 449.474
Populus deltoides | LI2B |  | RACiR 500 to 0 | 315.4183 | 29.2533 | 24.9832 | 1000.01 | 0.0066624 | 0.29255 | 423.83 | 447.102
Populus deltoides | LI2B |  | RACiR 500 to 0 | 313.1825 | 29.2123 | 24.9854 | 999.997 | 0.0066608 | 0.29224 | 421.517 | 444.738
Populus deltoides | LI2B |  | RACiR 500 to 0 | 311.5431 | 29.1385 | 24.9882 | 1000.03 | 0.0066474 | 0.29139 | 419.884 | 443.021
Populus deltoides | LI2B |  | RACiR 500 to 0 | 309.4053 | 29.0494 | 24.991 | 1000.03 | 0.0066397 | 0.29082 | 417.593 | 440.626
Populus deltoides | LI2B |  | RACiR 500 to 0 | 308.0805 | 28.7471 | 24.986 | 999.671 | 0.0066245 | 0.29034 | 415.344 | 438.061
Populus deltoides | LI2B |  | RACiR 500 to 0 | 306.2018 | 28.7696 | 24.9896 | 999.685 | 0.0066153 | 0.28979 | 413.697 | 436.424
Populus deltoides | LI2B |  | RACiR 500 to 0 | 304.5596 | 28.5433 | 24.9914 | 1000.06 | 0.0066019 | 0.28924 | 411.415 | 433.903
Populus deltoides | LI2B |  | RACiR 500 to 0 | 302.4297 | 28.4956 | 24.9969 | 1000.11 | 0.0066026 | 0.28917 | 409.094 | 431.52
Populus deltoides | LI2B |  | RACiR 500 to 0 | 301.9941 | 28.2024 | 24.9942 | 1000.1 | 0.0066073 | 0.28964 | 407.458 | 429.587
Populus deltoides | LI2B |  | RACiR 500 to 0 | 300.5076 | 28.0279 | 24.9911 | 1000.05 | 0.006613 | 0.29018 | 405.137 | 427.084
Populus deltoides | LI2B |  | RACiR 500 to 0 | 298.7581 | 27.8802 | 24.9926 | 1000.03 | 0.0066159 | 0.2903 | 402.793 | 424.585
Populus deltoides | LI2B |  | RACiR 500 to 0 | 297.251 | 27.8256 | 24.9971 | 999.927 | 0.0066171 | 0.29018 | 401.106 | 422.839
Populus deltoides | LI2B |  | RACiR 500 to 0 | 295.2436 | 27.7493 | 24.9984 | 999.885 | 0.0066188 | 0.29024 | 398.765 | 420.41
Populus deltoides | LI2B |  | RACiR 500 to 0 | 293.5496 | 27.5927 | 24.9973 | 1000.04 | 0.0066195 | 0.2904 | 396.431 | 417.911
Populus deltoides | LI2B |  | RACiR 500 to 0 | 291.7861 | 27.6569 | 24.9944 | 999.988 | 0.0066205 | 0.29062 | 394.77 | 416.305
Populus deltoides | LI2B |  | RACiR 500 to 0 | 289.8274 | 27.5679 | 24.9976 | 1000.02 | 0.006623 | 0.29065 | 392.449 | 413.887
Populus deltoides | LI2B |  | RACiR 500 to 0 | 287.3562 | 27.6265 | 25.0028 | 1000.07 | 0.0066254 | 0.29057 | 390.158 | 411.643
Populus deltoides | LI2B |  | RACiR 500 to 0 | 286.1642 | 27.5321 | 25.0027 | 1000.05 | 0.006627 | 0.29072 | 388.558 | 409.943
Populus deltoides | LI2B |  | RACiR 500 to 0 | 284.8825 | 27.2923 | 25.0006 | 1000.02 | 0.0066289 | 0.29095 | 386.332 | 407.468
Populus deltoides | LI2B |  | RACiR 500 to 0 | 283.9475 | 26.9479 | 24.9993 | 999.968 | 0.006633 | 0.29125 | 384.078 | 404.861
Populus deltoides | LI2B |  | RACiR 500 to 0 | 283.3352 | 26.6623 | 24.9984 | 999.993 | 0.0066335 | 0.29129 | 382.449 | 402.943
Populus deltoides | LI2B |  | RACiR 500 to 0 | 280.8809 | 26.7068 | 25.0005 | 999.891 | 0.0066366 | 0.2913 | 380.091 | 400.618
Populus deltoides | LI2B |  | RACiR 500 to 0 | 278.8079 | 26.6322 | 25.0011 | 1000.04 | 0.0066338 | 0.2911 | 377.779 | 398.224
Populus deltoides | LI2B |  | RACiR 500 to 0 | 276.9702 | 26.6923 | 25.0027 | 1000.03 | 0.0066336 | 0.29102 | 376.133 | 396.63
Populus deltoides | LI2B |  | RACiR 500 to 0 | 275.5783 | 26.4691 | 24.9991 | 999.999 | 0.0066335 | 0.29119 | 373.874 | 394.139
Populus deltoides | LI2B |  | RACiR 500 to 0 | 273.8541 | 26.3513 | 24.9968 | 999.936 | 0.0066367 | 0.2915 | 371.6 | 391.735
Populus deltoides | LI2B |  | RACiR 500 to 0 | 273.1265 | 26.0963 | 24.9987 | 1000.04 | 0.0066377 | 0.29149 | 369.976 | 389.852
Populus deltoides | LI2B |  | RACiR 500 to 0 | 271.2677 | 25.9882 | 24.9983 | 1000.01 | 0.0066395 | 0.29166 | 367.645 | 387.403
Populus deltoides | LI2B |  | RACiR 500 to 0 | 269.3926 | 25.8686 | 25.0013 | 1000.04 | 0.0066407 | 0.29161 | 365.33 | 384.959
Populus deltoides | LI2B |  | RACiR 500 to 0 | 268.0154 | 25.8214 | 24.9979 | 1000.01 | 0.0066409 | 0.2918 | 363.695 | 383.271
Populus deltoides | LI2B |  | RACiR 500 to 0 | 266.0619 | 25.7465 | 25.0003 | 999.97 | 0.0066428 | 0.29179 | 361.441 | 380.93
Populus deltoides | LI2B |  | RACiR 500 to 0 | 264.601 | 25.5341 | 25.0008 | 999.996 | 0.0066463 | 0.29192 | 359.17 | 378.441
Populus deltoides | LI2B |  | RACiR 500 to 0 | 263.1897 | 25.5594 | 24.9997 | 1000.02 | 0.006665 | 0.29284 | 357.535 | 376.829
Populus deltoides | LI2B |  | RACiR 500 to 0 | 261.7794 | 25.3946 | 25.003 | 999.994 | 0.0066894 | 0.29374 | 355.259 | 374.39
Populus deltoides | LI2B |  | RACiR 500 to 0 | 260.1506 | 25.3072 | 24.9988 | 1000.02 | 0.0067079 | 0.29461 | 353.031 | 372.07
Populus deltoides | LI2B |  | RACiR 500 to 0 | 259.352 | 25.0748 | 24.9947 | 999.991 | 0.0067054 | 0.29454 | 351.436 | 370.235
Populus deltoides | LI2B |  | RACiR 500 to 0 | 257.6262 | 24.9125 | 24.9877 | 1000.04 | 0.0066982 | 0.29434 | 349.17 | 367.793
Populus deltoides | LI2B |  | RACiR 500 to 0 | 255.7947 | 24.7563 | 24.9873 | 1000.04 | 0.0066927 | 0.29396 | 346.871 | 365.326
Populus deltoides | LI2B |  | RACiR 500 to 0 | 254.4307 | 24.6566 | 24.9916 | 1000.07 | 0.0066901 | 0.29361 | 345.241 | 363.59
Populus deltoides | LI2B |  | RACiR 500 to 0 | 252.7149 | 24.4979 | 24.9917 | 999.995 | 0.0066888 | 0.2935 | 342.971 | 361.149
Populus deltoides | LI2B |  | RACiR 500 to 0 | 251.2898 | 24.227 | 24.9963 | 1000.01 | 0.0066862 | 0.29316 | 340.682 | 358.58
Populus deltoides | LI2B |  | RACiR 500 to 0 | 249.6945 | 24.2106 | 24.9969 | 1000.02 | 0.006685 | 0.29309 | 339.014 | 356.887
Populus deltoides | LI2B |  | RACiR 500 to 0 | 247.9748 | 24.0407 | 24.999 | 1000.03 | 0.0066819 | 0.29288 | 336.729 | 354.424
Populus deltoides | LI2B |  | RACiR 500 to 0 | 245.7757 | 24.03 | 25.0006 | 999.92 | 0.0066831 | 0.29285 | 334.451 | 352.121
Populus deltoides | LI2B |  | RACiR 500 to 0 | 244.5974 | 23.9025 | 25.0016 | 999.962 | 0.0066825 | 0.29274 | 332.84 | 350.376
Populus deltoides | LI2B |  | RACiR 500 to 0 | 242.8253 | 23.7746 | 25.0025 | 1000.03 | 0.0066844 | 0.29276 | 330.58 | 347.979
Populus deltoides | LI2B |  | RACiR 500 to 0 | 241.1625 | 23.6377 | 24.9984 | 1000.02 | 0.0066859 | 0.293 | 328.336 | 345.589
Populus deltoides | LI2B |  | RACiR 500 to 0 | 240.1498 | 23.4977 | 24.9924 | 1000.09 | 0.006685 | 0.29325 | 326.739 | 343.848
Populus deltoides | LI2B |  | RACiR 500 to 0 | 238.5224 | 23.3241 | 24.9947 | 999.992 | 0.0066861 | 0.29321 | 324.489 | 341.412
Populus deltoides | LI2B |  | RACiR 500 to 0 | 237.4853 | 22.9678 | 24.996 | 999.937 | 0.0066842 | 0.29304 | 322.248 | 338.807
Populus deltoides | LI2B |  | RACiR 500 to 0 | 235.9523 | 22.9178 | 25.0051 | 999.953 | 0.0066858 | 0.29266 | 320.618 | 337.118
Populus deltoides | LI2B |  | RACiR 500 to 0 | 234.3403 | 22.7236 | 25.0046 | 1000.06 | 0.0066845 | 0.2926 | 318.315 | 334.613
Populus deltoides | LI2B |  | RACiR 500 to 0 | 232.2038 | 22.6971 | 25.0048 | 999.981 | 0.0066858 | 0.29264 | 316.026 | 332.285
Populus deltoides | LI2B |  | RACiR 500 to 0 | 231.268 | 22.5123 | 25.002 | 1000 | 0.0066861 | 0.29281 | 314.382 | 330.452
Populus deltoides | LI2B |  | RACiR 500 to 0 | 229.5277 | 22.3742 | 25.0023 | 999.941 | 0.0066879 | 0.29291 | 312.096 | 328.018
Populus deltoides | LI2B |  | RACiR 500 to 0 | 228.1789 | 22.1311 | 25.0019 | 999.934 | 0.0066902 | 0.29312 | 309.822 | 325.492
Populus deltoides | LI2B |  | RACiR 500 to 0 | 227.0298 | 21.9945 | 25.0021 | 1000.04 | 0.00669 | 0.29312 | 308.172 | 323.7
Populus deltoides | LI2B |  | RACiR 500 to 0 | 225.3983 | 21.824 | 25.0007 | 999.968 | 0.0066923 | 0.29332 | 305.863 | 321.209
Populus deltoides | LI2B |  | RACiR 500 to 0 | 223.9356 | 21.606 | 24.9997 | 999.972 | 0.0066946 | 0.29351 | 303.563 | 318.682
Populus deltoides | LI2B |  | RACiR 500 to 0 | 222.0005 | 21.6854 | 25.0025 | 1000 | 0.006695 | 0.29339 | 301.892 | 317.082
Populus deltoides | LI2B |  | RACiR 500 to 0 | 220.3219 | 21.5176 | 25.0039 | 999.987 | 0.0066941 | 0.29329 | 299.623 | 314.637
Populus deltoides | LI2B |  | RACiR 500 to 0 | 218.3509 | 21.445 | 25.004 | 1000.01 | 0.0066953 | 0.29334 | 297.342 | 312.271
Populus deltoides | LI2B |  | RACiR 500 to 0 | 217.8702 | 21.1507 | 24.9971 | 999.985 | 0.0066956 | 0.29372 | 295.733 | 310.363
Populus deltoides | LI2B |  | RACiR 500 to 0 | 216.2974 | 20.9766 | 24.994 | 999.954 | 0.0066987 | 0.29404 | 293.443 | 307.89
Populus deltoides | LI2B |  | RACiR 500 to 0 | 214.7831 | 20.7898 | 24.9898 | 999.984 | 0.0067018 | 0.29442 | 291.155 | 305.407
Populus deltoides | LI2B |  | RACiR 500 to 0 | 213.3977 | 20.7152 | 24.9948 | 999.941 | 0.006703 | 0.29424 | 289.528 | 303.699
Populus deltoides | LI2B |  | RACiR 500 to 0 | 211.3278 | 20.6454 | 25.0048 | 999.952 | 0.0067053 | 0.29395 | 287.249 | 301.339
Populus deltoides | LI2B |  | RACiR 500 to 0 | 209.9599 | 20.4032 | 25.0164 | 1000.02 | 0.006719 | 0.29407 | 284.997 | 298.84
Populus deltoides | LI2B |  | RACiR 500 to 0 | 208.6136 | 20.4061 | 25.015 | 999.951 | 0.0067405 | 0.29513 | 283.375 | 297.219
Populus deltoides | LI2B |  | RACiR 500 to 0 | 207.4508 | 20.1982 | 25.0061 | 1000.01 | 0.0067663 | 0.29661 | 281.12 | 294.755
Populus deltoides | LI2B |  | RACiR 500 to 0 | 205.5775 | 20.1409 | 25.0067 | 999.97 | 0.0067823 | 0.29713 | 278.891 | 292.461
Populus deltoides | LI2B |  | RACiR 500 to 0 | 204.5633 | 19.9498 | 25.0073 | 999.963 | 0.0067783 | 0.29675 | 277.294 | 290.666
Populus deltoides | LI2B |  | RACiR 500 to 0 | 202.8645 | 19.78 | 25.0077 | 1000.27 | 0.0067743 | 0.29641 | 275.058 | 288.248
Populus deltoides | LI2B |  | RACiR 500 to 0 | 201.384 | 19.5447 | 25.0051 | 999.934 | 0.0067675 | 0.2961 | 272.809 | 285.754
Populus deltoides | LI2B |  | RACiR 500 to 0 | 200.2004 | 19.4109 | 25.0077 | 999.918 | 0.0067641 | 0.29579 | 271.214 | 284.018
Populus deltoides | LI2B |  | RACiR 500 to 0 | 198.4683 | 19.2592 | 25.0097 | 999.983 | 0.0067628 | 0.29559 | 268.97 | 281.611
Populus deltoides | LI2B |  | RACiR 500 to 0 | 197.0167 | 19.0423 | 25.0067 | 999.98 | 0.0067609 | 0.29563 | 266.731 | 279.146
Populus deltoides | LI2B |  | RACiR 500 to 0 | 195.3781 | 19.0468 | 25.0059 | 999.902 | 0.0067564 | 0.29547 | 265.105 | 277.515
Populus deltoides | LI2B |  | RACiR 500 to 0 | 193.6884 | 18.8613 | 25.007 | 999.918 | 0.0067451 | 0.29494 | 262.859 | 275.071
Populus deltoides | LI2B |  | RACiR 500 to 0 | 191.6691 | 18.7836 | 25.0113 | 999.967 | 0.0067386 | 0.29452 | 260.62 | 272.74
Populus deltoides | LI2B |  | RACiR 500 to 0 | 190.9783 | 18.5258 | 25.0138 | 1000.02 | 0.0067371 | 0.29443 | 259.049 | 270.906
Populus deltoides | LI2B |  | RACiR 500 to 0 | 189.52 | 18.3448 | 25.0117 | 1000.25 | 0.0067431 | 0.29492 | 256.826 | 268.494
Populus deltoides | LI2B |  | RACiR 500 to 0 | 188.1576 | 18.115 | 25.0113 | 999.945 | 0.0067466 | 0.29519 | 254.584 | 266.014
Populus deltoides | LI2B |  | RACiR 500 to 0 | 186.9778 | 18.0081 | 25.0096 | 999.959 | 0.0067487 | 0.29541 | 252.962 | 264.28
Populus deltoides | LI2B |  | RACiR 500 to 0 | 185.3264 | 17.8432 | 25.0131 | 1000.1 | 0.0067517 | 0.29545 | 250.702 | 261.844
Populus deltoides | LI2B |  | RACiR 500 to 0 | 184.0639 | 17.566 | 25.0114 | 999.922 | 0.0067517 | 0.29555 | 248.438 | 259.294
Populus deltoides | LI2B |  | RACiR 500 to 0 | 182.5963 | 17.5339 | 25.0104 | 999.907 | 0.0067523 | 0.29565 | 246.806 | 257.622
Populus deltoides | LI2B |  | RACiR 500 to 0 | 180.9679 | 17.3623 | 25.0105 | 999.942 | 0.0067535 | 0.29572 | 244.539 | 255.175
Populus deltoides | LI2B |  | RACiR 500 to 0 | 178.9322 | 17.3365 | 25.0067 | 999.896 | 0.0067567 | 0.29609 | 242.29 | 252.889
Populus deltoides | LI2B |  | RACiR 500 to 0 | 177.9486 | 17.1658 | 25.009 | 999.932 | 0.0067588 | 0.29613 | 240.693 | 251.116
Populus deltoides | LI2B |  | RACiR 500 to 0 | 176.5579 | 16.95 | 25.0064 | 1000.3 | 0.0067616 | 0.29642 | 238.473 | 248.671
Populus deltoides | LI2B |  | RACiR 500 to 0 | 175.0247 | 16.7622 | 25.0079 | 999.923 | 0.0067643 | 0.29655 | 236.239 | 246.241
Populus deltoides | LI2B |  | RACiR 500 to 0 | 173.8875 | 16.6465 | 25.0052 | 999.967 | 0.0067665 | 0.29681 | 234.628 | 244.508
Populus deltoides | LI2B |  | RACiR 500 to 0 | 172.2065 | 16.4884 | 25.0089 | 1000.11 | 0.0067689 | 0.29679 | 232.375 | 242.086
Populus deltoides | LI2B |  | RACiR 500 to 0 | 171.0177 | 16.1812 | 25.0123 | 999.906 | 0.00677 | 0.29673 | 230.126 | 239.522
Populus deltoides | LI2B |  | RACiR 500 to 0 | 169.6835 | 16.1022 | 25.0149 | 999.935 | 0.0067705 | 0.29664 | 228.509 | 237.818
Populus deltoides | LI2B |  | RACiR 500 to 0 | 168.3359 | 15.8464 | 25.0161 | 999.858 | 0.0067724 | 0.29671 | 226.247 | 235.291
Populus deltoides | LI2B |  | RACiR 500 to 0 | 166.3835 | 15.7702 | 25.0158 | 1000.18 | 0.0067751 | 0.29687 | 223.959 | 232.916
Populus deltoides | LI2B |  | RACiR 500 to 0 | 165.6395 | 15.5209 | 25.017 | 999.894 | 0.0067754 | 0.29685 | 222.353 | 231.055
Populus deltoides | LI2B |  | RACiR 500 to 0 | 164.1755 | 15.2976 | 25.0174 | 999.85 | 0.0067769 | 0.29694 | 220.078 | 228.547
Populus deltoides | LI2B |  | RACiR 500 to 0 | 162.8061 | 15.0503 | 25.0169 | 1000.17 | 0.0067809 | 0.2972 | 217.789 | 226.002
Populus deltoides | LI2B |  | RACiR 500 to 0 | 161.6514 | 14.9239 | 25.0138 | 999.88 | 0.0067811 | 0.29735 | 216.148 | 224.228
Populus deltoides | LI2B |  | RACiR 500 to 0 | 159.8657 | 14.7996 | 25.0132 | 999.938 | 0.0067875 | 0.29771 | 213.837 | 221.783
Populus deltoides | LI2B |  | RACiR 500 to 0 | 158.448 | 14.6111 | 25.0091 | 999.928 | 0.0068108 | 0.29897 | 211.534 | 219.287
Populus deltoides | LI2B |  | RACiR 500 to 0 | 156.7566 | 14.6746 | 25.0115 | 999.988 | 0.0068321 | 0.29978 | 209.885 | 217.697
Populus deltoides | LI2B |  | RACiR 500 to 0 | 155.2641 | 14.5159 | 25.0072 | 999.88 | 0.0068548 | 0.30087 | 207.644 | 215.293
Populus deltoides | LI2B |  | RACiR 500 to 0 | 153.3435 | 14.4119 | 25.0141 | 999.908 | 0.0068547 | 0.30031 | 205.428 | 212.962
Populus deltoides | LI2B |  | RACiR 500 to 0 | 152.6781 | 14.1273 | 25.0124 | 999.911 | 0.0068478 | 0.29991 | 203.854 | 211.096
Populus deltoides | LI2B |  | RACiR 500 to 0 | 151.003 | 13.9341 | 25.0148 | 1000.16 | 0.0068389 | 0.29921 | 201.602 | 208.639
Populus deltoides | LI2B |  | RACiR 500 to 0 | 149.5077 | 13.7156 | 25.0106 | 999.932 | 0.0068347 | 0.29914 | 199.344 | 206.151
Populus deltoides | LI2B |  | RACiR 500 to 0 | 148.3108 | 13.5983 | 25.0089 | 1000.19 | 0.0068305 | 0.29896 | 197.749 | 204.432
Populus deltoides | LI2B |  | RACiR 500 to 0 | 146.4736 | 13.4944 | 25.0045 | 999.861 | 0.0068287 | 0.29903 | 195.506 | 202.074
Populus deltoides | LI2B |  | RACiR 500 to 0 | 145.0675 | 13.2626 | 25.0014 | 999.913 | 0.0068272 | 0.29906 | 193.277 | 199.604
Populus deltoides | LI2B |  | RACiR 500 to 0 | 143.557 | 13.2559 | 24.9973 | 999.879 | 0.0068285 | 0.29931 | 191.669 | 197.981
Populus deltoides | LI2B |  | RACiR 500 to 0 | 142.2572 | 12.9965 | 24.9985 | 999.886 | 0.0068297 | 0.29932 | 189.461 | 195.505
Populus deltoides | LI2B |  | RACiR 500 to 0 | 140.3946 | 12.8982 | 24.9999 | 1000.14 | 0.0068299 | 0.29929 | 187.229 | 193.164
Populus deltoides | LI2B |  | RACiR 500 to 0 | 139.7218 | 12.6217 | 25.0035 | 999.89 | 0.0068274 | 0.29901 | 185.648 | 191.3
Populus deltoides | LI2B |  | RACiR 500 to 0 | 138.3239 | 12.3849 | 25.0017 | 999.905 | 0.0068269 | 0.29909 | 183.405 | 188.81
Populus deltoides | LI2B |  | RACiR 500 to 0 | 136.8832 | 12.1564 | 25.0049 | 1000.22 | 0.0068299 | 0.29911 | 181.156 | 186.323
Populus deltoides | LI2B |  | RACiR 500 to 0 | 135.4934 | 12.0872 | 25.0116 | 999.915 | 0.0068302 | 0.2988 | 179.547 | 184.638
Populus deltoides | LI2B |  | RACiR 500 to 0 | 133.4033 | 12.0671 | 25.0084 | 1000.03 | 0.0068338 | 0.29916 | 177.288 | 182.349
Populus deltoides | LI2B |  | RACiR 500 to 0 | 131.9173 | 11.8784 | 25.0023 | 999.906 | 0.006835 | 0.29954 | 175.076 | 179.939
Populus deltoides | LI2B |  | RACiR 500 to 0 | 130.5328 | 11.8628 | 24.9899 | 999.923 | 0.0068384 | 0.30033 | 173.493 | 178.333
Populus deltoides | LI2B |  | RACiR 500 to 0 | 129.3134 | 11.5776 | 24.9953 | 999.906 | 0.0068379 | 0.30006 | 171.323 | 175.869
Populus deltoides | LI2B |  | RACiR 500 to 0 | 127.5614 | 11.4358 | 25.0028 | 1000.22 | 0.0068391 | 0.29976 | 169.095 | 173.489
Populus deltoides | LI2B |  | RACiR 500 to 0 | 126.8896 | 11.163 | 25.0093 | 1000.13 | 0.0068393 | 0.29946 | 167.529 | 171.644
Populus deltoides | LI2B |  | RACiR 500 to 0 | 125.2886 | 10.9811 | 25.0122 | 999.833 | 0.0068401 | 0.29938 | 165.288 | 169.211
Populus deltoides | LI2B |  | RACiR 500 to 0 | 123.9069 | 10.7397 | 25.0126 | 999.934 | 0.0068409 | 0.29942 | 163.053 | 166.725
Populus deltoides | LI2B |  | RACiR 500 to 0 | 122.7218 | 10.6175 | 25.0144 | 999.854 | 0.0068404 | 0.29932 | 161.441 | 164.984
Populus deltoides | LI2B |  | RACiR 500 to 0 | 121.0851 | 10.4484 | 25.0195 | 1000.06 | 0.0068451 | 0.2993 | 159.199 | 162.563
Populus deltoides | LI2B |  | RACiR 500 to 0 | 119.948 | 10.1393 | 25.0177 | 1000.17 | 0.0068476 | 0.29951 | 156.966 | 160.012
Populus deltoides | LI2B |  | RACiR 500 to 0 | 118.5552 | 10.0943 | 25.0155 | 1000.18 | 0.006851 | 0.29982 | 155.352 | 158.346
Populus deltoides | LI2B |  | RACiR 500 to 0 | 117.5336 | 9.7532 | 25.0108 | 1000.17 | 0.0068517 | 0.30013 | 153.119 | 155.763
Populus deltoides | LI2B |  | RACiR 500 to 0 | 115.6828 | 9.6438 | 25.0112 | 999.841 | 0.0068541 | 0.30026 | 150.843 | 153.367
Populus deltoides | LI2B |  | RACiR 500 to 0 | 114.9743 | 9.3725 | 25.015 | 1000.23 | 0.0068529 | 0.30001 | 149.233 | 151.479
Populus deltoides | LI2B |  | RACiR 500 to 0 | 113.2114 | 9.2321 | 25.0136 | 1000.13 | 0.0068529 | 0.30007 | 146.948 | 149.043
Populus deltoides | LI2B |  | RACiR 500 to 0 | 111.4539 | 9.0873 | 25.0159 | 1000.16 | 0.0068518 | 0.29991 | 144.68 | 146.62
Populus deltoides | LI2B |  | RACiR 500 to 0 | 110.1229 | 9.0213 | 25.0137 | 999.867 | 0.0068531 | 0.30012 | 143.073 | 144.94
Populus deltoides | LI2B |  | RACiR 500 to 0 | 108.4264 | 8.8793 | 25.0124 | 1000.04 | 0.0068544 | 0.30025 | 140.844 | 142.559
Populus deltoides | LI2B |  | RACiR 500 to 0 | 107.2357 | 8.5971 | 25.0126 | 1000.17 | 0.0068677 | 0.30089 | 138.614 | 140.039
Populus deltoides | LI2B |  | RACiR 500 to 0 | 106.0215 | 8.5183 | 25.0152 | 1000.22 | 0.0068894 | 0.30178 | 137.023 | 138.365
Populus deltoides | LI2B |  | RACiR 500 to 0 | 104.7744 | 8.2672 | 25.0166 | 1000.17 | 0.0069181 | 0.30295 | 134.799 | 135.884
Populus deltoides | LI2B |  | RACiR 500 to 0 | 103.0556 | 8.1424 | 25.0149 | 999.891 | 0.006934 | 0.30356 | 132.57 | 133.522
Populus deltoides | LI2B |  | RACiR 500 to 0 | 102.1714 | 7.9291 | 25.0134 | 1000.19 | 0.0069286 | 0.30321 | 130.986 | 131.717
Populus deltoides | LI2B |  | RACiR 500 to 0 | 100.4318 | 7.763 | 25.0173 | 1000.22 | 0.0069151 | 0.3022 | 128.742 | 129.295
Populus deltoides | LI2B |  | RACiR 500 to 0 | 98.5398 | 7.6428 | 25.0131 | 999.918 | 0.0068873 | 0.30105 | 126.501 | 126.92
Populus deltoides | LI2B |  | RACiR 500 to 0 | 97.3763 | 7.5041 | 25.0133 | 1000.19 | 0.0068688 | 0.30024 | 124.913 | 125.184
Populus deltoides | LI2B |  | RACiR 500 to 0 | 95.8818 | 7.2901 | 25.0116 | 999.985 | 0.006851 | 0.29962 | 122.712 | 122.757
Populus deltoides | LI2B |  | RACiR 500 to 0 | 94.6633 | 7.0083 | 25.0157 | 1000.14 | 0.006854 | 0.29973 | 120.506 | 120.26
Populus deltoides | LI2B |  | RACiR 500 to 0 | 93.2311 | 6.981 | 25.0124 | 1000.22 | 0.006859 | 0.30019 | 118.91 | 118.63
Populus deltoides | LI2B |  | RACiR 500 to 0 | 91.9119 | 6.7363 | 25.0085 | 999.894 | 0.0068634 | 0.3007 | 116.695 | 116.161
Populus deltoides | LI2B |  | RACiR 500 to 0 | 90.3896 | 6.5515 | 25.0068 | 1000.27 | 0.0068683 | 0.30109 | 114.486 | 113.758
Populus deltoides | LI2B |  | RACiR 500 to 0 | 89.9007 | 6.2289 | 25.0105 | 1000.21 | 0.0068715 | 0.30111 | 112.901 | 111.844
Populus deltoides | LI2B |  | RACiR 500 to 0 | 88.438 | 6.0045 | 25.012 | 1000.17 | 0.0068747 | 0.30126 | 110.641 | 109.35
Populus deltoides | LI2B |  | RACiR 500 to 0 | 86.9716 | 5.7735 | 25.0121 | 1000.21 | 0.0068763 | 0.30136 | 108.358 | 106.826
Populus deltoides | LI2B |  | RACiR 500 to 0 | 85.8205 | 5.6457 | 25.0114 | 999.849 | 0.0068769 | 0.30146 | 106.745 | 105.078
Populus deltoides | LI2B |  | RACiR 500 to 0 | 84.0976 | 5.4948 | 25.0136 | 1000.02 | 0.0068771 | 0.30137 | 104.482 | 102.654
Populus deltoides | LI2B |  | RACiR 500 to 0 | 82.7083 | 5.2438 | 25.0159 | 1000.17 | 0.0068785 | 0.30138 | 102.218 | 100.129
Populus deltoides | LI2B |  | RACiR 500 to 0 | 81.1988 | 5.2227 | 25.017 | 1000.14 | 0.0068812 | 0.30151 | 100.595 | 98.478
Populus deltoides | LI2B |  | RACiR 500 to 0 | 79.6624 | 5.0256 | 25.0123 | 999.857 | 0.0068834 | 0.30191 | 98.3378 | 96.0135
Populus deltoides | LI2B |  | RACiR 500 to 0 | 77.8331 | 4.9239 | 25.01 | 1000.19 | 0.0068884 | 0.30231 | 96.1024 | 93.6672
Populus deltoides | LI2B |  | RACiR 500 to 0 | 77.0446 | 4.6934 | 25.0076 | 1000.18 | 0.0068872 | 0.30236 | 94.5213 | 91.8483
Populus deltoides | LI2B |  | RACiR 500 to 0 | 75.4357 | 4.5185 | 25.0104 | 1000.19 | 0.0068896 | 0.30236 | 92.2909 | 89.4334
Populus deltoides | LI2B |  | RACiR 500 to 0 | 74.0011 | 4.2948 | 25.0065 | 1000.21 | 0.0068901 | 0.30259 | 90.0624 | 86.9713
Populus deltoides | LI2B |  | RACiR 500 to 0 | 72.8704 | 4.1647 | 25.0072 | 999.839 | 0.0068906 | 0.30259 | 88.4708 | 85.2422
Populus deltoides | LI2B |  | RACiR 500 to 0 | 71.2185 | 4.0029 | 25.0079 | 1000.04 | 0.0068928 | 0.30269 | 86.2356 | 82.8356
Populus deltoides | LI2B |  | RACiR 500 to 0 | 69.9118 | 3.7354 | 25.0136 | 1000.19 | 0.0068921 | 0.30238 | 84.0192 | 80.3413
Populus deltoides | LI2B |  | RACiR 500 to 0 | 68.43 | 3.7089 | 25.0152 | 1000.2 | 0.0068927 | 0.30235 | 82.4165 | 78.7051
Populus deltoides | LI2B |  | RACiR 500 to 0 | 66.9263 | 3.5102 | 25.0099 | 1000.17 | 0.0068926 | 0.30262 | 80.2012 | 76.2807
Populus deltoides | LI2B |  | RACiR 500 to 0 | 65.305 | 3.3513 | 25.0074 | 999.864 | 0.0068939 | 0.30277 | 78.0054 | 73.9173
Populus deltoides | LI2B |  | RACiR 500 to 0 | 64.5685 | 3.1112 | 25.0048 | 999.844 | 0.0068972 | 0.30303 | 76.44 | 72.1049
Populus deltoides | LI2B |  | RACiR 500 to 0 | 63.2471 | 2.8539 | 25.0078 | 1000.22 | 0.0069093 | 0.30344 | 74.2173 | 69.616
Populus deltoides | LI2B |  | RACiR 500 to 0 | 61.9477 | 2.5803 | 25.0111 | 1000.22 | 0.0069193 | 0.3037 | 71.9697 | 67.0852
Populus deltoides | LI2B |  | RACiR 500 to 0 | 60.8224 | 2.4422 | 25.0108 | 999.845 | 0.0069229 | 0.30381 | 70.3555 | 65.3256
Populus deltoides | LI2B |  | RACiR 500 to 0 | 59.2883 | 2.2439 | 25.013 | 1000.04 | 0.0069313 | 0.30395 | 68.1235 | 62.8864
Populus deltoides | LI2B |  | RACiR 500 to 0 | 58.2532 | 1.8974 | 25.0071 | 1000.2 | 0.0069414 | 0.30458 | 65.8973 | 60.3041
Populus deltoides | LI2B |  | RACiR 500 to 0 | 56.87 | 1.8394 | 25.0021 | 1000.2 | 0.0069499 | 0.30511 | 64.2784 | 58.6209
Populus deltoides | LI2B |  | RACiR 500 to 0 | 55.733 | 1.5112 | 25.001 | 1000.23 | 0.0069497 | 0.30504 | 62.0285 | 56.0323
Populus deltoides | LI2B |  | RACiR 500 to 0 | 53.886 | 1.3939 | 25.002 | 999.896 | 0.0069454 | 0.30464 | 59.7571 | 53.6338
Populus deltoides | LI2B |  | RACiR 500 to 0 | 53.092 | 1.1516 | 25.0072 | 999.855 | 0.0069426 | 0.3042 | 58.1471 | 51.7735
Populus deltoides | LI2B |  | RACiR 500 to 0 | 51.4117 | 0.98903 | 25.0024 | 1000.2 | 0.0069402 | 0.30425 | 55.8862 | 49.3401
Populus deltoides | LI2B |  | RACiR 500 to 0 | 49.6149 | 0.86327 | 25.001 | 1000.22 | 0.0069368 | 0.30409 | 53.6313 | 46.9487
Populus deltoides | LI2B |  | RACiR 500 to 0 | 48.4741 | 0.7364 | 24.9965 | 999.886 | 0.0069332 | 0.30411 | 52.0414 | 45.2237
Populus deltoides | LI2B |  | RACiR 500 to 0 | 47.0028 | 0.5212 | 24.9952 | 1000.05 | 0.0069312 | 0.30401 | 49.8209 | 42.7789
Populus deltoides | LI2B |  | RACiR 500 to 0 | 45.9358 | 0.17904 | 24.9933 | 1000.15 | 0.00693 | 0.30405 | 47.5905 | 40.1955
Populus deltoides | LI2B |  | RACiR 500 to 0 | 44.6531 | 0.092132 | 24.9994 | 1000.15 | 0.0069317 | 0.30383 | 45.9903 | 38.5025
Populus deltoides | LI2B |  | RACiR 500 to 0 | 43.2662 | -0.15742 | 24.9971 | 1000.19 | 0.0069347 | 0.30414 | 43.7405 | 35.9923
Populus deltoides | LI2B |  | RACiR 500 to 0 | 41.4434 | -0.27354 | 24.9951 | 999.864 | 0.0069385 | 0.30447 | 41.4899 | 33.6163
Populus deltoides | LI2B |  | RACiR 500 to 0 | 40.6263 | -0.50414 | 24.988 | 999.854 | 0.0069397 | 0.30491 | 39.8892 | 31.7769
Populus deltoides | LI2B |  | RACiR 500 to 0 | 39.0303 | -0.68953 | 24.9886 | 1000.17 | 0.0069433 | 0.30509 | 37.6427 | 29.3354
Populus deltoides | LI2B |  | RACiR 500 to 0 | 37.6664 | -0.94565 | 24.9888 | 1000.2 | 0.0069446 | 0.30519 | 35.3991 | 26.8258
Populus deltoides | LI2B |  | RACiR 300 to 800 | 262.7196 | 23.6848 | 24.9675 | 1000.27 | 0.0081026 | 0.35878 | 335.49 | 370.818
Populus deltoides | LI2B |  | RACiR 300 to 800 | 263.8639 | 23.7888 | 24.967 | 999.879 | 0.0081007 | 0.3587 | 336.967 | 372.407
Populus deltoides | LI2B |  | RACiR 300 to 800 | 265.4664 | 23.7828 | 24.9698 | 999.794 | 0.0081 | 0.35851 | 338.628 | 374.068
Populus deltoides | LI2B |  | RACiR 300 to 800 | 266.5844 | 23.915 | 24.9679 | 999.904 | 0.0080999 | 0.35863 | 340.12 | 375.699
Populus deltoides | LI2B |  | RACiR 300 to 800 | 267.8754 | 24.0263 | 24.9673 | 999.976 | 0.0081001 | 0.35869 | 341.743 | 377.442
Populus deltoides | LI2B |  | RACiR 300 to 800 | 269.2777 | 24.028 | 24.9718 | 999.901 | 0.0081013 | 0.35849 | 343.224 | 378.93
Populus deltoides | LI2B |  | RACiR 300 to 800 | 270.2803 | 24.2458 | 24.9709 | 1000.22 | 0.0081014 | 0.35855 | 344.852 | 380.783
Populus deltoides | LI2B |  | RACiR 300 to 800 | 271.3194 | 24.3975 | 24.9693 | 999.939 | 0.0081007 | 0.35863 | 346.325 | 382.418
Populus deltoides | LI2B |  | RACiR 300 to 800 | 272.8468 | 24.4716 | 24.9607 | 999.919 | 0.0081021 | 0.35923 | 347.971 | 384.14
Populus deltoides | LI2B |  | RACiR 300 to 800 | 273.8261 | 24.6504 | 24.9574 | 1000.25 | 0.0081011 | 0.35937 | 349.445 | 385.799
Populus deltoides | LI2B |  | RACiR 300 to 800 | 275.1762 | 24.7531 | 24.956 | 999.909 | 0.0081 | 0.35941 | 351.105 | 387.572
Populus deltoides | LI2B |  | RACiR 300 to 800 | 276.6051 | 24.7626 | 24.9571 | 1000.31 | 0.0081022 | 0.35947 | 352.582 | 389.062
Populus deltoides | LI2B |  | RACiR 300 to 800 | 277.6303 | 24.9268 | 24.9664 | 999.884 | 0.0081022 | 0.35891 | 354.214 | 390.867
Populus deltoides | LI2B |  | RACiR 300 to 800 | 278.884 | 25.0033 | 24.9652 | 999.86 | 0.0081008 | 0.35892 | 355.707 | 392.449
Populus deltoides | LI2B |  | RACiR 300 to 800 | 280.4752 | 25.0217 | 24.9645 | 1000.22 | 0.0081046 | 0.3592 | 357.33 | 394.092
Populus deltoides | LI2B |  | RACiR 300 to 800 | 281.3983 | 25.1993 | 24.9667 | 999.911 | 0.0081049 | 0.35911 | 358.795 | 395.743
Populus deltoides | LI2B |  | RACiR 300 to 800 | 282.5515 | 25.3698 | 24.9636 | 999.998 | 0.0081037 | 0.35922 | 360.429 | 397.555
Populus deltoides | LI2B |  | RACiR 300 to 800 | 283.8827 | 25.4344 | 24.9605 | 999.906 | 0.0081046 | 0.35947 | 361.917 | 399.111
Populus deltoides | LI2B |  | RACiR 300 to 800 | 284.8989 | 25.6654 | 24.9587 | 1000.23 | 0.0081056 | 0.35967 | 363.564 | 400.999
Populus deltoides | LI2B |  | RACiR 300 to 800 | 286.2535 | 25.7273 | 24.9539 | 999.854 | 0.0081046 | 0.35991 | 365.07 | 402.576
Populus deltoides | LI2B |  | RACiR 300 to 800 | 288.1029 | 25.6919 | 24.9497 | 999.988 | 0.0081083 | 0.36039 | 366.758 | 404.234
Populus deltoides | LI2B |  | RACiR 300 to 800 | 289.0986 | 25.8372 | 24.9539 | 999.944 | 0.0081079 | 0.36012 | 368.245 | 405.873
Populus deltoides | LI2B |  | RACiR 300 to 800 | 290.4319 | 25.9293 | 24.9571 | 999.904 | 0.0081078 | 0.35993 | 369.909 | 407.638
Populus deltoides | LI2B |  | RACiR 300 to 800 | 291.8474 | 25.9342 | 24.9596 | 999.963 | 0.0081082 | 0.35982 | 371.395 | 409.134
Populus deltoides | LI2B |  | RACiR 300 to 800 | 292.7564 | 26.158 | 24.9656 | 1000.19 | 0.0081081 | 0.35946 | 373.035 | 411.005
Populus deltoides | LI2B |  | RACiR 300 to 800 | 293.823 | 26.2833 | 24.9675 | 999.934 | 0.0081065 | 0.35929 | 374.517 | 412.622
Populus deltoides | LI2B |  | RACiR 300 to 800 | 295.3246 | 26.3218 | 24.9692 | 999.963 | 0.0081075 | 0.35926 | 376.169 | 414.317
Populus deltoides | LI2B |  | RACiR 300 to 800 | 296.186 | 26.5417 | 24.9692 | 999.892 | 0.0081088 | 0.35936 | 377.647 | 416.022
Populus deltoides | LI2B |  | RACiR 300 to 800 | 297.6076 | 26.6313 | 24.9666 | 1000.22 | 0.0081087 | 0.35955 | 379.31 | 417.785
Populus deltoides | LI2B |  | RACiR 300 to 800 | 299.203 | 26.6125 | 24.9638 | 999.959 | 0.0081112 | 0.35985 | 380.821 | 419.279
Populus deltoides | LI2B |  | RACiR 300 to 800 | 300.4266 | 26.7827 | 24.9613 | 999.915 | 0.0081129 | 0.36013 | 382.488 | 421.126
Populus deltoides | LI2B |  | RACiR 300 to 800 | 301.7391 | 26.8477 | 24.9599 | 999.963 | 0.0081122 | 0.3602 | 383.996 | 422.71
Populus deltoides | LI2B |  | RACiR 300 to 800 | 303.4318 | 26.8395 | 24.958 | 999.923 | 0.008115 | 0.36049 | 385.641 | 424.353
Populus deltoides | LI2B |  | RACiR 300 to 800 | 304.5087 | 26.984 | 24.9578 | 999.95 | 0.0081162 | 0.3606 | 387.122 | 425.985
Populus deltoides | LI2B |  | RACiR 300 to 800 | 305.9244 | 27.0517 | 24.9588 | 999.904 | 0.0081161 | 0.36056 | 388.77 | 427.711
Populus deltoides | LI2B |  | RACiR 300 to 800 | 307.2248 | 27.0904 | 24.9633 | 999.968 | 0.0081189 | 0.36045 | 390.236 | 429.218
Populus deltoides | LI2B |  | RACiR 300 to 800 | 308.1666 | 27.3066 | 24.9706 | 999.93 | 0.0081231 | 0.36028 | 391.852 | 431.059
Populus deltoides | LI2B |  | RACiR 300 to 800 | 310.1964 | 27.4344 | 24.9745 | 999.925 | 0.0081418 | 0.36093 | 394.15 | 433.52
Populus deltoides | LI2B |  | RACiR 300 to 800 | 310.1991 | 27.4344 | 24.9745 | 999.925 | 0.0081418 | 0.36095 | 394.15 | 433.52
Populus deltoides | LI2B |  | RACiR 300 to 800 | 313.1598 | 27.2422 | 24.9791 | 999.911 | 0.0081671 | 0.36179 | 396.468 | 435.65
Populus deltoides | LI2B |  | RACiR 300 to 800 | 314.6563 | 27.3068 | 24.9774 | 999.914 | 0.0081788 | 0.36232 | 398.067 | 437.328
Populus deltoides | LI2B |  | RACiR 300 to 800 | 315.6678 | 27.4649 | 24.9713 | 1000.25 | 0.0081785 | 0.36251 | 399.499 | 438.921
Populus deltoides | LI2B |  | RACiR 300 to 800 | 316.2331 | 27.8038 | 24.9705 | 999.953 | 0.0081738 | 0.36214 | 401.104 | 440.864
Populus deltoides | LI2B |  | RACiR 300 to 800 | 316.3428 | 28.2493 | 24.9704 | 999.91 | 0.0081655 | 0.36157 | 402.591 | 442.809
Populus deltoides | LI2B |  | RACiR 300 to 800 | 317.5835 | 28.386 | 24.967 | 999.968 | 0.0081597 | 0.36136 | 404.286 | 444.642
Populus deltoides | LI2B |  | RACiR 300 to 800 | 318.8708 | 28.4469 | 24.9687 | 1000.26 | 0.0081563 | 0.36101 | 405.855 | 446.278
Populus deltoides | LI2B |  | RACiR 300 to 800 | 320.4597 | 28.4567 | 24.9696 | 999.973 | 0.0081532 | 0.36076 | 407.566 | 448.007
Populus deltoides | LI2B |  | RACiR 300 to 800 | 322.0216 | 28.4398 | 24.9668 | 999.937 | 0.0081511 | 0.36086 | 409.091 | 449.52
Populus deltoides | LI2B |  | RACiR 300 to 800 | 323.0271 | 28.6536 | 24.9684 | 999.913 | 0.0081483 | 0.36064 | 410.769 | 451.419
Populus deltoides | LI2B |  | RACiR 300 to 800 | 324.2823 | 28.7163 | 24.9715 | 1000.19 | 0.0081454 | 0.36034 | 412.3 | 453.02
Populus deltoides | LI2B |  | RACiR 300 to 800 | 326.0755 | 28.6726 | 24.9694 | 999.965 | 0.0081447 | 0.36047 | 413.979 | 454.659
Populus deltoides | LI2B |  | RACiR 300 to 800 | 327.2512 | 28.7936 | 24.9671 | 999.99 | 0.0081446 | 0.36063 | 415.481 | 456.289
Populus deltoides | LI2B |  | RACiR 300 to 800 | 328.8684 | 28.811 | 24.968 | 999.908 | 0.0081468 | 0.36074 | 417.158 | 457.996
Populus deltoides | LI2B |  | RACiR 300 to 800 | 330.5497 | 28.7736 | 24.9653 | 999.932 | 0.0081503 | 0.36116 | 418.673 | 459.479
Populus deltoides | LI2B |  | RACiR 300 to 800 | 331.6525 | 28.9623 | 24.9655 | 1000.23 | 0.0081502 | 0.36119 | 420.323 | 461.327
Populus deltoides | LI2B |  | RACiR 300 to 800 | 332.7207 | 29.0913 | 24.9726 | 999.976 | 0.008154 | 0.36106 | 421.812 | 462.956
Populus deltoides | LI2B |  | RACiR 300 to 800 | 334.1396 | 29.1668 | 24.9739 | 999.959 | 0.0081553 | 0.36106 | 423.477 | 464.7
Populus deltoides | LI2B |  | RACiR 300 to 800 | 334.9471 | 29.3787 | 24.98 | 999.988 | 0.0081567 | 0.3608 | 424.963 | 466.406
Populus deltoides | LI2B |  | RACiR 300 to 800 | 336.3721 | 29.4471 | 24.9834 | 999.97 | 0.008158 | 0.36068 | 426.644 | 468.166
Populus deltoides | LI2B |  | RACiR 300 to 800 | 337.9844 | 29.4175 | 24.9818 | 999.935 | 0.0081591 | 0.36084 | 428.171 | 469.665
Populus deltoides | LI2B |  | RACiR 300 to 800 | 339.3017 | 29.5383 | 24.9821 | 999.985 | 0.0081587 | 0.36086 | 429.852 | 471.478
Populus deltoides | LI2B |  | RACiR 300 to 800 | 340.6582 | 29.5935 | 24.9798 | 999.957 | 0.0081591 | 0.36102 | 431.353 | 473.041
Populus deltoides | LI2B |  | RACiR 300 to 800 | 342.3756 | 29.5782 | 24.9785 | 999.956 | 0.0081614 | 0.36127 | 433.008 | 474.688
Populus deltoides | LI2B |  | RACiR 300 to 800 | 343.5802 | 29.6793 | 24.9774 | 999.969 | 0.0081619 | 0.36137 | 434.497 | 476.285
Populus deltoides | LI2B |  | RACiR 300 to 800 | 344.9926 | 29.7483 | 24.9777 | 999.981 | 0.0081623 | 0.36139 | 436.129 | 477.996
Populus deltoides | LI2B |  | RACiR 300 to 800 | 346.3738 | 29.7681 | 24.9795 | 1000.02 | 0.0081643 | 0.36141 | 437.594 | 479.484
Populus deltoides | LI2B |  | RACiR 300 to 800 | 347.4373 | 29.9547 | 24.9814 | 999.93 | 0.0081625 | 0.36122 | 439.252 | 481.34
Populus deltoides | LI2B |  | RACiR 300 to 800 | 348.3647 | 30.1098 | 24.9885 | 999.976 | 0.008163 | 0.36085 | 440.73 | 482.982
Populus deltoides | LI2B |  | RACiR 300 to 800 | 350.0246 | 30.1226 | 24.9879 | 1000.01 | 0.0081651 | 0.36103 | 442.42 | 484.69
Populus deltoides | LI2B |  | RACiR 300 to 800 | 351.1272 | 30.2655 | 24.9873 | 999.955 | 0.0081658 | 0.36113 | 443.925 | 486.345
Populus deltoides | LI2B |  | RACiR 300 to 800 | 352.6839 | 30.3195 | 24.9854 | 1000.23 | 0.0081674 | 0.36138 | 445.605 | 488.091
Populus deltoides | LI2B |  | RACiR 300 to 800 | 354.3598 | 30.2892 | 24.9794 | 999.987 | 0.0081702 | 0.3619 | 447.101 | 489.557
Populus deltoides | LI2B |  | RACiR 300 to 800 | 355.4607 | 30.4492 | 24.9862 | 999.974 | 0.0081717 | 0.36163 | 448.747 | 491.375
Populus deltoides | LI2B |  | RACiR 300 to 800 | 356.754 | 30.5168 | 24.9873 | 999.967 | 0.008173 | 0.36168 | 450.248 | 492.954
Populus deltoides | LI2B |  | RACiR 300 to 800 | 358.4138 | 30.4939 | 24.9894 | 999.966 | 0.0081736 | 0.36159 | 451.904 | 494.589
Populus deltoides | LI2B |  | RACiR 300 to 800 | 359.379 | 30.6664 | 24.9937 | 999.972 | 0.0081746 | 0.36147 | 453.408 | 496.279
Populus deltoides | LI2B |  | RACiR 300 to 800 | 360.6451 | 30.8123 | 24.9927 | 999.984 | 0.0081764 | 0.36165 | 455.065 | 498.088
Populus deltoides | LI2B |  | RACiR 300 to 800 | 362.0177 | 30.8487 | 24.9948 | 1000.32 | 0.0081776 | 0.36161 | 456.583 | 499.646
Populus deltoides | LI2B |  | RACiR 300 to 800 | 362.942 | 31.098 | 24.9992 | 999.996 | 0.0081797 | 0.3615 | 458.257 | 501.582
Populus deltoides | LI2B |  | RACiR 300 to 800 | 364.4225 | 31.1436 | 24.9943 | 999.979 | 0.008181 | 0.36188 | 459.8 | 503.178
Populus deltoides | LI2B |  | RACiR 300 to 800 | 366.2747 | 31.0836 | 24.9954 | 999.931 | 0.0081826 | 0.36193 | 461.515 | 504.84
Populus deltoides | LI2B |  | RACiR 300 to 800 | 367.4284 | 31.2045 | 24.9959 | 999.975 | 0.0081831 | 0.36195 | 463.027 | 506.482
Populus deltoides | LI2B |  | RACiR 300 to 800 | 368.9892 | 31.2509 | 24.9959 | 999.999 | 0.0081847 | 0.3621 | 464.716 | 508.228
Populus deltoides | LI2B |  | RACiR 300 to 800 | 370.5085 | 31.2122 | 25.003 | 999.99 | 0.0081876 | 0.36187 | 466.226 | 509.702
Populus deltoides | LI2B |  | RACiR 300 to 800 | 371.4968 | 31.4321 | 25.0046 | 1000.01 | 0.0081865 | 0.36174 | 467.887 | 511.593
Populus deltoides | LI2B |  | RACiR 300 to 800 | 372.8059 | 31.5158 | 25.0017 | 999.944 | 0.0081865 | 0.36192 | 469.41 | 513.208
Populus deltoides | LI2B |  | RACiR 300 to 800 | 374.557 | 31.506 | 24.9991 | 1000.29 | 0.0081883 | 0.36219 | 471.103 | 514.898
Populus deltoides | LI2B |  | RACiR 300 to 800 | 375.5468 | 31.6962 | 24.9985 | 999.992 | 0.0081889 | 0.36226 | 472.63 | 516.621
Populus deltoides | LI2B |  | RACiR 300 to 800 | 377.2127 | 31.7146 | 24.9987 | 999.964 | 0.0081897 | 0.36232 | 474.369 | 518.391
Populus deltoides | LI2B |  | RACiR 300 to 800 | 378.883 | 31.6455 | 25.0014 | 1000.01 | 0.0081917 | 0.36228 | 475.899 | 519.854
Populus deltoides | LI2B |  | RACiR 300 to 800 | 380.161 | 31.7764 | 25.0029 | 999.972 | 0.0081917 | 0.36222 | 477.587 | 521.683
Populus deltoides | LI2B |  | RACiR 300 to 800 | 381.5813 | 31.8169 | 25 | 1000.01 | 0.0081912 | 0.3624 | 479.104 | 523.251
Populus deltoides | LI2B |  | RACiR 300 to 800 | 383.4312 | 31.7492 | 25.001 | 999.994 | 0.0081936 | 0.3625 | 480.781 | 524.865
Populus deltoides | LI2B |  | RACiR 300 to 800 | 384.5844 | 31.8757 | 25.0004 | 999.956 | 0.0081946 | 0.36262 | 482.283 | 526.501
Populus deltoides | LI2B |  | RACiR 300 to 800 | 386.0711 | 31.9314 | 25.0039 | 999.957 | 0.0081996 | 0.36269 | 483.946 | 528.232
Populus deltoides | LI2B |  | RACiR 300 to 800 | 387.7579 | 31.9151 | 25.0037 | 1000.27 | 0.0082143 | 0.36339 | 485.453 | 529.733
Populus deltoides | LI2B |  | RACiR 300 to 800 | 389.0129 | 32.0998 | 25.0059 | 1000 | 0.0082293 | 0.36395 | 487.116 | 531.598
Populus deltoides | LI2B |  | RACiR 300 to 800 | 390.2599 | 32.2005 | 25.0091 | 999.984 | 0.0082405 | 0.3642 | 488.618 | 533.219
Populus deltoides | LI2B |  | RACiR 300 to 800 | 392.0655 | 32.2069 | 25.0049 | 999.936 | 0.008252 | 0.36484 | 490.321 | 534.936
Populus deltoides | LI2B |  | RACiR 300 to 800 | 392.9986 | 32.3752 | 25.0059 | 999.912 | 0.00825 | 0.36454 | 491.821 | 536.613
Populus deltoides | LI2B |  | RACiR 300 to 800 | 394.6285 | 32.384 | 24.9997 | 999.954 | 0.0082441 | 0.36447 | 493.526 | 538.334
Populus deltoides | LI2B |  | RACiR 300 to 800 | 396.0974 | 32.3673 | 24.9997 | 999.997 | 0.0082409 | 0.3642 | 495.052 | 539.841
Populus deltoides | LI2B |  | RACiR 300 to 800 | 397.326 | 32.4846 | 25.0031 | 999.963 | 0.0082376 | 0.36379 | 496.744 | 541.661
Populus deltoides | LI2B |  | RACiR 300 to 800 | 398.6484 | 32.5356 | 25.0031 | 1000.03 | 0.0082357 | 0.36366 | 498.272 | 543.247
Populus deltoides | LI2B |  | RACiR 300 to 800 | 400.4696 | 32.4695 | 25.0035 | 1000.01 | 0.0082335 | 0.3635 | 499.992 | 544.904
Populus deltoides | LI2B |  | RACiR 300 to 800 | 401.5367 | 32.6213 | 25.0023 | 999.987 | 0.0082317 | 0.36346 | 501.515 | 546.587
Populus deltoides | LI2B |  | RACiR 300 to 800 | 402.8708 | 32.7314 | 25.0017 | 1000.01 | 0.0082292 | 0.36336 | 503.212 | 548.406
Populus deltoides | LI2B |  | RACiR 300 to 800 | 404.4168 | 32.7166 | 25.0021 | 1000 | 0.0082294 | 0.36333 | 504.761 | 549.941
Populus deltoides | LI2B |  | RACiR 300 to 800 | 405.4702 | 32.9174 | 25.0047 | 999.981 | 0.0082274 | 0.36307 | 506.467 | 551.858
Populus deltoides | LI2B |  | RACiR 300 to 800 | 406.8916 | 32.9661 | 25.0012 | 1000.03 | 0.0082266 | 0.36323 | 508.011 | 553.461
Populus deltoides | LI2B |  | RACiR 300 to 800 | 408.7818 | 32.908 | 24.9997 | 1000.01 | 0.0082294 | 0.36349 | 509.712 | 555.106
Populus deltoides | LI2B |  | RACiR 300 to 800 | 410.066 | 33.0015 | 24.9975 | 999.944 | 0.0082298 | 0.36366 | 511.238 | 556.734
Populus deltoides | LI2B |  | RACiR 300 to 800 | 411.848 | 32.9906 | 24.9941 | 999.953 | 0.0082307 | 0.36395 | 512.95 | 558.45
Populus deltoides | LI2B |  | RACiR 300 to 800 | 413.6003 | 32.9126 | 24.9929 | 999.965 | 0.0082309 | 0.36402 | 514.506 | 559.931
Populus deltoides | LI2B |  | RACiR 300 to 800 | 414.7854 | 33.1094 | 24.9898 | 999.945 | 0.0082316 | 0.36426 | 516.201 | 561.83
Populus deltoides | LI2B |  | RACiR 300 to 800 | 416.1562 | 33.1629 | 24.9901 | 999.979 | 0.0082309 | 0.36423 | 517.759 | 563.456
Populus deltoides | LI2B |  | RACiR 300 to 800 | 418.0276 | 33.1199 | 24.9863 | 999.899 | 0.008231 | 0.36445 | 519.491 | 565.146
Populus deltoides | LI2B |  | RACiR 300 to 800 | 418.9838 | 33.3017 | 24.9907 | 999.971 | 0.0082322 | 0.36427 | 521.026 | 566.869
Populus deltoides | LI2B |  | RACiR 300 to 800 | 420.6219 | 33.3203 | 24.9911 | 1000.29 | 0.0082315 | 0.36423 | 522.763 | 568.636
Populus deltoides | LI2B |  | RACiR 300 to 800 | 422.4035 | 33.2087 | 24.995 | 999.976 | 0.008232 | 0.36402 | 524.337 | 570.099
Populus deltoides | LI2B |  | RACiR 300 to 800 | 423.7735 | 33.3006 | 25.0016 | 999.945 | 0.0082329 | 0.36372 | 526.08 | 571.945
Populus deltoides | LI2B |  | RACiR 300 to 800 | 425.3301 | 33.3187 | 24.9965 | 1000 | 0.0082325 | 0.36402 | 527.636 | 573.53
Populus deltoides | LI2B |  | RACiR 300 to 800 | 427.3785 | 33.1928 | 24.9966 | 999.991 | 0.0082339 | 0.36411 | 529.358 | 575.133
Populus deltoides | LI2B |  | RACiR 300 to 800 | 428.7292 | 33.2782 | 24.9927 | 1000.01 | 0.008235 | 0.36443 | 530.887 | 576.757
Populus deltoides | LI2B |  | RACiR 300 to 800 | 430.5616 | 33.2461 | 24.9886 | 999.981 | 0.0082351 | 0.36469 | 532.6 | 578.451
Populus deltoides | LI2B |  | RACiR 300 to 800 | 432.1448 | 33.2122 | 24.9896 | 1000.01 | 0.0082364 | 0.36466 | 534.133 | 579.953
Populus deltoides | LI2B |  | RACiR 300 to 800 | 433.265 | 33.3923 | 24.993 | 999.982 | 0.0082368 | 0.36453 | 535.818 | 581.826
Populus deltoides | LI2B |  | RACiR 300 to 800 | 434.5134 | 33.4622 | 24.9974 | 999.924 | 0.0082372 | 0.3643 | 537.355 | 583.442
Populus deltoides | LI2B |  | RACiR 300 to 800 | 436.0889 | 33.488 | 24.9987 | 999.932 | 0.0082373 | 0.36425 | 539.051 | 585.17
Populus deltoides | LI2B |  | RACiR 300 to 800 | 437.2833 | 33.596 | 24.9998 | 1000.25 | 0.0082371 | 0.36423 | 540.579 | 586.814
Populus deltoides | LI2B |  | RACiR 300 to 800 | 439.123 | 33.575 | 24.9947 | 999.961 | 0.0082369 | 0.36455 | 542.311 | 588.534
Populus deltoides | LI2B |  | RACiR 300 to 800 | 441.0281 | 33.4761 | 24.9892 | 999.886 | 0.0082392 | 0.36501 | 543.857 | 589.984
Populus deltoides | LI2B |  | RACiR 300 to 800 | 442.5361 | 33.5462 | 24.9887 | 999.953 | 0.0082402 | 0.3651 | 545.569 | 591.775
Populus deltoides | LI2B |  | RACiR 300 to 800 | 443.9094 | 33.5865 | 24.9902 | 999.974 | 0.0082398 | 0.36499 | 547.117 | 593.375
Populus deltoides | LI2B |  | RACiR 300 to 800 | 445.7047 | 33.5403 | 24.9929 | 999.971 | 0.008242 | 0.36493 | 548.845 | 595.06
Populus deltoides | LI2B |  | RACiR 300 to 800 | 446.8747 | 33.668 | 24.9934 | 1000 | 0.0082407 | 0.36485 | 550.415 | 596.771
Populus deltoides | LI2B |  | RACiR 300 to 800 | 448.1668 | 33.7824 | 24.9966 | 999.975 | 0.0082402 | 0.36464 | 552.115 | 598.598
Populus deltoides | LI2B |  | RACiR 300 to 800 | 450.0469 | 33.6902 | 24.9944 | 1000.26 | 0.0082424 | 0.3649 | 553.711 | 600.102
Populus deltoides | LI2B |  | RACiR 300 to 800 | 451.2827 | 33.8379 | 24.9967 | 1000.02 | 0.008241 | 0.3647 | 555.44 | 601.99
Populus deltoides | LI2B |  | RACiR 300 to 800 | 452.6168 | 33.8813 | 25.0015 | 1000 | 0.0082406 | 0.3644 | 557.012 | 603.617
Populus deltoides | LI2B |  | RACiR 300 to 800 | 454.6213 | 33.7768 | 25.0038 | 1000.02 | 0.0082427 | 0.36439 | 558.778 | 605.279
Populus deltoides | LI2B |  | RACiR 300 to 800 | 455.9203 | 33.8497 | 25.0065 | 999.976 | 0.0082423 | 0.36425 | 560.348 | 606.934
Populus deltoides | LI2B |  | RACiR 300 to 800 | 457.6484 | 33.8487 | 25.0068 | 999.98 | 0.0082436 | 0.36433 | 562.091 | 608.686
Populus deltoides | LI2B |  | RACiR 300 to 800 | 459.7141 | 33.6843 | 25.005 | 999.914 | 0.0082461 | 0.36461 | 563.672 | 610.109
Populus deltoides | LI2B |  | RACiR 300 to 800 | 461.2171 | 33.8081 | 24.9973 | 999.984 | 0.0082464 | 0.36512 | 565.405 | 611.976
Populus deltoides | LI2B |  | RACiR 300 to 800 | 462.7099 | 33.8449 | 24.9945 | 999.98 | 0.0082464 | 0.36529 | 566.986 | 613.603
Populus deltoides | LI2B |  | RACiR 300 to 800 | 464.6939 | 33.763 | 24.9916 | 1000.27 | 0.0082474 | 0.3655 | 568.728 | 615.267
Populus deltoides | LI2B |  | RACiR 300 to 800 | 465.8813 | 33.9113 | 24.9925 | 999.977 | 0.0082543 | 0.36577 | 570.285 | 616.985
Populus deltoides | LI2B |  | RACiR 300 to 800 | 467.6797 | 33.9345 | 24.9961 | 999.997 | 0.0082732 | 0.3664 | 572.034 | 618.781
Populus deltoides | LI2B |  | RACiR 300 to 800 | 469.6337 | 33.8658 | 24.9961 | 999.911 | 0.0082932 | 0.36728 | 573.62 | 620.309
Populus deltoides | LI2B |  | RACiR 300 to 800 | 471.1821 | 33.9831 | 24.9966 | 999.96 | 0.008311 | 0.36794 | 575.362 | 622.191
Populus deltoides | LI2B |  | RACiR 300 to 800 | 472.7413 | 34.0193 | 24.9922 | 1000.09 | 0.0083196 | 0.36841 | 576.934 | 623.813
Populus deltoides | LI2B |  | RACiR 300 to 800 | 474.6743 | 33.8992 | 24.995 | 1000.04 | 0.008317 | 0.36792 | 578.718 | 625.479
Populus deltoides | LI2B |  | RACiR 300 to 800 | 475.9859 | 33.9372 | 24.9938 | 999.961 | 0.0083097 | 0.36746 | 580.285 | 627.092
Populus deltoides | LI2B |  | RACiR 300 to 800 | 477.8318 | 33.8552 | 24.9916 | 999.974 | 0.0083021 | 0.36706 | 582.052 | 628.784
Populus deltoides | LI2B |  | RACiR 300 to 800 | 479.6652 | 33.7162 | 24.9937 | 999.976 | 0.0082985 | 0.36669 | 583.646 | 630.241
Populus deltoides | LI2B |  | RACiR 300 to 800 | 481.0533 | 33.8379 | 24.9875 | 999.959 | 0.0082954 | 0.36684 | 585.353 | 632.079
Populus deltoides | LI2B |  | RACiR 300 to 800 | 482.3691 | 33.8975 | 24.9874 | 999.984 | 0.0082932 | 0.36671 | 586.897 | 633.691
Populus deltoides | LI2B |  | RACiR 300 to 800 | 484.3129 | 33.8427 | 24.9819 | 999.963 | 0.0082941 | 0.36707 | 588.632 | 635.371
Populus deltoides | LI2B |  | RACiR 300 to 800 | 485.6494 | 33.9445 | 24.9769 | 999.955 | 0.0082936 | 0.36731 | 590.208 | 637.059
Populus deltoides | LI2B |  | RACiR 300 to 800 | 487.306 | 33.9422 | 24.9782 | 999.943 | 0.0082924 | 0.36712 | 591.949 | 638.809
Populus deltoides | LI2B |  | RACiR 300 to 800 | 488.9977 | 33.8707 | 24.9802 | 999.985 | 0.0082933 | 0.36703 | 593.51 | 640.302
Populus deltoides | LI2B |  | RACiR 300 to 800 | 490.3894 | 33.9413 | 24.9868 | 999.905 | 0.0082919 | 0.36656 | 595.262 | 642.135
Populus deltoides | LI2B |  | RACiR 300 to 800 | 491.7943 | 33.9457 | 24.9911 | 999.951 | 0.0082914 | 0.36628 | 596.793 | 643.68
Populus deltoides | LI2B |  | RACiR 300 to 800 | 493.8725 | 33.824 | 24.9887 | 999.972 | 0.0082927 | 0.36651 | 598.518 | 645.288
Populus deltoides | LI2B |  | RACiR 300 to 800 | 495.1428 | 33.9525 | 24.9838 | 1000 | 0.0082949 | 0.36695 | 600.046 | 646.953
Populus deltoides | LI2B |  | RACiR 300 to 800 | 496.6643 | 34.0233 | 24.9823 | 999.94 | 0.0082937 | 0.36699 | 601.784 | 648.772
Populus deltoides | LI2B |  | RACiR 300 to 800 | 498.2745 | 34.0163 | 24.9817 | 1000.27 | 0.0082965 | 0.36721 | 603.354 | 650.34
Populus deltoides | LI2B |  | RACiR 300 to 800 | 499.5601 | 34.1552 | 24.9849 | 999.998 | 0.008296 | 0.36701 | 605.109 | 652.244
Populus deltoides | LI2B |  | RACiR 300 to 800 | 501.1015 | 34.1224 | 24.9925 | 999.959 | 0.0082954 | 0.36654 | 606.73 | 653.842
Populus deltoides | LI2B |  | RACiR 300 to 800 | 503.3443 | 33.9708 | 24.9891 | 999.901 | 0.0082974 | 0.36689 | 608.509 | 655.477
Populus deltoides | LI2B |  | RACiR 300 to 800 | 504.8035 | 34.0267 | 24.9851 | 999.877 | 0.0082962 | 0.3671 | 610.093 | 657.127
Populus deltoides | LI2B |  | RACiR 300 to 800 | 506.5964 | 33.9939 | 24.9867 | 999.981 | 0.0082969 | 0.36705 | 611.85 | 658.859
Populus deltoides | LI2B |  | RACiR 300 to 800 | 508.5918 | 33.8769 | 24.9801 | 1000.31 | 0.0082972 | 0.36747 | 613.45 | 660.349
Populus deltoides | LI2B |  | RACiR 0 to 500 | 95.4228 | 3.5924 | 24.9546 | 1000.17 | 0.0070708 | 0.31169 | 109.27 | 123.58
Populus deltoides | LI2B |  | RACiR 0 to 500 | 96.4269 | 3.7262 | 24.9519 | 1000.13 | 0.0070713 | 0.3119 | 110.722 | 125.174
Populus deltoides | LI2B |  | RACiR 0 to 500 | 97.9324 | 3.7591 | 24.9522 | 1000.18 | 0.0070725 | 0.31198 | 112.365 | 126.859
Populus deltoides | LI2B |  | RACiR 0 to 500 | 98.7607 | 3.9444 | 24.9506 | 1000.2 | 0.0070722 | 0.31209 | 113.808 | 128.494
Populus deltoides | LI2B |  | RACiR 0 to 500 | 100.0583 | 4.0327 | 24.9515 | 999.844 | 0.0070726 | 0.3121 | 115.421 | 130.205
Populus deltoides | LI2B |  | RACiR 0 to 500 | 101.5444 | 4.0164 | 24.9511 | 999.814 | 0.0070737 | 0.31218 | 116.885 | 131.659
Populus deltoides | LI2B |  | RACiR 0 to 500 | 102.5198 | 4.198 | 24.9527 | 1000.17 | 0.0070746 | 0.31215 | 118.473 | 133.437
Populus deltoides | LI2B |  | RACiR 0 to 500 | 103.6748 | 4.2762 | 24.9584 | 999.837 | 0.0070759 | 0.31196 | 119.918 | 134.97
Populus deltoides | LI2B |  | RACiR 0 to 500 | 105.186 | 4.291 | 24.9579 | 999.884 | 0.0070776 | 0.31208 | 121.506 | 136.58
Populus deltoides | LI2B |  | RACiR 0 to 500 | 106.1255 | 4.428 | 24.9616 | 1000.22 | 0.0070798 | 0.31204 | 122.914 | 138.134
Populus deltoides | LI2B |  | RACiR 0 to 500 | 107.3667 | 4.5161 | 24.9626 | 1000.2 | 0.0070819 | 0.31215 | 124.464 | 139.781
Populus deltoides | LI2B |  | RACiR 0 to 500 | 108.5967 | 4.5596 | 24.9588 | 999.831 | 0.0070836 | 0.31246 | 125.847 | 141.214
Populus deltoides | LI2B |  | RACiR 0 to 500 | 109.4298 | 4.7766 | 24.9573 | 1000.16 | 0.0070838 | 0.31257 | 127.395 | 142.988
Populus deltoides | LI2B |  | RACiR 0 to 500 | 110.287 | 4.9352 | 24.9586 | 1000.17 | 0.007085 | 0.3126 | 128.784 | 144.543
Populus deltoides | LI2B |  | RACiR 0 to 500 | 111.5276 | 5.0298 | 24.956 | 1000.14 | 0.0070874 | 0.31288 | 130.343 | 146.205
Populus deltoides | LI2B |  | RACiR 0 to 500 | 112.3499 | 5.2068 | 24.9583 | 1000.17 | 0.0070886 | 0.31287 | 131.757 | 147.804
Populus deltoides | LI2B |  | RACiR 0 to 500 | 113.588 | 5.3022 | 24.9602 | 999.858 | 0.0070914 | 0.31298 | 133.326 | 149.478
Populus deltoides | LI2B |  | RACiR 0 to 500 | 114.8738 | 5.3327 | 24.9594 | 999.824 | 0.0070938 | 0.31318 | 134.728 | 150.917
Populus deltoides | LI2B |  | RACiR 0 to 500 | 115.8047 | 5.5135 | 24.964 | 1000.22 | 0.0070941 | 0.31299 | 136.277 | 152.654
Populus deltoides | LI2B |  | RACiR 0 to 500 | 116.8113 | 5.6279 | 24.9648 | 999.84 | 0.0070953 | 0.31306 | 137.672 | 154.171
Populus deltoides | LI2B |  | RACiR 0 to 500 | 118.378 | 5.6066 | 24.965 | 999.866 | 0.0070954 | 0.31306 | 139.206 | 155.693
Populus deltoides | LI2B |  | RACiR 0 to 500 | 119.5547 | 5.6682 | 24.9624 | 1000.2 | 0.0070976 | 0.31333 | 140.591 | 157.147
Populus deltoides | LI2B |  | RACiR 0 to 500 | 120.2406 | 5.9173 | 24.9602 | 1000.21 | 0.0070992 | 0.31357 | 142.081 | 158.896
Populus deltoides | LI2B |  | RACiR 0 to 500 | 121.0038 | 6.1015 | 24.9576 | 999.798 | 0.0071023 | 0.31386 | 143.436 | 160.437
Populus deltoides | LI2B |  | RACiR 0 to 500 | 121.4528 | 6.4365 | 24.9598 | 1000.13 | 0.007101 | 0.31373 | 144.984 | 162.331
Populus deltoides | LI2B |  | RACiR 0 to 500 | 121.762 | 6.7766 | 24.9593 | 1000.03 | 0.0071011 | 0.31382 | 146.39 | 164.087
Populus deltoides | LI2B |  | RACiR 0 to 500 | 123.2964 | 6.798 | 24.9576 | 1000.13 | 0.0071048 | 0.31413 | 148.005 | 165.73
Populus deltoides | LI2B |  | RACiR 0 to 500 | 124.3994 | 6.9032 | 24.9569 | 1000.15 | 0.0071058 | 0.31425 | 149.463 | 167.302
Populus deltoides | LI2B |  | RACiR 0 to 500 | 125.7148 | 6.9796 | 24.9554 | 1000.19 | 0.0071067 | 0.31441 | 151.042 | 168.968
Populus deltoides | LI2B |  | RACiR 0 to 500 | 127.1398 | 6.973 | 24.9596 | 1000.12 | 0.0071089 | 0.31431 | 152.487 | 170.411
Populus deltoides | LI2B |  | RACiR 0 to 500 | 127.8789 | 7.2182 | 24.9649 | 999.78 | 0.0071096 | 0.3141 | 154.05 | 172.229
Populus deltoides | LI2B |  | RACiR 0 to 500 | 128.8976 | 7.3446 | 24.968 | 1000.17 | 0.0071225 | 0.31456 | 155.466 | 173.782
Populus deltoides | LI2B |  | RACiR 0 to 500 | 130.3568 | 7.3943 | 24.9712 | 999.886 | 0.0071462 | 0.31551 | 157.048 | 175.423
Populus deltoides | LI2B |  | RACiR 0 to 500 | 131.0572 | 7.618 | 24.9747 | 1000.17 | 0.007164 | 0.31607 | 158.44 | 177.051
Populus deltoides | LI2B |  | RACiR 0 to 500 | 132.2824 | 7.7395 | 24.9755 | 1000.17 | 0.0071838 | 0.31685 | 160.022 | 178.767
Populus deltoides | LI2B |  | RACiR 0 to 500 | 133.4889 | 7.7977 | 24.9749 | 999.846 | 0.0071897 | 0.31699 | 161.433 | 180.244
Populus deltoides | LI2B |  | RACiR 0 to 500 | 134.3235 | 8.0134 | 24.9754 | 1000.14 | 0.0071857 | 0.31662 | 163.007 | 182.041
Populus deltoides | LI2B |  | RACiR 0 to 500 | 135.394 | 8.1131 | 24.971 | 1000.07 | 0.0071806 | 0.31647 | 164.432 | 183.573
Populus deltoides | LI2B |  | RACiR 0 to 500 | 136.8159 | 8.1456 | 24.9684 | 1000.16 | 0.0071753 | 0.31623 | 166.01 | 185.19
Populus deltoides | LI2B |  | RACiR 0 to 500 | 137.6999 | 8.3031 | 24.966 | 1000.19 | 0.0071722 | 0.31617 | 167.422 | 186.768
Populus deltoides | LI2B |  | RACiR 0 to 500 | 138.8881 | 8.4151 | 24.9639 | 1000.1 | 0.0071708 | 0.31619 | 168.993 | 188.46
Populus deltoides | LI2B |  | RACiR 0 to 500 | 140.006 | 8.4946 | 24.9645 | 1000.18 | 0.0071703 | 0.31612 | 170.398 | 189.95
Populus deltoides | LI2B |  | RACiR 0 to 500 | 140.7122 | 8.7495 | 24.9632 | 999.845 | 0.0071676 | 0.31603 | 171.944 | 191.76
Populus deltoides | LI2B |  | RACiR 0 to 500 | 141.5559 | 8.9208 | 24.9611 | 1000.16 | 0.0071672 | 0.31609 | 173.348 | 193.343
Populus deltoides | LI2B |  | RACiR 0 to 500 | 142.7655 | 9.0257 | 24.9605 | 1000.15 | 0.0071654 | 0.316 | 174.93 | 195.037
Populus deltoides | LI2B |  | RACiR 0 to 500 | 143.5205 | 9.2196 | 24.9613 | 999.764 | 0.0071593 | 0.31569 | 176.352 | 196.66
Populus deltoides | LI2B |  | RACiR 0 to 500 | 144.7037 | 9.3342 | 24.9611 | 999.803 | 0.0071522 | 0.31544 | 177.953 | 198.384
Populus deltoides | LI2B |  | RACiR 0 to 500 | 146.0697 | 9.3435 | 24.9617 | 1000.13 | 0.0071467 | 0.31522 | 179.4 | 199.847
Populus deltoides | LI2B |  | RACiR 0 to 500 | 147.0522 | 9.5215 | 24.9618 | 1000.13 | 0.0071413 | 0.31508 | 180.989 | 201.622
Populus deltoides | LI2B |  | RACiR 0 to 500 | 148.1424 | 9.621 | 24.9631 | 1000.11 | 0.0071412 | 0.3151 | 182.421 | 203.161
Populus deltoides | LI2B |  | RACiR 0 to 500 | 149.5426 | 9.6659 | 24.9678 | 1000.2 | 0.0071454 | 0.31519 | 183.99 | 204.783
Populus deltoides | LI2B |  | RACiR 0 to 500 | 150.2273 | 9.8861 | 24.9732 | 1000.18 | 0.007148 | 0.31515 | 185.403 | 206.425
Populus deltoides | LI2B |  | RACiR 0 to 500 | 151.2867 | 10.0477 | 24.9749 | 999.882 | 0.0071518 | 0.31535 | 186.985 | 208.177
Populus deltoides | LI2B |  | RACiR 0 to 500 | 152.2884 | 10.1821 | 24.9735 | 999.811 | 0.0071564 | 0.31573 | 188.399 | 209.732
Populus deltoides | LI2B |  | RACiR 0 to 500 | 152.9204 | 10.4874 | 24.9706 | 1000.21 | 0.007159 | 0.31606 | 189.986 | 211.634
Populus deltoides | LI2B |  | RACiR 0 to 500 | 153.7944 | 10.666 | 24.9705 | 999.794 | 0.0071626 | 0.31629 | 191.425 | 213.261
Populus deltoides | LI2B |  | RACiR 0 to 500 | 155.0802 | 10.7624 | 24.9725 | 1000.15 | 0.0071662 | 0.31638 | 193.04 | 214.981
Populus deltoides | LI2B |  | RACiR 0 to 500 | 155.981 | 10.9421 | 24.9725 | 999.817 | 0.0071694 | 0.31661 | 194.51 | 216.64
Populus deltoides | LI2B |  | RACiR 0 to 500 | 157.356 | 11.0146 | 24.9728 | 999.791 | 0.0071727 | 0.31682 | 196.124 | 218.337
Populus deltoides | LI2B |  | RACiR 0 to 500 | 158.7834 | 11.0152 | 24.9757 | 1000.22 | 0.0071753 | 0.31683 | 197.586 | 219.806
Populus deltoides | LI2B |  | RACiR 0 to 500 | 159.6505 | 11.2367 | 24.9785 | 1000.15 | 0.0071769 | 0.31683 | 199.181 | 221.633
Populus deltoides | LI2B |  | RACiR 0 to 500 | 160.5613 | 11.3933 | 24.9801 | 999.794 | 0.0071785 | 0.31688 | 200.607 | 223.225
Populus deltoides | LI2B |  | RACiR 0 to 500 | 161.861 | 11.4868 | 24.9815 | 999.945 | 0.0071814 | 0.31699 | 202.222 | 224.94
Populus deltoides | LI2B |  | RACiR 0 to 500 | 162.5824 | 11.7262 | 24.9769 | 999.844 | 0.0071837 | 0.3174 | 203.671 | 226.637
Populus deltoides | LI2B |  | RACiR 0 to 500 | 163.6829 | 11.8861 | 24.9759 | 999.83 | 0.0071853 | 0.3176 | 205.282 | 228.418
Populus deltoides | LI2B |  | RACiR 0 to 500 | 165.0418 | 11.9066 | 24.9801 | 999.845 | 0.0071877 | 0.31755 | 206.744 | 229.908
Populus deltoides | LI2B |  | RACiR 0 to 500 | 166.0488 | 12.0903 | 24.9826 | 1000.23 | 0.0071907 | 0.31764 | 208.349 | 231.707
Populus deltoides | LI2B |  | RACiR 0 to 500 | 167.0827 | 12.2078 | 24.9863 | 999.877 | 0.0071921 | 0.31756 | 209.793 | 233.278
Populus deltoides | LI2B |  | RACiR 0 to 500 | 168.5511 | 12.2426 | 24.9881 | 999.85 | 0.0071941 | 0.3176 | 211.401 | 234.928
Populus deltoides | LI2B |  | RACiR 0 to 500 | 169.5095 | 12.3989 | 24.9878 | 1000.23 | 0.0071962 | 0.31777 | 212.857 | 236.548
Populus deltoides | LI2B |  | RACiR 0 to 500 | 170.7186 | 12.5213 | 24.9862 | 1000.22 | 0.0071969 | 0.31793 | 214.462 | 238.287
Populus deltoides | LI2B |  | RACiR 0 to 500 | 171.9233 | 12.5954 | 24.9851 | 999.854 | 0.0071997 | 0.31817 | 215.898 | 239.801
Populus deltoides | LI2B |  | RACiR 0 to 500 | 172.7115 | 12.8572 | 24.9853 | 1000.19 | 0.0072072 | 0.31856 | 217.486 | 241.661
Populus deltoides | LI2B |  | RACiR 0 to 500 | 173.5432 | 13.053 | 24.9864 | 1000.09 | 0.0072154 | 0.3189 | 218.914 | 243.296
Populus deltoides | LI2B |  | RACiR 0 to 500 | 174.8902 | 13.1458 | 24.9883 | 1000.12 | 0.0072247 | 0.31924 | 220.543 | 245.028
Populus deltoides | LI2B |  | RACiR 0 to 500 | 175.619 | 13.3714 | 24.9915 | 1000.16 | 0.0072328 | 0.31944 | 221.979 | 246.7
Populus deltoides | LI2B |  | RACiR 0 to 500 | 176.8979 | 13.4722 | 24.9947 | 999.825 | 0.0072369 | 0.31945 | 223.607 | 248.441
Populus deltoides | LI2B |  | RACiR 0 to 500 | 178.3623 | 13.4692 | 24.9955 | 999.837 | 0.0072397 | 0.31952 | 225.087 | 249.925
Populus deltoides | LI2B |  | RACiR 0 to 500 | 179.4802 | 13.6564 | 24.9978 | 1000.2 | 0.0072649 | 0.32052 | 226.69 | 251.73
Populus deltoides | LI2B |  | RACiR 0 to 500 | 180.8088 | 13.7405 | 24.9968 | 999.827 | 0.0072938 | 0.32183 | 228.138 | 253.279
Populus deltoides | LI2B |  | RACiR 0 to 500 | 182.4236 | 13.7872 | 24.9926 | 999.827 | 0.0073244 | 0.32324 | 229.747 | 254.947
Populus deltoides | LI2B |  | RACiR 0 to 500 | 183.1222 | 14.0597 | 24.9876 | 1000.19 | 0.0073482 | 0.32436 | 231.161 | 256.647
Populus deltoides | LI2B |  | RACiR 0 to 500 | 184.1239 | 14.2505 | 24.9822 | 1000.16 | 0.0073512 | 0.32447 | 232.766 | 258.453
Populus deltoides | LI2B |  | RACiR 0 to 500 | 185.1621 | 14.3552 | 24.9796 | 999.855 | 0.0073431 | 0.32395 | 234.229 | 260.024
Populus deltoides | LI2B |  | RACiR 0 to 500 | 185.82 | 14.6313 | 24.9761 | 999.772 | 0.0073323 | 0.32336 | 235.848 | 261.927
Populus deltoides | LI2B |  | RACiR 0 to 500 | 186.9542 | 14.7171 | 24.9722 | 999.922 | 0.0073239 | 0.323 | 237.329 | 263.503
Populus deltoides | LI2B |  | RACiR 0 to 500 | 188.4099 | 14.7413 | 24.9738 | 999.822 | 0.0073194 | 0.32257 | 238.958 | 265.161
Populus deltoides | LI2B |  | RACiR 0 to 500 | 189.1479 | 14.936 | 24.9772 | 999.817 | 0.0073141 | 0.32205 | 240.404 | 266.81
Populus deltoides | LI2B |  | RACiR 0 to 500 | 190.378 | 15.0297 | 24.9811 | 1000.12 | 0.0073094 | 0.32158 | 242.029 | 268.539
Populus deltoides | LI2B |  | RACiR 0 to 500 | 191.5405 | 15.0854 | 24.9883 | 1000.14 | 0.0073076 | 0.32107 | 243.475 | 270.044
Populus deltoides | LI2B |  | RACiR 0 to 500 | 192.2672 | 15.3471 | 24.9899 | 999.888 | 0.0073045 | 0.3208 | 245.086 | 271.928
Populus deltoides | LI2B |  | RACiR 0 to 500 | 193.1988 | 15.5086 | 24.9886 | 1000.15 | 0.0073036 | 0.3208 | 246.548 | 273.559
Populus deltoides | LI2B |  | RACiR 0 to 500 | 194.5416 | 15.5888 | 24.9881 | 999.844 | 0.0073036 | 0.32077 | 248.18 | 275.278
Populus deltoides | LI2B |  | RACiR 0 to 500 | 195.3731 | 15.8013 | 24.982 | 1000.17 | 0.0073019 | 0.32101 | 249.658 | 276.977
Populus deltoides | LI2B |  | RACiR 0 to 500 | 196.5611 | 15.9314 | 24.9785 | 1000.18 | 0.0072972 | 0.32096 | 251.288 | 278.745
Populus deltoides | LI2B |  | RACiR 0 to 500 | 197.8825 | 15.9646 | 24.9764 | 999.875 | 0.0072934 | 0.32093 | 252.748 | 280.245
Populus deltoides | LI2B |  | RACiR 0 to 500 | 198.8777 | 16.1516 | 24.9751 | 999.81 | 0.0072891 | 0.32083 | 254.369 | 282.059
Populus deltoides | LI2B |  | RACiR 0 to 500 | 199.8662 | 16.2734 | 24.9807 | 999.932 | 0.0072865 | 0.3205 | 255.822 | 283.644
Populus deltoides | LI2B |  | RACiR 0 to 500 | 201.3477 | 16.3117 | 24.9837 | 999.819 | 0.0072897 | 0.32057 | 257.448 | 285.315
Populus deltoides | LI2B |  | RACiR 0 to 500 | 202.3328 | 16.4477 | 24.9889 | 999.825 | 0.007291 | 0.3204 | 258.915 | 286.927
Populus deltoides | LI2B |  | RACiR 0 to 500 | 203.5227 | 16.579 | 24.9897 | 999.794 | 0.0072935 | 0.32054 | 260.525 | 288.681
Populus deltoides | LI2B |  | RACiR 0 to 500 | 204.843 | 16.6273 | 24.9882 | 999.837 | 0.0072967 | 0.32082 | 261.98 | 290.19
Populus deltoides | LI2B |  | RACiR 0 to 500 | 205.7869 | 16.8387 | 24.9882 | 1000.25 | 0.0072977 | 0.32092 | 263.593 | 292.031
Populus deltoides | LI2B |  | RACiR 0 to 500 | 206.697 | 17.0225 | 24.9832 | 999.888 | 0.0073006 | 0.3214 | 265.019 | 293.648
Populus deltoides | LI2B |  | RACiR 0 to 500 | 207.9976 | 17.1231 | 24.9846 | 1000.17 | 0.0073054 | 0.32164 | 266.624 | 295.361
Populus deltoides | LI2B |  | RACiR 0 to 500 | 208.8289 | 17.3107 | 24.9884 | 999.859 | 0.0073075 | 0.32163 | 268.068 | 297.002
Populus deltoides | LI2B |  | RACiR 0 to 500 | 210.053 | 17.4247 | 24.9913 | 999.886 | 0.0073098 | 0.32166 | 269.676 | 298.728
Populus deltoides | LI2B |  | RACiR 0 to 500 | 211.5394 | 17.4127 | 24.9924 | 1000.2 | 0.0073122 | 0.3218 | 271.133 | 300.182
Populus deltoides | LI2B |  | RACiR 0 to 500 | 212.5339 | 17.6055 | 24.9883 | 999.869 | 0.0073122 | 0.32204 | 272.712 | 301.965
Populus deltoides | LI2B |  | RACiR 0 to 500 | 213.6157 | 17.7346 | 24.9825 | 999.882 | 0.0073144 | 0.32246 | 274.146 | 303.536
Populus deltoides | LI2B |  | RACiR 0 to 500 | 215.0721 | 17.7921 | 24.9778 | 999.863 | 0.0073153 | 0.32276 | 275.759 | 305.217
Populus deltoides | LI2B |  | RACiR 0 to 500 | 215.9062 | 17.9723 | 24.9811 | 999.841 | 0.0073153 | 0.32259 | 277.21 | 306.858
Populus deltoides | LI2B |  | RACiR 0 to 500 | 217.0194 | 18.104 | 24.9874 | 999.841 | 0.0073145 | 0.32223 | 278.831 | 308.62
Populus deltoides | LI2B |  | RACiR 0 to 500 | 218.2902 | 18.152 | 24.9912 | 999.859 | 0.0073162 | 0.32217 | 280.295 | 310.138
Populus deltoides | LI2B |  | RACiR 0 to 500 | 219.1429 | 18.3826 | 24.9934 | 1000.22 | 0.007319 | 0.32226 | 281.876 | 311.96
Populus deltoides | LI2B |  | RACiR 0 to 500 | 220.0808 | 18.5516 | 24.9917 | 999.844 | 0.0073208 | 0.32248 | 283.323 | 313.587
Populus deltoides | LI2B |  | RACiR 0 to 500 | 221.3911 | 18.6616 | 24.9898 | 999.823 | 0.0073256 | 0.32288 | 284.933 | 315.315
Populus deltoides | LI2B |  | RACiR 0 to 500 | 222.147 | 18.8826 | 24.9902 | 1000.14 | 0.0073273 | 0.323 | 286.375 | 316.987
Populus deltoides | LI2B |  | RACiR 0 to 500 | 223.3058 | 19.0195 | 24.9938 | 1000.2 | 0.0073285 | 0.32293 | 288.005 | 318.767
Populus deltoides | LI2B |  | RACiR 0 to 500 | 224.5537 | 19.0867 | 24.9956 | 999.84 | 0.0073314 | 0.323 | 289.479 | 320.31
Populus deltoides | LI2B |  | RACiR 0 to 500 | 225.3119 | 19.356 | 24.9962 | 999.876 | 0.0073315 | 0.32303 | 291.092 | 322.205
Populus deltoides | LI2B |  | RACiR 0 to 500 | 226.3225 | 19.4972 | 24.9976 | 999.989 | 0.0073327 | 0.32305 | 292.565 | 323.828
Populus deltoides | LI2B |  | RACiR 0 to 500 | 227.7131 | 19.5835 | 24.996 | 999.859 | 0.0073349 | 0.32326 | 294.215 | 325.571
Populus deltoides | LI2B |  | RACiR 0 to 500 | 228.4554 | 19.8192 | 24.9958 | 999.853 | 0.0073362 | 0.32339 | 295.686 | 327.288
Populus deltoides | LI2B |  | RACiR 0 to 500 | 229.862 | 19.924 | 24.9965 | 999.815 | 0.0073482 | 0.32394 | 297.346 | 329.066
Populus deltoides | LI2B |  | RACiR 0 to 500 | 231.5264 | 19.9178 | 24.9945 | 999.861 | 0.0073676 | 0.32495 | 298.834 | 330.561
Populus deltoides | LI2B |  | RACiR 0 to 500 | 232.875 | 20.0767 | 24.9905 | 1000.13 | 0.0073889 | 0.32609 | 300.486 | 332.386
Populus deltoides | LI2B |  | RACiR 0 to 500 | 234.3289 | 20.1314 | 24.987 | 999.828 | 0.0074058 | 0.32696 | 301.974 | 333.947
Populus deltoides | LI2B |  | RACiR 0 to 500 | 236.0494 | 20.1264 | 24.9847 | 999.794 | 0.0074175 | 0.32748 | 303.619 | 335.596
Populus deltoides | LI2B |  | RACiR 0 to 500 | 237.0197 | 20.2682 | 24.9837 | 1000.21 | 0.0074158 | 0.32732 | 305.082 | 337.208
Populus deltoides | LI2B |  | RACiR 0 to 500 | 238.1711 | 20.3724 | 24.9881 | 1000.18 | 0.0074102 | 0.32671 | 306.706 | 338.949
Populus deltoides | LI2B |  | RACiR 0 to 500 | 239.3028 | 20.4394 | 24.9926 | 999.811 | 0.0074082 | 0.32629 | 308.159 | 340.471
Populus deltoides | LI2B |  | RACiR 0 to 500 | 240.0806 | 20.6731 | 24.9966 | 1000.14 | 0.0074049 | 0.32585 | 309.771 | 342.328
Populus deltoides | LI2B |  | RACiR 0 to 500 | 240.9382 | 20.8462 | 24.9982 | 1000.08 | 0.0074015 | 0.32555 | 311.25 | 343.99
Populus deltoides | LI2B |  | RACiR 0 to 500 | 242.3327 | 20.9134 | 24.9975 | 1000.12 | 0.0074004 | 0.32547 | 312.902 | 345.714
Populus deltoides | LI2B |  | RACiR 0 to 500 | 243.1826 | 21.1004 | 24.997 | 1000.2 | 0.0073979 | 0.32535 | 314.375 | 347.383
Populus deltoides | LI2B |  | RACiR 0 to 500 | 244.5233 | 21.1944 | 24.9956 | 999.863 | 0.0073974 | 0.32539 | 316.03 | 349.142
Populus deltoides | LI2B |  | RACiR 0 to 500 | 245.8698 | 21.22 | 24.9967 | 999.81 | 0.0073982 | 0.32537 | 317.492 | 350.634
Populus deltoides | LI2B |  | RACiR 0 to 500 | 246.8132 | 21.4332 | 24.9952 | 1000.1 | 0.0073994 | 0.32554 | 319.084 | 352.449
Populus deltoides | LI2B |  | RACiR 0 to 500 | 247.9725 | 21.5308 | 24.9925 | 999.776 | 0.0073992 | 0.32568 | 320.54 | 354.012
Populus deltoides | LI2B |  | RACiR 0 to 500 | 249.5387 | 21.5385 | 24.9925 | 999.867 | 0.0073992 | 0.32571 | 322.161 | 355.65
Populus deltoides | LI2B |  | RACiR 0 to 500 | 250.413 | 21.7339 | 24.9915 | 1000.26 | 0.0073998 | 0.32583 | 323.634 | 357.327
Populus deltoides | LI2B |  | RACiR 0 to 500 | 251.4823 | 21.9131 | 24.9912 | 1000.19 | 0.0074012 | 0.32596 | 325.257 | 359.139
Populus deltoides | LI2B |  | RACiR 0 to 500 | 252.7277 | 22.002 | 24.9892 | 999.86 | 0.0074036 | 0.32621 | 326.75 | 360.727
Populus deltoides | LI2B |  | RACiR 0 to 500 | 253.6488 | 22.2552 | 24.9846 | 999.793 | 0.0074051 | 0.32656 | 328.399 | 362.641
Populus deltoides | LI2B |  | RACiR 0 to 500 | 254.8081 | 22.3642 | 24.9838 | 999.98 | 0.007407 | 0.32673 | 329.884 | 364.246
Populus deltoides | LI2B |  | RACiR 0 to 500 | 256.5153 | 22.3583 | 24.9825 | 999.819 | 0.0074091 | 0.32694 | 331.566 | 365.928
Populus deltoides | LI2B |  | RACiR 0 to 500 | 257.7328 | 22.452 | 24.9819 | 999.849 | 0.0074105 | 0.32709 | 333.068 | 367.534
Populus deltoides | LI2B |  | RACiR 0 to 500 | 259.091 | 22.5295 | 24.986 | 999.897 | 0.0074134 | 0.32707 | 334.703 | 369.259
Populus deltoides | LI2B |  | RACiR 0 to 500 | 260.4285 | 22.5574 | 24.9899 | 999.842 | 0.0074163 | 0.32704 | 336.167 | 370.755
Populus deltoides | LI2B |  | RACiR 0 to 500 | 261.39 | 22.7597 | 24.9934 | 1000.12 | 0.0074173 | 0.32695 | 337.799 | 372.605
Populus deltoides | LI2B |  | RACiR 0 to 500 | 262.3059 | 22.929 | 24.9949 | 999.834 | 0.0074179 | 0.32694 | 339.262 | 374.247
Populus deltoides | LI2B |  | RACiR 0 to 500 | 263.7539 | 22.9944 | 24.9927 | 999.865 | 0.0074184 | 0.32708 | 340.913 | 375.969
Populus deltoides | LI2B |  | RACiR 0 to 500 | 264.7362 | 23.1779 | 24.9888 | 1000.14 | 0.0074186 | 0.32736 | 342.419 | 377.672
Populus deltoides | LI2B |  | RACiR 0 to 500 | 266.114 | 23.2693 | 24.9907 | 1000.2 | 0.007421 | 0.32743 | 344.096 | 379.449
Populus deltoides | LI2B |  | RACiR 0 to 500 | 267.6852 | 23.2449 | 24.9937 | 999.842 | 0.0074244 | 0.32748 | 345.619 | 380.952
Populus deltoides | LI2B |  | RACiR 0 to 500 | 268.8264 | 23.406 | 24.9981 | 999.869 | 0.0074278 | 0.3275 | 347.282 | 382.79
Populus deltoides | LI2B |  | RACiR 0 to 500 | 270.1345 | 23.4839 | 24.9988 | 999.945 | 0.0074349 | 0.32784 | 348.784 | 384.379
Populus deltoides | LI2B |  | RACiR 0 to 500 | 271.8819 | 23.4796 | 24.9962 | 999.816 | 0.0074427 | 0.32839 | 350.432 | 386.033
Populus deltoides | LI2B |  | RACiR 0 to 500 | 273.0667 | 23.5885 | 24.9959 | 999.87 | 0.0074479 | 0.32864 | 351.924 | 387.648
Populus deltoides | LI2B |  | RACiR 0 to 500 | 274.6047 | 23.646 | 24.9937 | 999.892 | 0.007454 | 0.32903 | 353.586 | 389.379
Populus deltoides | LI2B |  | RACiR 0 to 500 | 275.9674 | 23.67 | 24.9951 | 999.86 | 0.0074564 | 0.32908 | 355.044 | 390.869
Populus deltoides | LI2B |  | RACiR 0 to 500 | 276.8531 | 23.8873 | 24.9981 | 1000.21 | 0.0074562 | 0.3289 | 356.662 | 392.714
Populus deltoides | LI2B |  | RACiR 0 to 500 | 277.7747 | 24.0445 | 25.0029 | 999.819 | 0.0074585 | 0.32879 | 358.118 | 394.339
Populus deltoides | LI2B |  | RACiR 0 to 500 | 279.3766 | 24.0639 | 25.0038 | 999.876 | 0.0074629 | 0.32896 | 359.777 | 396.026
Populus deltoides | LI2B |  | RACiR 0 to 500 | 280.2872 | 24.2433 | 25.0069 | 1000.17 | 0.007467 | 0.329 | 361.255 | 397.693
Populus deltoides | LI2B |  | RACiR 0 to 500 | 281.8776 | 24.3032 | 25.0065 | 1000.19 | 0.0074808 | 0.32965 | 362.919 | 399.433
Populus deltoides | LI2B |  | RACiR 0 to 500 | 283.7522 | 24.262 | 25.0022 | 999.865 | 0.0075022 | 0.33081 | 364.443 | 400.928
Populus deltoides | LI2B |  | RACiR 0 to 500 | 284.9688 | 24.442 | 25.0035 | 999.909 | 0.0075222 | 0.33156 | 366.073 | 402.758
Populus deltoides | LI2B |  | RACiR 0 to 500 | 286.3655 | 24.5116 | 24.9999 | 1000.11 | 0.0075373 | 0.33226 | 367.556 | 404.329
Populus deltoides | LI2B |  | RACiR 0 to 500 | 288.0319 | 24.5151 | 24.9977 | 1000.21 | 0.0075475 | 0.3326 | 369.197 | 405.982
Populus deltoides | LI2B |  | RACiR 0 to 500 | 288.7656 | 24.7105 | 24.998 | 1000.17 | 0.0075406 | 0.33204 | 370.673 | 407.661
Populus deltoides | LI2B |  | RACiR 0 to 500 | 289.8383 | 24.853 | 24.9946 | 999.85 | 0.0075306 | 0.33151 | 372.329 | 409.469
Populus deltoides | LI2B |  | RACiR 0 to 500 | 291.039 | 24.9188 | 24.9934 | 999.881 | 0.0075258 | 0.33122 | 373.826 | 411.033
Populus deltoides | LI2B |  | RACiR 0 to 500 | 291.8896 | 25.1531 | 24.9927 | 1000.19 | 0.0075195 | 0.33083 | 375.505 | 412.954
Populus deltoides | LI2B |  | RACiR 0 to 500 | 293.1218 | 25.2258 | 24.9904 | 999.881 | 0.007515 | 0.33065 | 377.03 | 414.559
Populus deltoides | LI2B |  | RACiR 0 to 500 | 294.9319 | 25.1655 | 24.9896 | 999.897 | 0.007513 | 0.33055 | 378.72 | 416.196
Populus deltoides | LI2B |  | RACiR 0 to 500 | 295.998 | 25.2813 | 24.9896 | 999.87 | 0.0075098 | 0.33032 | 380.22 | 417.822
Populus deltoides | LI2B |  | RACiR 0 to 500 | 297.4874 | 25.3221 | 24.9884 | 999.819 | 0.007507 | 0.33023 | 381.889 | 419.542
Populus deltoides | LI2B |  | RACiR 0 to 500 | 298.9827 | 25.3182 | 24.9868 | 1000.19 | 0.0075075 | 0.33036 | 383.373 | 421.03
Populus deltoides | LI2B |  | RACiR 0 to 500 | 299.8967 | 25.5516 | 24.9836 | 999.892 | 0.0075058 | 0.33043 | 385.005 | 422.906
Populus deltoides | LI2B |  | RACiR 0 to 500 | 301.0755 | 25.6613 | 24.9803 | 1000.12 | 0.0075046 | 0.33057 | 386.507 | 424.529
Populus deltoides | LI2B |  | RACiR 0 to 500 | 302.5558 | 25.7169 | 24.9793 | 1000.19 | 0.0075061 | 0.33068 | 388.165 | 426.244
Populus deltoides | LI2B |  | RACiR 0 to 500 | 303.4916 | 25.9089 | 24.9758 | 1000.25 | 0.0075054 | 0.33083 | 389.67 | 427.95
Populus deltoides | LI2B |  | RACiR 0 to 500 | 304.9892 | 25.9681 | 24.9746 | 999.859 | 0.0075045 | 0.33084 | 391.379 | 429.729
Populus deltoides | LI2B |  | RACiR 0 to 500 | 306.6044 | 25.9406 | 24.9726 | 1000.24 | 0.0075049 | 0.33095 | 392.917 | 431.245
Populus deltoides | LI2B |  | RACiR 0 to 500 | 307.8327 | 26.0975 | 24.9675 | 999.895 | 0.007503 | 0.33114 | 394.602 | 433.102
Populus deltoides | LI2B |  | RACiR 0 to 500 | 309.2529 | 26.1404 | 24.9626 | 1000.18 | 0.0075039 | 0.33145 | 396.103 | 434.657
Populus deltoides | LI2B |  | RACiR 0 to 500 | 311.0727 | 26.1031 | 24.9571 | 999.913 | 0.0075041 | 0.33175 | 397.773 | 436.296
Populus deltoides | LI2B |  | RACiR 0 to 500 | 312.1727 | 26.2299 | 24.954 | 999.905 | 0.0075031 | 0.33188 | 399.249 | 437.908
Populus deltoides | LI2B |  | RACiR 0 to 500 | 313.6876 | 26.2664 | 24.9537 | 999.846 | 0.0075036 | 0.33193 | 400.896 | 439.604
Populus deltoides | LI2B |  | RACiR 0 to 500 | 315.1356 | 26.259 | 24.9563 | 1000.18 | 0.0075042 | 0.33184 | 402.379 | 441.086
Populus deltoides | LI2B |  | RACiR 0 to 500 | 316.0932 | 26.4799 | 24.9565 | 999.898 | 0.0075055 | 0.33194 | 404.006 | 442.943
Populus deltoides | LI2B |  | RACiR 0 to 500 | 317.2148 | 26.6042 | 24.9548 | 1000.14 | 0.0075064 | 0.3321 | 405.492 | 444.564
Populus deltoides | LI2B |  | RACiR 0 to 500 | 318.6513 | 26.6762 | 24.9539 | 1000.26 | 0.0075074 | 0.33223 | 407.146 | 446.295
Populus deltoides | LI2B |  | RACiR 0 to 500 | 319.6382 | 26.8349 | 24.9533 | 999.919 | 0.0075059 | 0.3322 | 408.645 | 447.964
Populus deltoides | LI2B |  | RACiR 0 to 500 | 321.0895 | 26.9019 | 24.9521 | 1000.22 | 0.0075057 | 0.33226 | 410.316 | 449.712
Populus deltoides | LI2B |  | RACiR 0 to 500 | 322.5408 | 26.899 | 24.955 | 999.871 | 0.0075059 | 0.33211 | 411.835 | 451.232
Populus deltoides | LI2B |  | RACiR 0 to 500 | 323.7154 | 27.0406 | 24.9556 | 1000.25 | 0.0075049 | 0.33202 | 413.491 | 453.044
Populus deltoides | LI2B |  | RACiR 0 to 500 | 324.9706 | 27.1219 | 24.953 | 999.885 | 0.0075053 | 0.33218 | 414.979 | 454.622
Populus deltoides | LI2B |  | RACiR 0 to 500 | 326.7032 | 27.0846 | 24.9539 | 999.894 | 0.0075058 | 0.33214 | 416.647 | 456.258
Populus deltoides | LI2B |  | RACiR 0 to 500 | 328.0298 | 27.1447 | 24.9514 | 1000.14 | 0.0075061 | 0.33232 | 418.138 | 457.82
Populus deltoides | LI2B |  | RACiR 0 to 500 | 329.4366 | 27.2181 | 24.9496 | 999.882 | 0.0075053 | 0.33241 | 419.775 | 459.544
Populus deltoides | LI2B |  | RACiR 0 to 500 | 330.7594 | 27.2552 | 24.9502 | 999.848 | 0.007506 | 0.33242 | 421.241 | 461.05
Populus deltoides | LI2B |  | RACiR 0 to 500 | 331.7375 | 27.4798 | 24.9492 | 999.885 | 0.0075079 | 0.33262 | 422.869 | 462.917
Populus deltoides | LI2B |  | RACiR 0 to 500 | 332.7101 | 27.6271 | 24.9509 | 999.986 | 0.0075074 | 0.33251 | 424.343 | 464.548
Populus deltoides | LI2B |  | RACiR 0 to 500 | 334.2617 | 27.6673 | 24.9512 | 999.895 | 0.0075101 | 0.33267 | 426.01 | 466.261
Populus deltoides | LI2B |  | RACiR 0 to 500 | 335.3612 | 27.8085 | 24.9505 | 1000.3 | 0.0075115 | 0.33282 | 427.523 | 467.926
Populus deltoides | LI2C |  | RACiR 800 to 300 | 570.9281 | 37.5276 | 24.9923 | 999.712 | 0.0068237 | 0.2939 | 713.526 | 745.619
Populus deltoides | LI2C |  | RACiR 800 to 300 | 567.6368 | 37.5438 | 24.9956 | 1000.09 | 0.006826 | 0.29385 | 710.245 | 742.344
Populus deltoides | LI2C |  | RACiR 800 to 300 | 565.2731 | 37.2824 | 25.0026 | 999.867 | 0.0068278 | 0.29363 | 707.035 | 738.867
Populus deltoides | LI2C |  | RACiR 800 to 300 | 562.1307 | 37.2247 | 25.0084 | 1000.06 | 0.0068279 | 0.29336 | 703.753 | 735.514
Populus deltoides | LI2C |  | RACiR 800 to 300 | 558.8958 | 37.2278 | 25.0112 | 1000.06 | 0.0068293 | 0.29329 | 700.488 | 732.238
Populus deltoides | LI2C |  | RACiR 800 to 300 | 556.3509 | 37.0808 | 25.0051 | 1000.15 | 0.0068308 | 0.29366 | 697.204 | 728.799
Populus deltoides | LI2C |  | RACiR 800 to 300 | 552.9836 | 37.1787 | 25.0002 | 1000.03 | 0.0068325 | 0.29396 | 693.95 | 725.632
Populus deltoides | LI2C |  | RACiR 800 to 300 | 549.667 | 37.2154 | 25.0042 | 999.747 | 0.0068331 | 0.29378 | 690.768 | 722.474
Populus deltoides | LI2C |  | RACiR 800 to 300 | 547.4234 | 36.9342 | 25.0054 | 1000.1 | 0.0068329 | 0.29369 | 687.547 | 718.963
Populus deltoides | LI2C |  | RACiR 800 to 300 | 544.5508 | 36.8922 | 24.9965 | 999.983 | 0.0068337 | 0.29413 | 684.255 | 715.617
Populus deltoides | LI2C |  | RACiR 800 to 300 | 541.306 | 36.9555 | 24.9883 | 1000.08 | 0.0068346 | 0.29456 | 680.948 | 712.359
Populus deltoides | LI2C |  | RACiR 800 to 300 | 538.559 | 36.8051 | 24.9932 | 1000.06 | 0.006835 | 0.29436 | 677.715 | 708.971
Populus deltoides | LI2C |  | RACiR 800 to 300 | 535.4233 | 36.7949 | 24.9965 | 1000.11 | 0.006836 | 0.29422 | 674.539 | 705.771
Populus deltoides | LI2C |  | RACiR 800 to 300 | 532.6446 | 36.6621 | 24.999 | 999.725 | 0.0068348 | 0.29402 | 671.33 | 702.416
Populus deltoides | LI2C |  | RACiR 800 to 300 | 530.0528 | 36.4392 | 25.0036 | 1000.05 | 0.0068337 | 0.29372 | 668.051 | 698.904
Populus deltoides | LI2C |  | RACiR 800 to 300 | 526.649 | 36.5217 | 24.9977 | 1000.06 | 0.0068342 | 0.29399 | 664.725 | 695.647
Populus deltoides | LI2C |  | RACiR 800 to 300 | 523.0671 | 36.6623 | 24.9942 | 1000.07 | 0.0068347 | 0.29416 | 661.46 | 692.51
Populus deltoides | LI2C |  | RACiR 800 to 300 | 520.6546 | 36.3853 | 25.0005 | 1000.07 | 0.0068324 | 0.29373 | 658.24 | 689.003
Populus deltoides | LI2C |  | RACiR 800 to 300 | 518.0843 | 36.1837 | 25.0057 | 1000.07 | 0.006833 | 0.29349 | 655.026 | 685.575
Populus deltoides | LI2C |  | RACiR 800 to 300 | 514.6816 | 36.2405 | 25.0089 | 1000.07 | 0.0068348 | 0.29344 | 651.767 | 682.361
Populus deltoides | LI2C |  | RACiR 800 to 300 | 511.6498 | 36.1534 | 25.0124 | 1000.02 | 0.0068342 | 0.29322 | 648.466 | 678.963
Populus deltoides | LI2C |  | RACiR 800 to 300 | 508.0657 | 36.2917 | 25.0097 | 999.746 | 0.0068357 | 0.29342 | 645.184 | 675.806
Populus deltoides | LI2C |  | RACiR 800 to 300 | 504.8727 | 36.382 | 25.0024 | 1000.06 | 0.0068398 | 0.29397 | 641.975 | 672.676
Populus deltoides | LI2C |  | RACiR 800 to 300 | 502.6864 | 36.1097 | 25.0015 | 1000.03 | 0.0068391 | 0.29397 | 638.8 | 669.222
Populus deltoides | LI2C |  | RACiR 800 to 300 | 499.7891 | 35.9928 | 25.005 | 1000.12 | 0.0068379 | 0.29372 | 635.548 | 665.839
Populus deltoides | LI2C |  | RACiR 800 to 300 | 496.59 | 36.0127 | 25.0034 | 1000.08 | 0.0068386 | 0.29383 | 632.292 | 662.591
Populus deltoides | LI2C |  | RACiR 800 to 300 | 493.8022 | 35.892 | 25.0039 | 1000.04 | 0.0068375 | 0.29373 | 629.065 | 659.233
Populus deltoides | LI2C |  | RACiR 800 to 300 | 490.1538 | 35.988 | 25.01 | 999.695 | 0.0068367 | 0.29337 | 625.826 | 656.076
Populus deltoides | LI2C |  | RACiR 800 to 300 | 486.8267 | 35.9831 | 25.0184 | 1000.09 | 0.0068362 | 0.29291 | 622.619 | 652.852
Populus deltoides | LI2C |  | RACiR 800 to 300 | 484.6818 | 35.6802 | 25.0204 | 1000.06 | 0.0068357 | 0.29276 | 619.445 | 649.366
Populus deltoides | LI2C |  | RACiR 800 to 300 | 481.7309 | 35.661 | 25.0152 | 1000.08 | 0.006838 | 0.29312 | 616.198 | 646.088
Populus deltoides | LI2C |  | RACiR 800 to 300 | 478.5368 | 35.6723 | 25.0168 | 999.79 | 0.0068393 | 0.29312 | 612.967 | 642.855
Populus deltoides | LI2C |  | RACiR 800 to 300 | 476.0647 | 35.3498 | 25.0186 | 1000.06 | 0.0068192 | 0.29215 | 609.734 | 639.279
Populus deltoides | LI2C |  | RACiR 800 to 300 | 472.8764 | 35.2235 | 25.0213 | 999.805 | 0.0067981 | 0.29117 | 606.447 | 635.842
Populus deltoides | LI2C |  | RACiR 800 to 300 | 469.3021 | 35.3355 | 25.025 | 1000.02 | 0.006801 | 0.29133 | 603.111 | 632.606
Populus deltoides | LI2C |  | RACiR 800 to 300 | 466.6017 | 35.1806 | 25.0267 | 1000.14 | 0.0067976 | 0.29128 | 599.826 | 629.155
Populus deltoides | LI2C |  | RACiR 800 to 300 | 462.3968 | 35.2557 | 25.0287 | 999.813 | 0.0067532 | 0.28941 | 596.572 | 625.936
Populus deltoides | LI2C |  | RACiR 800 to 300 | 458.4137 | 35.3869 | 25.0288 | 1000.08 | 0.0067224 | 0.28847 | 593.337 | 622.802
Populus deltoides | LI2C |  | RACiR 800 to 300 | 456.5712 | 35.1528 | 25.0308 | 1000.04 | 0.0067397 | 0.28966 | 590.109 | 619.342
Populus deltoides | LI2C |  | RACiR 800 to 300 | 454.058 | 35.0837 | 25.0289 | 999.79 | 0.0067534 | 0.29062 | 586.871 | 616.029
Populus deltoides | LI2C |  | RACiR 800 to 300 | 451.1641 | 35.1069 | 25.0243 | 1000.05 | 0.0067622 | 0.29141 | 583.641 | 612.816
Populus deltoides | LI2C |  | RACiR 800 to 300 | 448.5823 | 34.9788 | 25.0243 | 1000.09 | 0.0067661 | 0.29169 | 580.433 | 609.473
Populus deltoides | LI2C |  | RACiR 800 to 300 | 445.632 | 34.9041 | 25.0283 | 999.785 | 0.0067672 | 0.29157 | 577.212 | 606.165
Populus deltoides | LI2C |  | RACiR 800 to 300 | 442.8563 | 34.746 | 25.0363 | 1000.14 | 0.0067689 | 0.29128 | 573.959 | 602.744
Populus deltoides | LI2C |  | RACiR 800 to 300 | 440.1938 | 34.5648 | 25.0412 | 1000.17 | 0.0067707 | 0.29114 | 570.668 | 599.263
Populus deltoides | LI2C |  | RACiR 800 to 300 | 436.9631 | 34.6062 | 25.038 | 1000.09 | 0.0067707 | 0.29128 | 567.446 | 596.071
Populus deltoides | LI2C |  | RACiR 800 to 300 | 433.6466 | 34.6609 | 25.0411 | 1000.11 | 0.0067721 | 0.29121 | 564.275 | 592.942
Populus deltoides | LI2C |  | RACiR 800 to 300 | 431.2515 | 34.4885 | 25.0378 | 1000.09 | 0.0067736 | 0.29144 | 561.12 | 589.607
Populus deltoides | LI2C |  | RACiR 800 to 300 | 428.5641 | 34.3897 | 25.0313 | 999.774 | 0.0067722 | 0.29166 | 557.928 | 586.305
Populus deltoides | LI2C |  | RACiR 800 to 300 | 425.5589 | 34.3079 | 25.0375 | 999.789 | 0.0067714 | 0.2913 | 554.727 | 583.008
Populus deltoides | LI2C |  | RACiR 800 to 300 | 423.1044 | 34.0957 | 25.0391 | 1000.08 | 0.0067709 | 0.29116 | 551.54 | 579.597
Populus deltoides | LI2C |  | RACiR 800 to 300 | 419.867 | 34.0976 | 25.0417 | 1000.11 | 0.0067687 | 0.29092 | 548.338 | 576.387
Populus deltoides | LI2C |  | RACiR 800 to 300 | 416.917 | 34.0465 | 25.0391 | 999.793 | 0.0067662 | 0.29086 | 545.162 | 573.145
Populus deltoides | LI2C |  | RACiR 800 to 300 | 415.1205 | 33.6553 | 25.038 | 1000.12 | 0.0067635 | 0.29075 | 542.005 | 569.588
Populus deltoides | LI2C |  | RACiR 800 to 300 | 412.4597 | 33.5298 | 25.0364 | 999.821 | 0.0067655 | 0.29092 | 538.773 | 566.217
Populus deltoides | LI2C |  | RACiR 800 to 300 | 409.2191 | 33.5823 | 25.0277 | 999.862 | 0.006765 | 0.29127 | 535.487 | 562.971
Populus deltoides | LI2C |  | RACiR 800 to 300 | 406.4645 | 33.4208 | 25.0304 | 999.958 | 0.0067603 | 0.29083 | 532.289 | 559.602
Populus deltoides | LI2C |  | RACiR 800 to 300 | 403.0235 | 33.4724 | 25.0378 | 999.765 | 0.0067596 | 0.29039 | 529.138 | 556.488
Populus deltoides | LI2C |  | RACiR 800 to 300 | 400.0435 | 33.4445 | 25.0355 | 999.84 | 0.0067592 | 0.29044 | 525.97 | 553.279
Populus deltoides | LI2C |  | RACiR 800 to 300 | 398.2712 | 33.093 | 25.0295 | 999.86 | 0.006759 | 0.29071 | 522.814 | 549.765
Populus deltoides | LI2C |  | RACiR 800 to 300 | 395.7107 | 32.9577 | 25.0252 | 999.833 | 0.0067588 | 0.29091 | 519.635 | 546.439
Populus deltoides | LI2C |  | RACiR 800 to 300 | 392.6927 | 32.9349 | 25.022 | 999.837 | 0.0067569 | 0.29093 | 516.456 | 543.225
Populus deltoides | LI2C |  | RACiR 800 to 300 | 390.551 | 32.606 | 25.0264 | 1000.02 | 0.0067544 | 0.29056 | 513.274 | 539.703
Populus deltoides | LI2C |  | RACiR 800 to 300 | 387.7599 | 32.469 | 25.0286 | 999.82 | 0.0067554 | 0.29045 | 509.989 | 536.268
Populus deltoides | LI2C |  | RACiR 800 to 300 | 384.5783 | 32.4987 | 25.0234 | 1000.09 | 0.006756 | 0.29071 | 506.727 | 533.025
Populus deltoides | LI2C |  | RACiR 800 to 300 | 382.0372 | 32.3228 | 25.0221 | 1000.1 | 0.0067534 | 0.29061 | 503.55 | 529.663
Populus deltoides | LI2C |  | RACiR 800 to 300 | 379.0157 | 32.3013 | 25.0188 | 999.827 | 0.0067515 | 0.2906 | 500.383 | 526.46
Populus deltoides | LI2C |  | RACiR 800 to 300 | 376.1632 | 32.2767 | 25.0111 | 999.881 | 0.0067497 | 0.29086 | 497.267 | 523.307
Populus deltoides | LI2C |  | RACiR 800 to 300 | 374.2904 | 31.9458 | 25.0059 | 999.976 | 0.0067471 | 0.29098 | 494.144 | 519.845
Populus deltoides | LI2C |  | RACiR 800 to 300 | 371.6944 | 31.7873 | 25.0077 | 1000.21 | 0.0067473 | 0.29084 | 490.994 | 516.523
Populus deltoides | LI2C |  | RACiR 800 to 300 | 368.7067 | 31.7299 | 25.0125 | 1000.15 | 0.0067475 | 0.29059 | 487.841 | 513.301
Populus deltoides | LI2C |  | RACiR 800 to 300 | 366.4826 | 31.499 | 25.0041 | 1000.16 | 0.0067446 | 0.29079 | 484.677 | 509.896
Populus deltoides | LI2C |  | RACiR 800 to 300 | 363.6818 | 31.3754 | 25.0035 | 999.859 | 0.0067391 | 0.29046 | 481.505 | 506.587
Populus deltoides | LI2C |  | RACiR 800 to 300 | 360.9858 | 31.1839 | 25.0096 | 1000.2 | 0.0067345 | 0.28985 | 478.322 | 503.2
Populus deltoides | LI2C |  | RACiR 800 to 300 | 358.9142 | 30.8793 | 25.0078 | 1000.05 | 0.0067333 | 0.28983 | 475.144 | 499.706
Populus deltoides | LI2C |  | RACiR 800 to 300 | 356.0716 | 30.8369 | 25.0002 | 1000.19 | 0.0067329 | 0.29017 | 471.951 | 496.459
Populus deltoides | LI2C |  | RACiR 800 to 300 | 352.4618 | 30.9038 | 24.9947 | 1000.17 | 0.0067102 | 0.28941 | 468.766 | 493.319
Populus deltoides | LI2C |  | RACiR 800 to 300 | 350.3726 | 30.5285 | 24.9956 | 999.91 | 0.0066875 | 0.28849 | 465.654 | 489.812
Populus deltoides | LI2C |  | RACiR 800 to 300 | 348.2108 | 30.304 | 24.9922 | 1000.19 | 0.0066897 | 0.28891 | 462.489 | 486.411
Populus deltoides | LI2C |  | RACiR 800 to 300 | 345.3451 | 30.2467 | 24.9962 | 999.907 | 0.0066945 | 0.28903 | 459.313 | 483.17
Populus deltoides | LI2C |  | RACiR 800 to 300 | 342.9084 | 30.0656 | 24.9983 | 1000 | 0.0066973 | 0.28911 | 456.157 | 479.826
Populus deltoides | LI2C |  | RACiR 800 to 300 | 339.7849 | 30.0963 | 24.998 | 1000.21 | 0.0067002 | 0.28928 | 453.004 | 476.693
Populus deltoides | LI2C |  | RACiR 800 to 300 | 336.9201 | 29.9892 | 25.0121 | 999.848 | 0.0067016 | 0.28873 | 449.916 | 473.486
Populus deltoides | LI2C |  | RACiR 800 to 300 | 335.1912 | 29.5985 | 25.0148 | 999.845 | 0.0067004 | 0.28856 | 446.843 | 470.015
Populus deltoides | LI2C |  | RACiR 800 to 300 | 332.9951 | 29.3677 | 25.013 | 1000.24 | 0.0067019 | 0.28872 | 443.726 | 466.656
Populus deltoides | LI2C |  | RACiR 800 to 300 | 329.4702 | 29.3306 | 25.0163 | 999.895 | 0.0066683 | 0.28714 | 440.559 | 463.425
Populus deltoides | LI2C |  | RACiR 800 to 300 | 326.4867 | 29.1423 | 25.0134 | 1000.02 | 0.0066225 | 0.28551 | 437.419 | 460.068
Populus deltoides | LI2C |  | RACiR 800 to 300 | 323.5716 | 29.1064 | 25.0188 | 1000.21 | 0.0066191 | 0.2855 | 434.305 | 456.905
Populus deltoides | LI2C |  | RACiR 800 to 300 | 321.1828 | 29.0095 | 25.0265 | 1000.14 | 0.0066347 | 0.28618 | 431.275 | 453.775
Populus deltoides | LI2C |  | RACiR 800 to 300 | 319.9416 | 28.5748 | 25.0267 | 999.819 | 0.0066437 | 0.28679 | 428.244 | 450.306
Populus deltoides | LI2C |  | RACiR 800 to 300 | 317.9654 | 28.3293 | 25.0229 | 1000.19 | 0.0066515 | 0.28745 | 425.118 | 446.925
Populus deltoides | LI2C |  | RACiR 800 to 300 | 314.9184 | 28.3255 | 25.026 | 1000.14 | 0.0066554 | 0.28755 | 421.951 | 443.745
Populus deltoides | LI2C |  | RACiR 800 to 300 | 312.5398 | 28.1258 | 25.0292 | 1000.07 | 0.0066556 | 0.28747 | 418.845 | 440.433
Populus deltoides | LI2C |  | RACiR 800 to 300 | 310.1778 | 27.9147 | 25.0341 | 1000.18 | 0.0066556 | 0.28723 | 415.773 | 437.139
Populus deltoides | LI2C |  | RACiR 800 to 300 | 307.9815 | 27.6421 | 25.0399 | 1000.2 | 0.006656 | 0.28694 | 412.675 | 433.755
Populus deltoides | LI2C |  | RACiR 800 to 300 | 306.0242 | 27.3141 | 25.0361 | 1000.14 | 0.0066531 | 0.28693 | 409.514 | 430.258
Populus deltoides | LI2C |  | RACiR 800 to 300 | 302.899 | 27.3057 | 25.0372 | 1000.14 | 0.0066536 | 0.28683 | 406.323 | 427.048
Populus deltoides | LI2C |  | RACiR 800 to 300 | 299.8153 | 27.2711 | 25.0451 | 1000.1 | 0.0066534 | 0.28641 | 403.198 | 423.876
Populus deltoides | LI2C |  | RACiR 800 to 300 | 297.9748 | 26.8944 | 25.0457 | 1000.01 | 0.0066473 | 0.28605 | 400.107 | 420.4
Populus deltoides | LI2C |  | RACiR 800 to 300 | 295.6175 | 26.6889 | 25.0423 | 1000.13 | 0.0066454 | 0.28608 | 396.954 | 417.03
Populus deltoides | LI2C |  | RACiR 800 to 300 | 292.9966 | 26.584 | 25.0362 | 999.871 | 0.0066458 | 0.28636 | 393.8 | 413.757
Populus deltoides | LI2C |  | RACiR 800 to 300 | 290.6181 | 26.4036 | 25.0297 | 1000.18 | 0.0066439 | 0.28654 | 390.662 | 410.43
Populus deltoides | LI2C |  | RACiR 800 to 300 | 288.0538 | 26.2457 | 25.0326 | 999.891 | 0.0066431 | 0.28631 | 387.56 | 407.159
Populus deltoides | LI2C |  | RACiR 800 to 300 | 285.8778 | 25.978 | 25.0332 | 999.838 | 0.0066415 | 0.28617 | 384.434 | 403.754
Populus deltoides | LI2C |  | RACiR 800 to 300 | 283.8468 | 25.694 | 25.0265 | 1000.04 | 0.0066391 | 0.28638 | 381.277 | 400.303
Populus deltoides | LI2C |  | RACiR 800 to 300 | 280.9587 | 25.6211 | 25.0269 | 1000.11 | 0.0066361 | 0.2862 | 378.123 | 397.064
Populus deltoides | LI2C |  | RACiR 800 to 300 | 277.9759 | 25.5824 | 25.0292 | 1000.14 | 0.0066357 | 0.286 | 375 | 393.889
Populus deltoides | LI2C |  | RACiR 800 to 300 | 275.6233 | 25.368 | 25.0307 | 1000.18 | 0.0066335 | 0.28573 | 371.923 | 390.589
Populus deltoides | LI2C |  | RACiR 800 to 300 | 273.3589 | 25.1671 | 25.0245 | 999.871 | 0.0066308 | 0.28579 | 368.869 | 387.323
Populus deltoides | LI2C |  | RACiR 800 to 300 | 271.2965 | 24.89 | 25.021 | 999.807 | 0.0066273 | 0.28575 | 365.79 | 383.957
Populus deltoides | LI2C |  | RACiR 800 to 300 | 269.0872 | 24.5945 | 25.0303 | 1000.01 | 0.0066254 | 0.2852 | 362.662 | 380.524
Populus deltoides | LI2C |  | RACiR 800 to 300 | 266.3505 | 24.5195 | 25.0222 | 1000.15 | 0.0066219 | 0.28532 | 359.549 | 377.323
Populus deltoides | LI2C |  | RACiR 800 to 300 | 263.716 | 24.4149 | 25.0158 | 999.815 | 0.0066181 | 0.28536 | 356.466 | 374.123
Populus deltoides | LI2C |  | RACiR 800 to 300 | 261.8622 | 24.0299 | 25.0255 | 999.834 | 0.006616 | 0.28479 | 353.385 | 370.65
Populus deltoides | LI2C |  | RACiR 800 to 300 | 259.6488 | 23.7794 | 25.0242 | 1000.17 | 0.0066137 | 0.28469 | 350.26 | 367.262
Populus deltoides | LI2C |  | RACiR 800 to 300 | 256.7884 | 23.734 | 25.0152 | 1000.18 | 0.0066112 | 0.28491 | 347.097 | 364.04
Populus deltoides | LI2C |  | RACiR 800 to 300 | 254.4098 | 23.5436 | 25.0107 | 1000.12 | 0.0066088 | 0.28497 | 343.967 | 360.712
Populus deltoides | LI2C |  | RACiR 800 to 300 | 251.8748 | 23.423 | 25.0049 | 1000.12 | 0.0066081 | 0.28515 | 340.885 | 357.497
Populus deltoides | LI2C |  | RACiR 800 to 300 | 249.2491 | 23.305 | 25.0027 | 999.845 | 0.0066043 | 0.285 | 337.823 | 354.305
Populus deltoides | LI2C |  | RACiR 800 to 300 | 247.5979 | 22.9047 | 25.0037 | 999.782 | 0.0066009 | 0.28476 | 334.782 | 350.856
Populus deltoides | LI2C |  | RACiR 800 to 300 | 245.3337 | 22.7045 | 24.9964 | 999.847 | 0.0065996 | 0.28498 | 331.687 | 347.549
Populus deltoides | LI2C |  | RACiR 800 to 300 | 242.5771 | 22.6212 | 24.9914 | 999.788 | 0.0065958 | 0.28493 | 328.579 | 344.346
Populus deltoides | LI2C |  | RACiR 800 to 300 | 240.3354 | 22.3624 | 24.9971 | 999.999 | 0.0065926 | 0.28447 | 325.503 | 341.002
Populus deltoides | LI2C |  | RACiR 800 to 300 | 237.9268 | 22.1593 | 25.0024 | 999.82 | 0.0065905 | 0.28406 | 322.437 | 337.721
Populus deltoides | LI2C |  | RACiR 800 to 300 | 235.7976 | 21.9072 | 25.0004 | 999.728 | 0.006591 | 0.28415 | 319.332 | 334.354
Populus deltoides | LI2C |  | RACiR 800 to 300 | 234.0953 | 21.5526 | 24.9954 | 999.821 | 0.0065936 | 0.28455 | 316.215 | 330.875
Populus deltoides | LI2C |  | RACiR 800 to 300 | 231.5381 | 21.4565 | 24.9865 | 999.813 | 0.0065946 | 0.28501 | 313.123 | 327.676
Populus deltoides | LI2C |  | RACiR 800 to 300 | 228.9345 | 21.3403 | 24.9866 | 1000.12 | 0.006595 | 0.28499 | 310.051 | 324.478
Populus deltoides | LI2C |  | RACiR 800 to 300 | 227.1719 | 20.949 | 24.9913 | 1000.14 | 0.0065933 | 0.28463 | 306.959 | 320.986
Populus deltoides | LI2C |  | RACiR 800 to 300 | 224.7201 | 20.7524 | 24.993 | 1000.17 | 0.0065915 | 0.28441 | 303.81 | 317.631
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 354.6951 | 26.6403 | 25.0701 | 1000.17 | 0.006789 | 0.30598 | 450.695 | 471.684
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 351.9998 | 26.5643 | 25.1268 | 999.806 | 0.0068429 | 0.30661 | 447.562 | 468.483
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 349.2677 | 26.4148 | 25.1988 | 1000.05 | 0.0068936 | 0.30625 | 444.453 | 465.235
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 346.3797 | 26.3215 | 25.2608 | 1000.06 | 0.006936 | 0.30583 | 441.38 | 462.071
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 343.443 | 26.2301 | 25.3091 | 1000.08 | 0.0069631 | 0.30513 | 438.328 | 458.925
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 341.4519 | 25.8795 | 25.3396 | 1000.14 | 0.0069819 | 0.30479 | 435.256 | 455.499
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 339.1212 | 25.6955 | 25.351 | 999.764 | 0.0069991 | 0.30531 | 432.124 | 452.174
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 336.0993 | 25.6907 | 25.3614 | 1000.06 | 0.0070121 | 0.3056 | 428.942 | 448.976
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 333.613 | 25.5148 | 25.3633 | 999.753 | 0.0070155 | 0.30574 | 425.779 | 445.624
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 330.9229 | 25.3672 | 25.3641 | 1000.01 | 0.007011 | 0.30541 | 422.631 | 442.312
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 328.5855 | 25.1469 | 25.3546 | 999.717 | 0.0070056 | 0.30549 | 419.479 | 438.921
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 326.3262 | 24.892 | 25.3436 | 1000.03 | 0.0069983 | 0.30553 | 416.304 | 435.474
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 323.2529 | 24.8684 | 25.3321 | 999.697 | 0.006988 | 0.30539 | 413.109 | 432.235
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 320.4927 | 24.7958 | 25.3099 | 1000.01 | 0.006975 | 0.30567 | 409.947 | 428.979
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 318.399 | 24.4903 | 25.2877 | 1000.02 | 0.0069567 | 0.30556 | 406.804 | 425.51
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 315.6803 | 24.325 | 25.271 | 1000.02 | 0.0069366 | 0.30503 | 403.605 | 422.121
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 312.464 | 24.3322 | 25.249 | 1000.05 | 0.0069135 | 0.30461 | 400.435 | 418.933
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 309.9092 | 24.1674 | 25.2163 | 1000.04 | 0.0068857 | 0.3044 | 397.301 | 415.611
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 307.023 | 24.1247 | 25.1841 | 1000.05 | 0.0068618 | 0.30444 | 394.167 | 412.408
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 304.0959 | 24.0188 | 25.1642 | 1000.03 | 0.0068322 | 0.30348 | 391.066 | 409.174
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 302.0555 | 23.6593 | 25.1358 | 1000.01 | 0.0067988 | 0.30269 | 387.968 | 405.692
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 299.5798 | 23.4828 | 25.0983 | 1000.1 | 0.0067684 | 0.30251 | 384.862 | 402.383
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 296.4271 | 23.4399 | 25.072 | 1000.09 | 0.0067344 | 0.30148 | 381.75 | 399.198
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 293.7077 | 23.2286 | 25.0483 | 1000.11 | 0.0066962 | 0.30008 | 378.617 | 395.827
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 291.0101 | 23.0408 | 25.0149 | 1000.01 | 0.0066549 | 0.29899 | 375.489 | 392.482
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 288.456 | 22.8436 | 24.9749 | 999.997 | 0.006614 | 0.29828 | 372.367 | 389.134
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 286.4075 | 22.5824 | 24.9446 | 999.705 | 0.0065989 | 0.29881 | 369.222 | 385.71
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 284.1535 | 22.5478 | 24.9212 | 1000.03 | 0.0066239 | 0.30169 | 366.029 | 382.475
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 282.0844 | 22.545 | 24.8943 | 1000.06 | 0.0066646 | 0.30577 | 362.844 | 379.287
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 281.1756 | 22.1663 | 24.8833 | 1000.03 | 0.0067119 | 0.30958 | 359.725 | 375.795
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 279.9274 | 21.8929 | 24.8798 | 1000.01 | 0.0067696 | 0.31374 | 356.574 | 372.375
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 277.641 | 21.8707 | 24.8702 | 1000.03 | 0.0068115 | 0.31719 | 353.368 | 369.146
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 275.2346 | 21.7488 | 24.8613 | 1000.01 | 0.0068244 | 0.31859 | 350.199 | 365.847
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 272.3781 | 21.6737 | 24.8636 | 1000.04 | 0.0068257 | 0.31849 | 347.068 | 362.625
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 269.5531 | 21.5416 | 24.8731 | 1000.02 | 0.006818 | 0.31738 | 344.018 | 359.426
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 267.5203 | 21.1968 | 24.8697 | 1000.06 | 0.0068082 | 0.31684 | 340.958 | 356.006
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 264.8297 | 20.9961 | 24.8791 | 1000.02 | 0.0068003 | 0.31569 | 337.829 | 352.657
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 261.6282 | 20.9183 | 24.8977 | 1000.05 | 0.0067898 | 0.31392 | 334.712 | 349.444
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 259.1077 | 20.6846 | 24.9046 | 1000.01 | 0.0067809 | 0.31295 | 331.604 | 346.087
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 256.3913 | 20.5533 | 24.9054 | 999.988 | 0.0067751 | 0.31245 | 328.515 | 342.85
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 253.8977 | 20.34 | 24.9094 | 999.988 | 0.0067694 | 0.31179 | 325.423 | 339.528
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 251.7545 | 19.9901 | 24.9193 | 999.995 | 0.006763 | 0.31082 | 322.314 | 336.055
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 248.9818 | 19.879 | 24.9214 | 999.941 | 0.0067611 | 0.31056 | 319.176 | 332.79
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 245.9537 | 19.8419 | 24.9236 | 1000 | 0.0067577 | 0.31015 | 316.051 | 329.611
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 243.9937 | 19.4938 | 24.9233 | 1000 | 0.0067534 | 0.30985 | 312.977 | 326.176
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 241.6031 | 19.2576 | 24.9302 | 1000.01 | 0.0067537 | 0.30947 | 309.849 | 322.797
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 238.6892 | 19.1889 | 24.9348 | 1000.01 | 0.0067529 | 0.30917 | 306.711 | 319.574
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 236.5534 | 18.9281 | 24.926 | 999.995 | 0.0067504 | 0.30943 | 303.616 | 316.205
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 233.7357 | 18.8532 | 24.9238 | 999.989 | 0.0067505 | 0.30948 | 300.48 | 312.979
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 231.0092 | 18.7424 | 24.9272 | 1000.01 | 0.0067501 | 0.30926 | 297.38 | 309.753
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 229.0039 | 18.3884 | 24.9353 | 999.939 | 0.0067459 | 0.30857 | 294.317 | 306.323
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 226.3528 | 18.223 | 24.9471 | 1000.01 | 0.0067438 | 0.30777 | 291.234 | 303.059
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 223.352 | 18.2061 | 24.9454 | 1000.02 | 0.006743 | 0.30777 | 288.11 | 299.903
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 221.2095 | 17.9459 | 24.9405 | 1000.03 | 0.0067422 | 0.30795 | 285.027 | 296.546
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 218.9152 | 17.7024 | 24.9491 | 1000.04 | 0.006743 | 0.30751 | 281.974 | 293.235
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 216.6186 | 17.4525 | 24.9568 | 999.974 | 0.0067445 | 0.30719 | 278.872 | 289.87
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 214.5352 | 17.1487 | 24.9569 | 1000.02 | 0.0067444 | 0.30719 | 275.742 | 286.423
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 211.7863 | 17.0696 | 24.9534 | 1000.01 | 0.0067454 | 0.30742 | 272.627 | 283.215
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 209.1306 | 16.9706 | 24.9499 | 999.941 | 0.0067463 | 0.30761 | 269.55 | 280.024
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 207.3773 | 16.5782 | 24.9514 | 1000.01 | 0.0067466 | 0.30755 | 266.478 | 276.547
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 205.0252 | 16.3301 | 24.9625 | 999.992 | 0.0067479 | 0.30702 | 263.361 | 273.168
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 202.0266 | 16.2763 | 24.9713 | 1000.01 | 0.0067491 | 0.30658 | 260.204 | 269.942
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 199.5237 | 16.0918 | 24.9753 | 999.987 | 0.0067502 | 0.30644 | 257.064 | 266.605
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 197.1378 | 15.9179 | 24.9754 | 999.993 | 0.0067535 | 0.30662 | 254.02 | 263.373
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 194.5026 | 15.8142 | 24.9735 | 999.99 | 0.0067552 | 0.30679 | 250.954 | 260.189
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 192.5305 | 15.4885 | 24.972 | 999.963 | 0.0067528 | 0.3067 | 247.879 | 256.776
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 190.3652 | 15.2305 | 24.974 | 1000.04 | 0.0067534 | 0.30665 | 244.824 | 253.449
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 187.5992 | 15.1236 | 24.9827 | 999.99 | 0.0067532 | 0.30615 | 241.733 | 250.236
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 185.7495 | 14.7364 | 24.9898 | 1000.02 | 0.0067526 | 0.3057 | 238.642 | 246.745
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 183.4466 | 14.4912 | 24.9923 | 999.929 | 0.0067503 | 0.30543 | 235.522 | 243.365
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 180.8598 | 14.3479 | 24.9901 | 1000.04 | 0.00675 | 0.30548 | 232.393 | 240.077
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 178.637 | 14.0791 | 24.9929 | 1000.03 | 0.0067504 | 0.30532 | 229.259 | 236.663
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 175.4124 | 14.1303 | 24.9987 | 999.945 | 0.0067538 | 0.30517 | 226.159 | 233.598
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 173.1159 | 13.9351 | 24.9959 | 999.956 | 0.0067544 | 0.30533 | 223.138 | 230.367
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 172.0216 | 13.3535 | 24.9957 | 999.963 | 0.0067525 | 0.30522 | 220.11 | 226.746
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 169.4387 | 13.1845 | 25.0012 | 1000.06 | 0.0067497 | 0.30477 | 216.979 | 223.431
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 166.6457 | 13.0965 | 25.0041 | 999.937 | 0.0067472 | 0.30445 | 213.88 | 220.229
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 164.5427 | 12.8212 | 25.0021 | 999.975 | 0.0067454 | 0.30439 | 210.823 | 216.883
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 162.0325 | 12.6539 | 25.004 | 999.929 | 0.0067438 | 0.30412 | 207.741 | 213.619
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 159.7539 | 12.4129 | 25.009 | 1000.04 | 0.0067435 | 0.30379 | 204.659 | 210.282
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 157.7215 | 12.0914 | 25.0107 | 999.969 | 0.0067393 | 0.30343 | 201.56 | 206.847
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 155.0779 | 11.9426 | 25.0158 | 999.962 | 0.0067287 | 0.30247 | 198.494 | 203.616
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 152.2767 | 11.8563 | 25.0189 | 1000 | 0.0067208 | 0.30174 | 195.442 | 200.462
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 150.2538 | 11.5802 | 25.0093 | 999.977 | 0.0067158 | 0.30186 | 192.428 | 197.158
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 148.1325 | 11.3404 | 24.9994 | 1000.25 | 0.0067142 | 0.30221 | 189.407 | 193.883
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 146.0004 | 11.0685 | 24.9997 | 999.985 | 0.0067124 | 0.30206 | 186.337 | 190.527
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 143.9369 | 10.7623 | 24.9999 | 999.925 | 0.0067089 | 0.30181 | 183.232 | 187.102
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 141.3106 | 10.6353 | 24.9989 | 999.985 | 0.0067053 | 0.30156 | 180.152 | 183.88
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 138.6437 | 10.5449 | 24.9932 | 1000.23 | 0.0067043 | 0.30177 | 177.094 | 180.717
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 136.9119 | 10.1652 | 24.9916 | 999.96 | 0.0067013 | 0.30167 | 174.062 | 177.292
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 134.712 | 9.9062 | 24.9984 | 999.95 | 0.0067012 | 0.30126 | 170.995 | 173.952
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 132.1844 | 9.7632 | 24.9931 | 999.953 | 0.006701 | 0.30148 | 167.903 | 170.703
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 130.105 | 9.4749 | 24.9886 | 999.937 | 0.006698 | 0.30151 | 164.805 | 167.303
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 127.5043 | 9.3384 | 24.9901 | 999.989 | 0.006697 | 0.30133 | 161.709 | 164.056
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 125.0453 | 9.1847 | 24.9866 | 999.921 | 0.0066949 | 0.30131 | 158.68 | 160.859
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 123.3389 | 8.7963 | 24.9844 | 999.93 | 0.0066888 | 0.30104 | 155.658 | 157.435
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 121.1624 | 8.5375 | 24.9809 | 999.943 | 0.0066869 | 0.30105 | 152.562 | 154.066
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 118.505 | 8.4281 | 24.9726 | 1000.07 | 0.0066842 | 0.30125 | 149.455 | 150.835
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 116.643 | 8.072 | 24.9717 | 999.901 | 0.0066801 | 0.30102 | 146.377 | 147.387
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 114.4101 | 7.8208 | 24.9788 | 1000.02 | 0.0066792 | 0.30056 | 143.294 | 144.039
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 112.0088 | 7.625 | 24.9801 | 999.927 | 0.006677 | 0.30033 | 140.201 | 140.736
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 110.0545 | 7.2912 | 24.9839 | 999.931 | 0.0066712 | 0.29974 | 137.13 | 137.317
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 107.2332 | 7.2118 | 24.9915 | 1000.02 | 0.0066668 | 0.29904 | 134.039 | 134.132
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 104.6168 | 7.093 | 24.9902 | 999.964 | 0.0066637 | 0.29887 | 130.976 | 130.936
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 103.0811 | 6.6655 | 24.9829 | 999.963 | 0.0066599 | 0.29897 | 127.949 | 127.468
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 100.879 | 6.4155 | 24.9833 | 999.951 | 0.0066564 | 0.29869 | 124.873 | 124.128
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 98.2101 | 6.2984 | 24.992 | 999.945 | 0.0066558 | 0.29817 | 121.789 | 120.913
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 96.1661 | 6.0082 | 24.9922 | 999.968 | 0.0066534 | 0.29798 | 118.728 | 117.548
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 93.9687 | 5.7592 | 24.994 | 999.966 | 0.0066492 | 0.29763 | 115.661 | 114.218
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 91.7916 | 5.496 | 25.0013 | 999.986 | 0.0066457 | 0.297 | 112.583 | 110.863
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 89.8925 | 5.1599 | 24.9963 | 999.959 | 0.0066419 | 0.29696 | 109.498 | 107.428
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 87.2877 | 5.0273 | 24.9966 | 1000.24 | 0.0066368 | 0.29661 | 106.404 | 104.187
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 84.5714 | 4.9323 | 25.0053 | 999.922 | 0.0066326 | 0.29586 | 103.349 | 101.023
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 82.5285 | 4.6524 | 25.006 | 999.989 | 0.0066303 | 0.29563 | 100.316 | 97.6965
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 80.4027 | 4.3971 | 25.01 | 999.925 | 0.0066273 | 0.29521 | 97.2923 | 94.4042
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 78.3649 | 4.1121 | 25.0123 | 999.946 | 0.0066245 | 0.2949 | 94.2477 | 91.0609
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 76.2826 | 3.8241 | 25.0161 | 999.954 | 0.0066203 | 0.29442 | 91.1535 | 87.6646
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 73.7798 | 3.6644 | 25.0216 | 999.936 | 0.0066168 | 0.29388 | 88.0709 | 84.4087
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 71.4378 | 3.4812 | 25.0129 | 999.943 | 0.0066154 | 0.29417 | 85.031 | 81.1721
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 69.7282 | 3.0931 | 25.0199 | 999.969 | 0.0066113 | 0.29356 | 81.9731 | 77.7122
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 67.7085 | 2.7925 | 25.0277 | 999.932 | 0.0066083 | 0.29295 | 78.8942 | 74.3195
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 65.1513 | 2.6483 | 25.0284 | 999.977 | 0.0066045 | 0.29265 | 75.791 | 71.058
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 63.0927 | 2.363 | 25.0342 | 999.953 | 0.0065999 | 0.29207 | 72.7174 | 67.6851
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 60.759 | 2.1705 | 25.033 | 999.958 | 0.0065968 | 0.2919 | 69.6687 | 64.4308
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 58.213 | 2.0356 | 25.033 | 999.989 | 0.0065928 | 0.2916 | 66.6049 | 61.2185
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 56.5137 | 1.6584 | 25.0274 | 999.962 | 0.0065888 | 0.29161 | 63.557 | 57.7793
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 54.5356 | 1.3593 | 25.0253 | 999.942 | 0.006586 | 0.29149 | 60.4984 | 54.4087
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 52.0721 | 1.1933 | 25.0278 | 999.95 | 0.0065838 | 0.29116 | 57.4083 | 51.1394
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 50.3387 | 0.81548 | 25.0343 | 999.962 | 0.0065803 | 0.2906 | 54.3299 | 47.669
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 48.3376 | 0.50829 | 25.0367 | 999.932 | 0.0065771 | 0.29026 | 51.2166 | 44.2352
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 45.9066 | 0.32204 | 25.0338 | 999.932 | 0.0065728 | 0.29008 | 48.0816 | 40.9011
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 43.9094 | 0.021692 | 25.0259 | 999.949 | 0.006569 | 0.29021 | 44.9903 | 37.495
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 41.4114 | -0.13278 | 25.0302 | 999.965 | 0.0065672 | 0.28984 | 41.8971 | 34.2338
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 38.9395 | -0.29247 | 25.0438 | 999.925 | 0.0065633 | 0.2889 | 38.8103 | 30.974
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 37.2626 | -0.6816 | 25.0433 | 1000.26 | 0.0065584 | 0.28859 | 35.7302 | 27.4905
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 35.1577 | -0.9512 | 25.0353 | 999.989 | 0.0065566 | 0.28882 | 32.6338 | 24.1116
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 32.5598 | -1.0803 | 25.0298 | 999.994 | 0.0065547 | 0.28893 | 29.5251 | 20.861
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 30.654 | -1.3966 | 25.0333 | 999.955 | 0.0065469 | 0.28828 | 26.4579 | 17.463
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 28.5372 | -1.6529 | 25.036 | 1000.02 | 0.0065435 | 0.28791 | 23.3862 | 14.1216
Populus deltoides | LI2C |  | RACiR 500 to 0 no VPD control | 26.1926 | -1.8493 | 25.0318 | 999.982 | 0.0065419 | 0.28799 | 20.2992 | 10.8254
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 307.0625 | 24.4542 | 25.2718 | 999.842 | 0.0059292 | 0.23994 | 417.94 | 427.74
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 301.7253 | 24.2944 | 25.2605 | 1000.19 | 0.0059213 | 0.2399 | 411.798 | 421.404
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 298.4955 | 23.6044 | 25.257 | 1000.13 | 0.0059161 | 0.23972 | 405.652 | 414.536
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 294.958 | 22.989 | 25.2457 | 999.89 | 0.0059088 | 0.23972 | 399.417 | 407.652
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 289.23 | 22.9027 | 25.2318 | 1000.2 | 0.0059032 | 0.2399 | 393.094 | 401.209
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 284.6261 | 22.5294 | 25.2233 | 1000.21 | 0.0058962 | 0.2398 | 386.831 | 394.54
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 279.1717 | 22.3812 | 25.2112 | 999.864 | 0.0058891 | 0.23983 | 380.595 | 388.122
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 273.564 | 22.2901 | 25.1946 | 999.859 | 0.0058808 | 0.23997 | 374.392 | 381.794
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 270.1293 | 21.6626 | 25.1833 | 999.85 | 0.005873 | 0.23994 | 368.225 | 374.967
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 266.0396 | 21.1775 | 25.173 | 1000.17 | 0.0058656 | 0.23988 | 362 | 368.224
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 260.5187 | 21.0513 | 25.1606 | 999.857 | 0.0058604 | 0.24002 | 355.748 | 361.813
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 255.9166 | 20.6908 | 25.1534 | 999.868 | 0.0058535 | 0.23989 | 349.556 | 355.228
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 250.3467 | 20.5737 | 25.1431 | 999.836 | 0.0058455 | 0.23981 | 343.375 | 348.897
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 245.256 | 20.3703 | 25.1263 | 999.807 | 0.0058361 | 0.23988 | 337.262 | 342.547
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 242.1955 | 19.6964 | 25.1032 | 999.914 | 0.0058274 | 0.2402 | 331.149 | 335.728
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 238.1064 | 19.2298 | 25.0925 | 999.847 | 0.0058223 | 0.24026 | 324.968 | 329.049
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 232.8056 | 19.0468 | 25.0797 | 1000.16 | 0.0058139 | 0.24025 | 318.758 | 322.623
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 229.5566 | 18.3839 | 25.0634 | 1000.09 | 0.0058056 | 0.2404 | 312.576 | 315.746
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 225.817 | 17.823 | 25.0532 | 999.83 | 0.0058018 | 0.24053 | 306.335 | 308.913
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 220.0799 | 17.7511 | 25.0362 | 999.756 | 0.0057973 | 0.24089 | 300.029 | 302.503
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 215.4229 | 17.4108 | 25.0231 | 999.954 | 0.0057912 | 0.24101 | 293.783 | 295.885
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 210.4649 | 17.1745 | 25.0097 | 999.825 | 0.0057876 | 0.24129 | 287.615 | 289.45
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 205.2857 | 16.9896 | 24.9987 | 999.83 | 0.0057823 | 0.24143 | 281.486 | 283.105
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 201.8907 | 16.3644 | 24.9931 | 999.805 | 0.00578 | 0.24147 | 275.372 | 276.336
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 198.1466 | 15.8282 | 24.9825 | 1000.19 | 0.0057785 | 0.24173 | 269.212 | 269.61
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 192.7809 | 15.699 | 24.9623 | 1000.09 | 0.0057752 | 0.24231 | 263.002 | 263.24
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 188.4619 | 15.3306 | 24.9493 | 1000.2 | 0.0057786 | 0.24298 | 256.852 | 256.693
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 184.4276 | 15.0464 | 24.9451 | 1000.12 | 0.005829 | 0.24608 | 250.725 | 250.265
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 180.5772 | 14.8128 | 24.9488 | 1000.16 | 0.0059116 | 0.25087 | 244.634 | 243.931
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 177.9218 | 14.1626 | 24.9586 | 1000.15 | 0.0059703 | 0.25401 | 238.573 | 237.204
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 174.2876 | 13.6001 | 24.9706 | 1000.16 | 0.0059908 | 0.25473 | 232.461 | 230.504
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 168.8361 | 13.3623 | 24.9949 | 1000.1 | 0.0059813 | 0.25315 | 226.287 | 224.059
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 164.7845 | 12.7325 | 25.0204 | 1000.1 | 0.0059604 | 0.25088 | 220.109 | 217.216
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 160.3908 | 12.2156 | 25.0355 | 999.807 | 0.0059385 | 0.24897 | 213.919 | 210.474
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 154.6232 | 12.0657 | 25.0479 | 1000.1 | 0.0059223 | 0.24752 | 207.705 | 204.076
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 150.1156 | 11.6113 | 25.0702 | 1000.21 | 0.0059154 | 0.24623 | 201.496 | 197.38
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 145.3062 | 11.2681 | 25.0849 | 1000.1 | 0.0059126 | 0.24546 | 195.314 | 190.824
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 139.7948 | 11.1175 | 25.1013 | 1000.13 | 0.005912 | 0.24477 | 189.187 | 184.516
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 135.8753 | 10.583 | 25.1164 | 999.876 | 0.0059103 | 0.24409 | 183.094 | 177.858
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 131.9515 | 10.0513 | 25.1316 | 1000.11 | 0.0059102 | 0.2435 | 176.979 | 171.182
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 127.0149 | 9.7773 | 25.1396 | 1000.14 | 0.0059123 | 0.24331 | 170.818 | 164.717
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 123.0322 | 9.2701 | 25.1485 | 999.79 | 0.0059143 | 0.24307 | 164.669 | 158.03
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 118.3124 | 8.9447 | 25.1607 | 1000.09 | 0.0059166 | 0.24271 | 158.54 | 151.546
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 113.126 | 8.7681 | 25.1543 | 999.811 | 0.0059173 | 0.24296 | 152.448 | 145.248
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 109.6608 | 8.164 | 25.1562 | 1000.01 | 0.0059165 | 0.24286 | 146.396 | 138.561
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 106.1729 | 7.5526 | 25.1645 | 999.797 | 0.0059189 | 0.24264 | 140.306 | 131.831
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 101.4308 | 7.2408 | 25.1681 | 1000.15 | 0.0059196 | 0.24252 | 134.162 | 125.345
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 98.0901 | 6.579 | 25.1715 | 1000.04 | 0.0059174 | 0.24227 | 128.007 | 118.497
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 94.2373 | 6.0379 | 25.1725 | 1000.1 | 0.0059178 | 0.24225 | 121.801 | 111.72
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 89.0233 | 5.82 | 25.175 | 999.805 | 0.0059182 | 0.24215 | 115.558 | 105.229
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 85.0207 | 5.313 | 25.176 | 999.906 | 0.0059178 | 0.24208 | 109.344 | 98.4767
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 80.1448 | 5.0305 | 25.1703 | 1000.07 | 0.0059173 | 0.24224 | 103.15 | 91.97
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 74.9507 | 4.8326 | 25.1734 | 1000.1 | 0.0059203 | 0.24226 | 97.0002 | 85.5923
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 71.8068 | 4.1385 | 25.1778 | 999.798 | 0.0059193 | 0.24203 | 90.8957 | 78.763
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 67.8701 | 3.6281 | 25.1707 | 1000.11 | 0.005918 | 0.2422 | 84.7141 | 72.0406
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 62.8369 | 3.382 | 25.16 | 999.82 | 0.0059169 | 0.24251 | 78.5069 | 65.5571
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 59.0959 | 2.8277 | 25.1511 | 999.903 | 0.005915 | 0.24274 | 72.3464 | 58.8114
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 54.3496 | 2.512 | 25.1536 | 1000.12 | 0.0059126 | 0.24249 | 66.1782 | 52.2963
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 49.5143 | 2.232 | 25.1514 | 1000.12 | 0.0059123 | 0.24251 | 60.0537 | 45.8634
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 46.5323 | 1.5071 | 25.1416 | 999.81 | 0.0059113 | 0.2428 | 53.9627 | 39.0169
Populus deltoides | LI2C |  | RACiR 500 to 0 2.5 min | 43.2261 | 0.84319 | 25.1419 | 1000.02 | 0.0059102 | 0.24268 | 47.8091 | 32.1696
Populus deltoides | LI2C |  | RACiR 500 to 0 | 326.4885 | 27.3788 | 24.982 | 999.995 | 0.0055534 | 0.24125 | 448.9 | 470.077
Populus deltoides | LI2C |  | RACiR 500 to 0 | 323.6312 | 27.2766 | 24.9788 | 999.988 | 0.0055445 | 0.24065 | 445.841 | 466.899
Populus deltoides | LI2C |  | RACiR 500 to 0 | 321.3606 | 27.0226 | 24.9758 | 999.996 | 0.0055313 | 0.23993 | 442.797 | 463.584
Populus deltoides | LI2C |  | RACiR 500 to 0 | 318.87 | 26.8167 | 24.9813 | 999.997 | 0.0055216 | 0.23913 | 439.769 | 460.339
Populus deltoides | LI2C |  | RACiR 500 to 0 | 316.0671 | 26.7331 | 24.9804 | 1000.06 | 0.0055162 | 0.23884 | 436.69 | 457.16
Populus deltoides | LI2C |  | RACiR 500 to 0 | 314.2695 | 26.4401 | 24.9728 | 1000.04 | 0.005511 | 0.2388 | 433.622 | 453.786
Populus deltoides | LI2C |  | RACiR 500 to 0 | 311.99 | 26.2302 | 24.9702 | 999.983 | 0.0055068 | 0.23859 | 430.497 | 450.441
Populus deltoides | LI2C |  | RACiR 500 to 0 | 309.254 | 26.1285 | 24.9736 | 999.697 | 0.0055067 | 0.23841 | 427.357 | 447.187
Populus deltoides | LI2C |  | RACiR 500 to 0 | 307.7292 | 25.7592 | 24.9728 | 1000.12 | 0.0055069 | 0.23844 | 424.213 | 443.663
Populus deltoides | LI2C |  | RACiR 500 to 0 | 305.0252 | 25.6504 | 24.9716 | 1000.03 | 0.0055039 | 0.23833 | 421.033 | 440.365
Populus deltoides | LI2C |  | RACiR 500 to 0 | 302.0407 | 25.6174 | 24.9752 | 1000.01 | 0.0055043 | 0.2382 | 417.906 | 437.19
Populus deltoides | LI2C |  | RACiR 500 to 0 | 300.0972 | 25.355 | 24.9755 | 999.656 | 0.0055044 | 0.23819 | 414.805 | 433.815
Populus deltoides | LI2C |  | RACiR 500 to 0 | 296.5521 | 25.2766 | 24.9795 | 1000.02 | 0.0054672 | 0.23642 | 411.657 | 430.567
Populus deltoides | LI2C |  | RACiR 500 to 0 | 293.4269 | 25.136 | 24.9861 | 999.982 | 0.0054334 | 0.2349 | 408.552 | 427.295
Populus deltoides | LI2C |  | RACiR 500 to 0 | 291.7844 | 24.8601 | 24.9853 | 999.975 | 0.0054393 | 0.2355 | 405.399 | 423.856
Populus deltoides | LI2C |  | RACiR 500 to 0 | 289.7978 | 24.6955 | 24.9824 | 1000.02 | 0.0054522 | 0.23642 | 402.232 | 420.521
Populus deltoides | LI2C |  | RACiR 500 to 0 | 287.4695 | 24.5833 | 24.9809 | 1000.05 | 0.0054622 | 0.23708 | 399.071 | 417.24
Populus deltoides | LI2C |  | RACiR 500 to 0 | 285.3599 | 24.4054 | 24.9793 | 1000.01 | 0.0054694 | 0.23754 | 395.944 | 413.927
Populus deltoides | LI2C |  | RACiR 500 to 0 | 283.0405 | 24.2693 | 24.9753 | 999.989 | 0.0054722 | 0.23788 | 392.837 | 410.678
Populus deltoides | LI2C |  | RACiR 500 to 0 | 280.7206 | 24.1176 | 24.978 | 1000.03 | 0.0054778 | 0.23809 | 389.728 | 407.404
Populus deltoides | LI2C |  | RACiR 500 to 0 | 278.6811 | 23.8568 | 24.9882 | 1000.02 | 0.0054806 | 0.23787 | 386.637 | 404.044
Populus deltoides | LI2C |  | RACiR 500 to 0 | 276.6821 | 23.6505 | 24.9842 | 1000.02 | 0.0054844 | 0.23828 | 383.536 | 400.73
Populus deltoides | LI2C |  | RACiR 500 to 0 | 274.3567 | 23.4751 | 24.99 | 1000.01 | 0.005489 | 0.23831 | 380.405 | 397.412
Populus deltoides | LI2C |  | RACiR 500 to 0 | 272.4354 | 23.1586 | 24.9965 | 1000.07 | 0.0054883 | 0.23799 | 377.237 | 393.916
Populus deltoides | LI2C |  | RACiR 500 to 0 | 269.6731 | 23.107 | 24.9873 | 1000.02 | 0.0054879 | 0.23831 | 374.046 | 390.665
Populus deltoides | LI2C |  | RACiR 500 to 0 | 266.6002 | 23.1148 | 24.9863 | 1000.04 | 0.005489 | 0.2384 | 370.898 | 387.511
Populus deltoides | LI2C |  | RACiR 500 to 0 | 264.6222 | 22.848 | 24.992 | 1000.01 | 0.0054906 | 0.23826 | 367.803 | 384.14
Populus deltoides | LI2C |  | RACiR 500 to 0 | 262.335 | 22.6719 | 24.9923 | 1000.08 | 0.0054909 | 0.23828 | 364.708 | 380.861
Populus deltoides | LI2C |  | RACiR 500 to 0 | 260.0303 | 22.511 | 24.9916 | 999.791 | 0.0054928 | 0.23843 | 361.605 | 377.585
Populus deltoides | LI2C |  | RACiR 500 to 0 | 258.3944 | 22.1853 | 24.9873 | 1000.08 | 0.0054931 | 0.23862 | 358.468 | 374.112
Populus deltoides | LI2C |  | RACiR 500 to 0 | 256.2162 | 21.9844 | 24.9852 | 999.954 | 0.005495 | 0.23881 | 355.311 | 370.745
Populus deltoides | LI2C |  | RACiR 500 to 0 | 253.336 | 21.9235 | 24.9951 | 1000.01 | 0.0054991 | 0.23862 | 352.19 | 367.551
Populus deltoides | LI2C |  | RACiR 500 to 0 | 251.1736 | 21.7027 | 24.9996 | 1000.08 | 0.0055009 | 0.23853 | 349.081 | 364.212
Populus deltoides | LI2C |  | RACiR 500 to 0 | 248.4417 | 21.6058 | 25.0028 | 1000.01 | 0.0055001 | 0.23832 | 345.959 | 360.983
Populus deltoides | LI2C |  | RACiR 500 to 0 | 245.8174 | 21.5041 | 25.0048 | 999.785 | 0.0055007 | 0.23824 | 342.875 | 357.785
Populus deltoides | LI2C |  | RACiR 500 to 0 | 244.3417 | 21.1521 | 24.9978 | 999.994 | 0.0054995 | 0.23844 | 339.788 | 354.337
Populus deltoides | LI2C |  | RACiR 500 to 0 | 242.2831 | 20.9063 | 24.9968 | 1000.03 | 0.0054977 | 0.23836 | 336.667 | 350.96
Populus deltoides | LI2C |  | RACiR 500 to 0 | 239.333 | 20.8396 | 25.0081 | 999.977 | 0.0054972 | 0.23789 | 333.554 | 347.767
Populus deltoides | LI2C |  | RACiR 500 to 0 | 237.0927 | 20.6425 | 25.01 | 1000.05 | 0.0054964 | 0.23778 | 330.467 | 344.473
Populus deltoides | LI2C |  | RACiR 500 to 0 | 234.6666 | 20.5215 | 25.0063 | 999.972 | 0.0054996 | 0.23805 | 327.364 | 341.24
Populus deltoides | LI2C |  | RACiR 500 to 0 | 232.2457 | 20.3784 | 25.0091 | 999.989 | 0.0055019 | 0.23804 | 324.287 | 338.008
Populus deltoides | LI2C |  | RACiR 500 to 0 | 230.556 | 20.0572 | 25.0087 | 999.955 | 0.0054993 | 0.23796 | 321.22 | 334.61
Populus deltoides | LI2C |  | RACiR 500 to 0 | 228.5162 | 19.8152 | 25.0103 | 1000.02 | 0.0055 | 0.23794 | 318.112 | 331.25
Populus deltoides | LI2C |  | RACiR 500 to 0 | 226.0935 | 19.677 | 25.0135 | 1000.03 | 0.0055043 | 0.23802 | 315.02 | 328.009
Populus deltoides | LI2C |  | RACiR 500 to 0 | 224.2388 | 19.3994 | 25.012 | 999.903 | 0.0055048 | 0.2381 | 311.914 | 324.615
Populus deltoides | LI2C |  | RACiR 500 to 0 | 222.3886 | 19.1019 | 25.0111 | 1000 | 0.0055021 | 0.23799 | 308.793 | 321.187
Populus deltoides | LI2C |  | RACiR 500 to 0 | 220.0248 | 18.945 | 25.0088 | 1000.05 | 0.0055042 | 0.23818 | 305.639 | 317.864
Populus deltoides | LI2C |  | RACiR 500 to 0 | 217.4896 | 18.806 | 25.0128 | 999.973 | 0.0055067 | 0.23815 | 302.467 | 314.543
Populus deltoides | LI2C |  | RACiR 500 to 0 | 214.9547 | 18.6908 | 25.0123 | 999.986 | 0.0055082 | 0.2382 | 299.367 | 311.318
Populus deltoides | LI2C |  | RACiR 500 to 0 | 212.9043 | 18.4828 | 25.0107 | 999.971 | 0.0055104 | 0.23837 | 296.329 | 308.061
Populus deltoides | LI2C |  | RACiR 500 to 0 | 210.9781 | 18.2161 | 25.014 | 999.992 | 0.0055117 | 0.23832 | 293.244 | 304.7
Populus deltoides | LI2C |  | RACiR 500 to 0 | 208.5189 | 18.0898 | 25.0074 | 1000.02 | 0.0055105 | 0.23851 | 290.123 | 301.443
Populus deltoides | LI2C |  | RACiR 500 to 0 | 206.5627 | 17.8522 | 25.005 | 999.968 | 0.0055116 | 0.23866 | 287.065 | 298.135
Populus deltoides | LI2C |  | RACiR 500 to 0 | 205.3276 | 17.421 | 25.0035 | 1000 | 0.0055115 | 0.23871 | 283.956 | 294.585
Populus deltoides | LI2C |  | RACiR 500 to 0 | 202.6402 | 17.3157 | 25.0043 | 999.953 | 0.005511 | 0.23864 | 280.783 | 291.297
Populus deltoides | LI2C |  | RACiR 500 to 0 | 199.6083 | 17.3138 | 25.0069 | 999.967 | 0.0055128 | 0.23863 | 277.675 | 288.175
Populus deltoides | LI2C |  | RACiR 500 to 0 | 197.7091 | 17.0701 | 25.0008 | 1000 | 0.0055137 | 0.2389 | 274.613 | 284.859
Populus deltoides | LI2C |  | RACiR 500 to 0 | 195.589 | 16.8642 | 24.9987 | 999.984 | 0.005514 | 0.23899 | 271.541 | 281.572
Populus deltoides | LI2C |  | RACiR 500 to 0 | 193.3612 | 16.6636 | 25.0088 | 999.99 | 0.0055168 | 0.23874 | 268.495 | 278.314
Populus deltoides | LI2C |  | RACiR 500 to 0 | 191.7982 | 16.2925 | 25.0168 | 999.999 | 0.0055181 | 0.23849 | 265.404 | 274.842
Populus deltoides | LI2C |  | RACiR 500 to 0 | 189.2492 | 16.1811 | 25.0122 | 1000.01 | 0.0055176 | 0.23863 | 262.277 | 271.594
Populus deltoides | LI2C |  | RACiR 500 to 0 | 186.5809 | 16.1323 | 25.0041 | 999.981 | 0.0055198 | 0.23905 | 259.214 | 268.47
Populus deltoides | LI2C |  | RACiR 500 to 0 | 184.9614 | 15.8289 | 24.997 | 1000.02 | 0.0055211 | 0.2394 | 256.167 | 265.11
Populus deltoides | LI2C |  | RACiR 500 to 0 | 182.8113 | 15.6232 | 24.9977 | 999.985 | 0.0055214 | 0.23941 | 253.094 | 261.822
Populus deltoides | LI2C |  | RACiR 500 to 0 | 180.7573 | 15.3826 | 25.0035 | 1000.03 | 0.0055208 | 0.23916 | 250.05 | 258.526
Populus deltoides | LI2C |  | RACiR 500 to 0 | 178.8688 | 15.0914 | 25.0095 | 999.99 | 0.0055209 | 0.23892 | 246.953 | 255.127
Populus deltoides | LI2C |  | RACiR 500 to 0 | 176.2354 | 14.9962 | 25.0098 | 999.975 | 0.0055235 | 0.23901 | 243.831 | 251.9
Populus deltoides | LI2C |  | RACiR 500 to 0 | 174.1547 | 14.8017 | 25.0045 | 999.948 | 0.0055248 | 0.23924 | 240.817 | 248.681
Populus deltoides | LI2C |  | RACiR 500 to 0 | 172.5806 | 14.4692 | 25.0021 | 999.957 | 0.0055264 | 0.23942 | 237.75 | 245.272
Populus deltoides | LI2C |  | RACiR 500 to 0 | 170.3556 | 14.2817 | 24.9984 | 999.985 | 0.0055261 | 0.23956 | 234.644 | 241.968
Populus deltoides | LI2C |  | RACiR 500 to 0 | 167.9724 | 14.1268 | 24.9992 | 999.995 | 0.0055248 | 0.23944 | 231.581 | 238.738
Populus deltoides | LI2C |  | RACiR 500 to 0 | 166.075 | 13.8379 | 25.0064 | 1000.03 | 0.0055259 | 0.23921 | 228.481 | 235.339
Populus deltoides | LI2C |  | RACiR 500 to 0 | 163.8676 | 13.6352 | 25.008 | 1000.01 | 0.0055282 | 0.23927 | 225.35 | 231.996
Populus deltoides | LI2C |  | RACiR 500 to 0 | 161.7023 | 13.4438 | 25.0042 | 999.95 | 0.0055296 | 0.23951 | 222.266 | 228.709
Populus deltoides | LI2C |  | RACiR 500 to 0 | 160.0894 | 13.1089 | 25.0044 | 999.982 | 0.0055306 | 0.23955 | 219.19 | 225.288
Populus deltoides | LI2C |  | RACiR 500 to 0 | 157.8719 | 12.9148 | 25.0024 | 999.968 | 0.0055316 | 0.23964 | 216.08 | 221.974
Populus deltoides | LI2C |  | RACiR 500 to 0 | 155.2299 | 12.8125 | 25.0026 | 999.917 | 0.0055316 | 0.23961 | 212.952 | 218.732
Populus deltoides | LI2C |  | RACiR 500 to 0 | 153.3767 | 12.5149 | 25.0084 | 999.953 | 0.005531 | 0.23936 | 209.858 | 215.33
Populus deltoides | LI2C |  | RACiR 500 to 0 | 151.1614 | 12.3344 | 25.003 | 1000.27 | 0.0055307 | 0.23954 | 206.785 | 212.066
Populus deltoides | LI2C |  | RACiR 500 to 0 | 148.7336 | 12.2055 | 24.9968 | 999.986 | 0.0055315 | 0.23983 | 203.689 | 208.83
Populus deltoides | LI2C |  | RACiR 500 to 0 | 147.18 | 11.8625 | 24.9958 | 999.987 | 0.0055349 | 0.24006 | 200.604 | 205.392
Populus deltoides | LI2C |  | RACiR 500 to 0 | 145.0489 | 11.6464 | 24.9972 | 1000 | 0.0055357 | 0.24002 | 197.523 | 202.085
Populus deltoides | LI2C |  | RACiR 500 to 0 | 142.4056 | 11.5535 | 24.9996 | 1000.06 | 0.0055371 | 0.23996 | 194.439 | 198.897
Populus deltoides | LI2C |  | RACiR 500 to 0 | 140.9235 | 11.1738 | 25.0024 | 999.95 | 0.0055363 | 0.23979 | 191.357 | 195.425
Populus deltoides | LI2C |  | RACiR 500 to 0 | 138.9552 | 10.919 | 25.0003 | 999.981 | 0.0055369 | 0.23992 | 188.242 | 192.045
Populus deltoides | LI2C |  | RACiR 500 to 0 | 136.1978 | 10.8628 | 25.0002 | 1000.04 | 0.0055415 | 0.24017 | 185.134 | 188.87
Populus deltoides | LI2C |  | RACiR 500 to 0 | 134.0292 | 10.6495 | 25.004 | 1000.27 | 0.0055416 | 0.24004 | 182.044 | 185.556
Populus deltoides | LI2C |  | RACiR 500 to 0 | 131.8142 | 10.4659 | 25.0004 | 999.903 | 0.0055409 | 0.24014 | 178.982 | 182.3
Populus deltoides | LI2C |  | RACiR 500 to 0 | 129.9259 | 10.215 | 25.0008 | 999.965 | 0.005545 | 0.24034 | 175.957 | 179.014
Populus deltoides | LI2C |  | RACiR 500 to 0 | 128.2484 | 9.8864 | 25.0016 | 999.988 | 0.0055455 | 0.24032 | 172.861 | 175.579
Populus deltoides | LI2C |  | RACiR 500 to 0 | 126.2462 | 9.6408 | 24.9993 | 1000.26 | 0.005545 | 0.24038 | 169.767 | 172.229
Populus deltoides | LI2C |  | RACiR 500 to 0 | 123.829 | 9.4897 | 25.0022 | 999.916 | 0.0055465 | 0.24032 | 166.668 | 168.968
Populus deltoides | LI2C |  | RACiR 500 to 0 | 121.9552 | 9.2105 | 25.0018 | 1000.02 | 0.0055452 | 0.24027 | 163.584 | 165.594
Populus deltoides | LI2C |  | RACiR 500 to 0 | 119.7327 | 9.0239 | 24.9991 | 1000.01 | 0.0055458 | 0.24041 | 160.501 | 162.314
Populus deltoides | LI2C |  | RACiR 500 to 0 | 117.3866 | 8.8606 | 25.0034 | 999.968 | 0.0055474 | 0.2403 | 157.432 | 159.071
Populus deltoides | LI2C |  | RACiR 500 to 0 | 115.509 | 8.5817 | 25.0107 | 999.966 | 0.0055499 | 0.24017 | 154.359 | 155.709
Populus deltoides | LI2C |  | RACiR 500 to 0 | 113.5835 | 8.3229 | 25.012 | 999.963 | 0.005552 | 0.24026 | 151.285 | 152.366
Populus deltoides | LI2C |  | RACiR 500 to 0 | 111.3734 | 8.1377 | 25.0095 | 999.956 | 0.005553 | 0.24037 | 148.226 | 149.111
Populus deltoides | LI2C |  | RACiR 500 to 0 | 109.5765 | 7.8362 | 25.0104 | 999.917 | 0.0055526 | 0.24028 | 145.132 | 145.705
Populus deltoides | LI2C |  | RACiR 500 to 0 | 107.4053 | 7.6378 | 25.0036 | 999.977 | 0.0055545 | 0.24067 | 142.019 | 142.383
Populus deltoides | LI2C |  | RACiR 500 to 0 | 104.9637 | 7.5085 | 24.9983 | 999.934 | 0.0055564 | 0.24098 | 138.933 | 139.157
Populus deltoides | LI2C |  | RACiR 500 to 0 | 103.1217 | 7.2208 | 25.0096 | 1000.33 | 0.0055565 | 0.24054 | 135.902 | 135.828
Populus deltoides | LI2C |  | RACiR 500 to 0 | 100.9188 | 7.0213 | 25.0193 | 999.959 | 0.0055582 | 0.24025 | 132.848 | 132.564
Populus deltoides | LI2C |  | RACiR 500 to 0 | 98.7587 | 6.8317 | 25.0126 | 999.989 | 0.0055604 | 0.24062 | 129.791 | 129.307
Populus deltoides | LI2C |  | RACiR 500 to 0 | 97.3221 | 6.455 | 25.0087 | 999.988 | 0.0055617 | 0.24087 | 126.707 | 125.836
Populus deltoides | LI2C |  | RACiR 500 to 0 | 95.5631 | 6.1423 | 25.0155 | 999.938 | 0.0055642 | 0.24077 | 123.607 | 122.413
Populus deltoides | LI2C |  | RACiR 500 to 0 | 92.9154 | 6.0472 | 25.0107 | 999.968 | 0.0055655 | 0.241 | 120.471 | 119.171
Populus deltoides | LI2C |  | RACiR 500 to 0 | 90.7396 | 5.845 | 25.0026 | 1000.29 | 0.0055664 | 0.24133 | 117.36 | 115.847
Populus deltoides | LI2C |  | RACiR 500 to 0 | 88.3806 | 5.6867 | 25.0081 | 1000.01 | 0.0055699 | 0.24128 | 114.289 | 112.607
Populus deltoides | LI2C |  | RACiR 500 to 0 | 85.9994 | 5.542 | 25.0104 | 999.957 | 0.0055736 | 0.24139 | 111.235 | 109.398
Populus deltoides | LI2C |  | RACiR 500 to 0 | 84.4261 | 5.204 | 25.0099 | 999.96 | 0.0055736 | 0.24141 | 108.206 | 106.021
Populus deltoides | LI2C |  | RACiR 500 to 0 | 82.4049 | 4.9636 | 25.0124 | 999.945 | 0.0055697 | 0.24111 | 105.158 | 102.721
Populus deltoides | LI2C |  | RACiR 500 to 0 | 80.1753 | 4.7778 | 25.0116 | 999.967 | 0.0055706 | 0.24117 | 102.091 | 99.4579
Populus deltoides | LI2C |  | RACiR 500 to 0 | 78.3216 | 4.5031 | 25.0073 | 1000.25 | 0.0055731 | 0.24149 | 99.0147 | 96.0963
Populus deltoides | LI2C |  | RACiR 500 to 0 | 76.0998 | 4.3181 | 25.0043 | 999.962 | 0.0055751 | 0.24172 | 95.9483 | 92.8342
Populus deltoides | LI2C |  | RACiR 500 to 0 | 73.9479 | 4.1169 | 25.0068 | 1000.01 | 0.0055766 | 0.24166 | 92.909 | 89.5838
Populus deltoides | LI2C |  | RACiR 500 to 0 | 72.2203 | 3.8054 | 25.0098 | 999.986 | 0.0055766 | 0.24155 | 89.8451 | 86.1974
Populus deltoides | LI2C |  | RACiR 500 to 0 | 70.2905 | 3.5433 | 25.013 | 1000.3 | 0.0055765 | 0.24142 | 86.7813 | 82.8605
Populus deltoides | LI2C |  | RACiR 500 to 0 | 68.1661 | 3.3255 | 25.013 | 999.954 | 0.0055777 | 0.24147 | 83.6915 | 79.5427
Populus deltoides | LI2C |  | RACiR 500 to 0 | 66.8151 | 2.9177 | 25.0122 | 999.937 | 0.0055791 | 0.24159 | 80.5945 | 76.0277
Populus deltoides | LI2C |  | RACiR 500 to 0 | 64.1398 | 2.8237 | 25.0154 | 999.964 | 0.0055779 | 0.24139 | 77.4739 | 72.8015
Populus deltoides | LI2C |  | RACiR 500 to 0 | 61.0276 | 2.848 | 25.0125 | 1000.24 | 0.0055803 | 0.2416 | 74.3795 | 69.7205
Populus deltoides | LI2C |  | RACiR 500 to 0 | 59.6827 | 2.459 | 25.0029 | 999.907 | 0.0055833 | 0.24215 | 71.3482 | 66.2899
Populus deltoides | LI2C |  | RACiR 500 to 0 | 58.1292 | 2.1068 | 25.0078 | 999.899 | 0.0055832 | 0.24194 | 68.2958 | 62.8745
Populus deltoides | LI2C |  | RACiR 500 to 0 | 55.9025 | 1.9105 | 25.0109 | 999.898 | 0.0055843 | 0.24184 | 65.2003 | 59.5727
Populus deltoides | LI2C |  | RACiR 500 to 0 | 54.3081 | 1.5645 | 25.0105 | 1000.03 | 0.0055841 | 0.24185 | 62.1212 | 56.1368
Populus deltoides | LI2C |  | RACiR 500 to 0 | 52.1677 | 1.3465 | 25.0141 | 999.917 | 0.0055838 | 0.24169 | 59.0236 | 52.8095
Populus deltoides | LI2C |  | RACiR 500 to 0 | 49.6233 | 1.2304 | 25.0107 | 1000.29 | 0.0055877 | 0.24203 | 55.9254 | 49.5849
Populus deltoides | LI2C |  | RACiR 500 to 0 | 48.1383 | 0.86164 | 25.0086 | 999.948 | 0.0055894 | 0.2422 | 52.8613 | 46.1421
Populus deltoides | LI2C |  | RACiR 500 to 0 | 46.1891 | 0.60182 | 25.0082 | 1000.01 | 0.0055872 | 0.24209 | 49.7826 | 42.7918
Populus deltoides | LI2C |  | RACiR 500 to 0 | 43.8773 | 0.4299 | 25.0093 | 999.955 | 0.0055896 | 0.24216 | 46.6978 | 39.5253
Populus deltoides | LI2C |  | RACiR 500 to 0 | 41.9252 | 0.16414 | 25.0087 | 999.966 | 0.0055915 | 0.24229 | 43.5894 | 36.1403
Populus deltoides | LI2C |  | RACiR 500 to 0 | 39.9292 | -0.089681 | 25.0045 | 1000.02 | 0.0055917 | 0.24246 | 40.4866 | 32.7716
Populus deltoides | LI2C |  | RACiR 500 to 0 | 38.0928 | -0.37581 | 25.0031 | 999.92 | 0.0055939 | 0.24261 | 37.4143 | 29.4037
Populus deltoides | LI2C |  | RACiR 500 to 0 | 36.5494 | -0.74397 | 25.005 | 999.968 | 0.0055973 | 0.24269 | 34.3009 | 25.912
Populus deltoides | LI2C |  | RACiR 500 to 0 | 34.436 | -0.97724 | 25.0095 | 1000.3 | 0.0055954 | 0.24242 | 31.1619 | 22.527
Populus deltoides | LI2C |  | RACiR 500 to 0 | 32.1601 | -1.1706 | 25.005 | 999.993 | 0.0055946 | 0.24254 | 28.0278 | 19.1896
Populus deltoides | LI2C |  | RACiR 500 to 0 | 30.0762 | -1.4062 | 25.0025 | 999.988 | 0.005596 | 0.24271 | 24.9164 | 15.832
Populus deltoides | LI2C |  | RACiR 300 to 800 | 229.3414 | 19.1622 | 24.9846 | 1000.02 | 0.0063198 | 0.27617 | 304.84 | 333.982
Populus deltoides | LI2C |  | RACiR 300 to 800 | 231.404 | 19.468 | 24.9849 | 999.955 | 0.0063413 | 0.27694 | 307.875 | 337.34
Populus deltoides | LI2C |  | RACiR 300 to 800 | 233.8848 | 19.6053 | 24.9843 | 1000.02 | 0.006344 | 0.2769 | 310.928 | 340.54
Populus deltoides | LI2C |  | RACiR 300 to 800 | 235.6412 | 19.9334 | 24.9778 | 999.996 | 0.0063375 | 0.27676 | 313.957 | 343.905
Populus deltoides | LI2C |  | RACiR 300 to 800 | 238.099 | 20.0866 | 24.9699 | 999.979 | 0.0063322 | 0.27676 | 317.027 | 347.137
Populus deltoides | LI2C |  | RACiR 300 to 800 | 240.5249 | 20.2656 | 24.9638 | 1000.03 | 0.0063334 | 0.27707 | 320.076 | 350.372
Populus deltoides | LI2C |  | RACiR 300 to 800 | 241.726 | 20.7546 | 24.9655 | 999.98 | 0.0063341 | 0.27701 | 323.108 | 353.907
Populus deltoides | LI2C |  | RACiR 300 to 800 | 244.0357 | 20.9601 | 24.9684 | 999.97 | 0.0063356 | 0.27695 | 326.241 | 357.254
Populus deltoides | LI2C |  | RACiR 300 to 800 | 247.228 | 21.0382 | 24.9687 | 999.942 | 0.0063692 | 0.27841 | 329.39 | 360.5
Populus deltoides | LI2C |  | RACiR 300 to 800 | 249.1711 | 21.4785 | 24.973 | 999.969 | 0.0064168 | 0.28012 | 332.509 | 364.087
Populus deltoides | LI2C |  | RACiR 300 to 800 | 251.2409 | 21.7631 | 24.9718 | 999.96 | 0.006426 | 0.28014 | 335.659 | 367.534
Populus deltoides | LI2C |  | RACiR 300 to 800 | 253.6666 | 21.8875 | 24.9687 | 999.976 | 0.0064111 | 0.27925 | 338.848 | 370.85
Populus deltoides | LI2C |  | RACiR 300 to 800 | 255.7686 | 22.1245 | 24.9647 | 1000.01 | 0.0064 | 0.2787 | 342.018 | 374.267
Populus deltoides | LI2C |  | RACiR 300 to 800 | 258.6606 | 22.1578 | 24.9617 | 1000.04 | 0.0063946 | 0.27845 | 345.17 | 377.46
Populus deltoides | LI2C |  | RACiR 300 to 800 | 261.132 | 22.2926 | 24.9669 | 999.983 | 0.0063952 | 0.27821 | 348.264 | 380.699
Populus deltoides | LI2C |  | RACiR 300 to 800 | 263.0116 | 22.6796 | 24.9683 | 1000.01 | 0.0064145 | 0.27896 | 351.376 | 384.215
Populus deltoides | LI2C |  | RACiR 300 to 800 | 265.5782 | 22.9075 | 24.962 | 1000.04 | 0.0064357 | 0.28 | 354.518 | 387.601
Populus deltoides | LI2C |  | RACiR 300 to 800 | 268.1117 | 23.0683 | 24.9549 | 1000.02 | 0.0064358 | 0.28012 | 357.651 | 390.905
Populus deltoides | LI2C |  | RACiR 300 to 800 | 270.0761 | 23.3373 | 24.951 | 1000.05 | 0.0064259 | 0.27967 | 360.768 | 394.302
Populus deltoides | LI2C |  | RACiR 300 to 800 | 272.7235 | 23.4323 | 24.9503 | 999.989 | 0.0064207 | 0.27934 | 363.924 | 397.559
Populus deltoides | LI2C |  | RACiR 300 to 800 | 274.7984 | 23.6892 | 24.9537 | 999.973 | 0.006419 | 0.27904 | 367.076 | 400.977
Populus deltoides | LI2C |  | RACiR 300 to 800 | 276.6191 | 24.0311 | 24.9571 | 999.98 | 0.0064196 | 0.2789 | 370.226 | 404.482
Populus deltoides | LI2C |  | RACiR 300 to 800 | 279.4102 | 24.1156 | 24.9598 | 1000 | 0.0064218 | 0.27891 | 373.391 | 407.741
Populus deltoides | LI2C |  | RACiR 300 to 800 | 282.5787 | 24.0954 | 24.9629 | 999.992 | 0.0064255 | 0.27898 | 376.539 | 410.879
Populus deltoides | LI2C |  | RACiR 300 to 800 | 284.7982 | 24.3272 | 24.9654 | 1000.01 | 0.0064287 | 0.27909 | 379.618 | 414.205
Populus deltoides | LI2C |  | RACiR 300 to 800 | 286.6577 | 24.6348 | 24.9707 | 1000 | 0.0064312 | 0.27899 | 382.673 | 417.574
Populus deltoides | LI2C |  | RACiR 300 to 800 | 288.9023 | 24.8943 | 24.9673 | 999.98 | 0.0064347 | 0.27935 | 385.787 | 420.96
Populus deltoides | LI2C |  | RACiR 300 to 800 | 291.1588 | 25.2015 | 24.9635 | 1000.01 | 0.0064491 | 0.28022 | 388.921 | 424.419
Populus deltoides | LI2C |  | RACiR 300 to 800 | 294.0504 | 25.3034 | 24.9716 | 1000 | 0.0064733 | 0.28083 | 392.059 | 427.677
Populus deltoides | LI2C |  | RACiR 300 to 800 | 296.9524 | 25.366 | 24.9751 | 1000.03 | 0.0064854 | 0.28103 | 395.199 | 430.894
Populus deltoides | LI2C |  | RACiR 300 to 800 | 298.5459 | 25.7551 | 24.9719 | 1000.02 | 0.0064791 | 0.28073 | 398.331 | 434.428
Populus deltoides | LI2C |  | RACiR 300 to 800 | 300.5405 | 26.0579 | 24.9695 | 1000.07 | 0.0064773 | 0.28068 | 401.48 | 437.885
Populus deltoides | LI2C |  | RACiR 300 to 800 | 303.46 | 26.1143 | 24.9684 | 1000.03 | 0.0064772 | 0.2807 | 404.664 | 441.132
Populus deltoides | LI2C |  | RACiR 300 to 800 | 306.1953 | 26.234 | 24.9643 | 1000.04 | 0.0064773 | 0.28093 | 407.811 | 444.417
Populus deltoides | LI2C |  | RACiR 300 to 800 | 309.0413 | 26.3241 | 24.962 | 999.993 | 0.0064804 | 0.28124 | 410.938 | 447.643
Populus deltoides | LI2C |  | RACiR 300 to 800 | 311.2484 | 26.5438 | 24.9668 | 999.96 | 0.006482 | 0.28109 | 414.041 | 450.974
Populus deltoides | LI2C |  | RACiR 300 to 800 | 312.8913 | 26.9467 | 24.9669 | 1000.04 | 0.006484 | 0.28117 | 417.145 | 454.493
Populus deltoides | LI2C |  | RACiR 300 to 800 | 315.238 | 27.1543 | 24.9692 | 1000 | 0.0064865 | 0.28119 | 420.292 | 457.861
Populus deltoides | LI2C |  | RACiR 300 to 800 | 318.1139 | 27.2248 | 24.9733 | 999.974 | 0.0064915 | 0.2813 | 423.452 | 461.105
Populus deltoides | LI2C |  | RACiR 300 to 800 | 320.0505 | 27.5731 | 24.9753 | 999.999 | 0.0064956 | 0.2815 | 426.617 | 464.633
Populus deltoides | LI2C |  | RACiR 300 to 800 | 322.1228 | 27.8789 | 24.9809 | 1000.04 | 0.0065042 | 0.28168 | 429.778 | 468.112
Populus deltoides | LI2C |  | RACiR 300 to 800 | 325.5725 | 28.008 | 24.9784 | 999.961 | 0.0065523 | 0.28386 | 432.996 | 471.489
Populus deltoides | LI2C |  | RACiR 300 to 800 | 328.714 | 28.1768 | 24.975 | 999.994 | 0.0065943 | 0.28558 | 436.209 | 474.901
Populus deltoides | LI2C |  | RACiR 300 to 800 | 331.8154 | 28.1099 | 24.9738 | 1000.01 | 0.0065883 | 0.28498 | 439.369 | 478.002
Populus deltoides | LI2C |  | RACiR 300 to 800 | 334.2803 | 28.1858 | 24.9749 | 999.994 | 0.0065772 | 0.28416 | 442.46 | 481.174
Populus deltoides | LI2C |  | RACiR 300 to 800 | 335.5224 | 28.659 | 24.9741 | 1000.03 | 0.0065712 | 0.28379 | 445.553 | 484.75
Populus deltoides | LI2C |  | RACiR 300 to 800 | 337.5676 | 28.9571 | 24.9718 | 999.99 | 0.0065696 | 0.28374 | 448.726 | 488.23
Populus deltoides | LI2C |  | RACiR 300 to 800 | 340.5774 | 28.972 | 24.9755 | 999.996 | 0.0065696 | 0.28355 | 451.936 | 491.463
Populus deltoides | LI2C |  | RACiR 300 to 800 | 343.1278 | 29.1188 | 24.9729 | 1000.01 | 0.0065668 | 0.28353 | 455.072 | 494.766
Populus deltoides | LI2C |  | RACiR 300 to 800 | 345.4935 | 29.337 | 24.9678 | 1000.01 | 0.0065674 | 0.28376 | 458.18 | 498.097
Populus deltoides | LI2C |  | RACiR 300 to 800 | 347.6676 | 29.6078 | 24.9694 | 1000.02 | 0.0065705 | 0.28386 | 461.336 | 501.532
Populus deltoides | LI2C |  | RACiR 300 to 800 | 349.9778 | 29.8506 | 24.9705 | 999.958 | 0.0065718 | 0.2839 | 464.55 | 505.005
Populus deltoides | LI2C |  | RACiR 300 to 800 | 352.8325 | 29.9358 | 24.9743 | 999.918 | 0.0065759 | 0.28396 | 467.756 | 508.306
Populus deltoides | LI2C |  | RACiR 300 to 800 | 355.771 | 30.0011 | 24.9783 | 1000.02 | 0.0065823 | 0.28415 | 470.93 | 511.557
Populus deltoides | LI2C |  | RACiR 300 to 800 | 357.6919 | 30.3471 | 24.9781 | 1000.03 | 0.0065845 | 0.28431 | 474.064 | 515.056
Populus deltoides | LI2C |  | RACiR 300 to 800 | 360.0572 | 30.5632 | 24.9782 | 999.972 | 0.006586 | 0.2844 | 477.217 | 518.434
Populus deltoides | LI2C |  | RACiR 300 to 800 | 363.0275 | 30.6348 | 24.9777 | 1000.05 | 0.0065893 | 0.28463 | 480.422 | 521.72
Populus deltoides | LI2C |  | RACiR 300 to 800 | 365.5308 | 30.809 | 24.9809 | 999.969 | 0.0065915 | 0.28463 | 483.606 | 525.096
Populus deltoides | LI2C |  | RACiR 300 to 800 | 368.5288 | 30.8352 | 24.9844 | 999.971 | 0.0065952 | 0.28469 | 486.747 | 528.273
Populus deltoides | LI2C |  | RACiR 300 to 800 | 371.3777 | 30.9103 | 24.9852 | 1000.04 | 0.0065985 | 0.28485 | 489.868 | 531.481
Populus deltoides | LI2C |  | RACiR 300 to 800 | 373.1376 | 31.2934 | 24.9875 | 1000.05 | 0.0066027 | 0.285 | 492.973 | 534.986
Populus deltoides | LI2C |  | RACiR 300 to 800 | 375.3534 | 31.5605 | 24.9881 | 1000.01 | 0.0066068 | 0.28521 | 496.104 | 538.393
Populus deltoides | LI2C |  | RACiR 300 to 800 | 378.4177 | 31.614 | 24.9848 | 999.697 | 0.0066098 | 0.28554 | 499.291 | 541.643
Populus deltoides | LI2C |  | RACiR 300 to 800 | 380.9328 | 31.7817 | 24.988 | 1000.04 | 0.0066122 | 0.28556 | 502.458 | 544.992
Populus deltoides | LI2C |  | RACiR 300 to 800 | 383.6575 | 31.8452 | 24.9957 | 1000.06 | 0.0066135 | 0.28527 | 505.597 | 548.208
Populus deltoides | LI2C |  | RACiR 300 to 800 | 386.0717 | 32.0137 | 25.0026 | 1000.02 | 0.0066182 | 0.28523 | 508.689 | 551.481
Populus deltoides | LI2C |  | RACiR 300 to 800 | 387.9028 | 32.4032 | 25.0012 | 999.707 | 0.0066218 | 0.28554 | 511.815 | 555.011
Populus deltoides | LI2C |  | RACiR 300 to 800 | 390.8546 | 32.4956 | 24.9986 | 1000.05 | 0.0066245 | 0.28582 | 515.043 | 558.34
Populus deltoides | LI2C |  | RACiR 300 to 800 | 393.9618 | 32.5251 | 24.9999 | 1000.08 | 0.0066271 | 0.28591 | 518.293 | 561.634
Populus deltoides | LI2C |  | RACiR 300 to 800 | 395.7854 | 32.9103 | 24.9941 | 999.867 | 0.0066246 | 0.28604 | 521.463 | 565.204
Populus deltoides | LI2C |  | RACiR 300 to 800 | 398.1825 | 33.0974 | 24.9964 | 1000.11 | 0.0066225 | 0.28578 | 524.689 | 568.622
Populus deltoides | LI2C |  | RACiR 300 to 800 | 401.4485 | 33.0753 | 24.9996 | 1000 | 0.0066243 | 0.28576 | 527.967 | 571.889
Populus deltoides | LI2C |  | RACiR 300 to 800 | 404.1476 | 33.2599 | 24.9935 | 1000.03 | 0.0066268 | 0.28622 | 531.178 | 575.304
Populus deltoides | LI2C |  | RACiR 300 to 800 | 407.2557 | 33.3023 | 24.9901 | 1000.06 | 0.0066285 | 0.28648 | 534.397 | 578.572
Populus deltoides | LI2C |  | RACiR 300 to 800 | 410.3538 | 33.3083 | 24.9932 | 999.957 | 0.0066311 | 0.28647 | 537.597 | 581.791
Populus deltoides | LI2C |  | RACiR 300 to 800 | 412.5744 | 33.5767 | 24.9943 | 1000.03 | 0.006633 | 0.28655 | 540.782 | 585.261
Populus deltoides | LI2C |  | RACiR 300 to 800 | 415.4013 | 33.7273 | 24.9878 | 1000.06 | 0.0066356 | 0.28704 | 543.992 | 588.63
Populus deltoides | LI2C |  | RACiR 300 to 800 | 418.4113 | 33.7881 | 24.983 | 1000.06 | 0.0066369 | 0.28732 | 547.161 | 591.871
Populus deltoides | LI2C |  | RACiR 300 to 800 | 420.8363 | 33.9826 | 24.9859 | 1000.04 | 0.0066401 | 0.28737 | 550.312 | 595.232
Populus deltoides | LI2C |  | RACiR 300 to 800 | 423.2555 | 34.1627 | 24.9904 | 1000.05 | 0.0066418 | 0.28727 | 553.467 | 598.578
Populus deltoides | LI2C |  | RACiR 300 to 800 | 425.7634 | 34.3704 | 24.9874 | 1000.06 | 0.0066426 | 0.28748 | 556.672 | 602.001
Populus deltoides | LI2C |  | RACiR 300 to 800 | 428.367 | 34.5713 | 24.9832 | 999.893 | 0.0066439 | 0.28778 | 559.904 | 605.448
Populus deltoides | LI2C |  | RACiR 300 to 800 | 431.4819 | 34.5907 | 24.9867 | 1000.03 | 0.0066476 | 0.28784 | 563.139 | 608.716
Populus deltoides | LI2C |  | RACiR 300 to 800 | 434.5119 | 34.6198 | 24.9905 | 999.725 | 0.0066516 | 0.2879 | 566.322 | 611.941
Populus deltoides | LI2C |  | RACiR 300 to 800 | 437.1511 | 34.7541 | 24.9959 | 1000.09 | 0.0066552 | 0.28787 | 569.515 | 615.288
Populus deltoides | LI2C |  | RACiR 300 to 800 | 440.4124 | 34.849 | 24.9955 | 1000.04 | 0.0066803 | 0.28902 | 572.696 | 618.584
Populus deltoides | LI2C |  | RACiR 300 to 800 | 443.4074 | 35.0054 | 24.9955 | 999.741 | 0.0067061 | 0.29007 | 575.868 | 621.939
Populus deltoides | LI2C |  | RACiR 300 to 800 | 445.7774 | 35.2368 | 24.9919 | 1000.08 | 0.0067074 | 0.29009 | 579.093 | 625.411
Populus deltoides | LI2C |  | RACiR 300 to 800 | 448.9834 | 35.2076 | 24.985 | 999.983 | 0.0066998 | 0.28989 | 582.352 | 628.647
Populus deltoides | LI2C |  | RACiR 300 to 800 | 451.9604 | 35.2009 | 24.9881 | 999.745 | 0.0066959 | 0.28946 | 585.568 | 631.863
Populus deltoides | LI2C |  | RACiR 300 to 800 | 453.9417 | 35.5061 | 24.9885 | 1000.03 | 0.0066918 | 0.2892 | 588.781 | 635.396
Populus deltoides | LI2C |  | RACiR 300 to 800 | 456.3985 | 35.6913 | 24.9942 | 999.716 | 0.0066931 | 0.28899 | 592.045 | 638.852
Populus deltoides | LI2C |  | RACiR 300 to 800 | 459.5684 | 35.6847 | 24.9983 | 1000.09 | 0.0066953 | 0.28892 | 595.303 | 642.117
Populus deltoides | LI2C |  | RACiR 300 to 800 | 462.5413 | 35.8063 | 24.9896 | 999.872 | 0.0066963 | 0.28941 | 598.539 | 645.492
Populus deltoides | LI2C |  | RACiR 300 to 800 | 466.163 | 35.7068 | 24.9864 | 1000.12 | 0.0066976 | 0.28965 | 601.788 | 648.651
Populus deltoides | LI2C |  | RACiR 300 to 800 | 468.8076 | 35.7673 | 24.9942 | 999.725 | 0.006696 | 0.28917 | 604.931 | 651.862
Populus deltoides | LI2C |  | RACiR 300 to 800 | 470.8629 | 36.0888 | 24.991 | 1000.06 | 0.0066955 | 0.28931 | 608.087 | 655.36
Populus deltoides | LI2C |  | RACiR 300 to 800 | 473.7752 | 36.246 | 24.9814 | 999.736 | 0.0066989 | 0.28995 | 611.318 | 658.755
Populus deltoides | LI2C |  | RACiR 300 to 800 | 477.2027 | 36.3362 | 24.987 | 1000.06 | 0.0067369 | 0.29136 | 614.538 | 662.098
Populus deltoides | LI2C |  | RACiR 300 to 800 | 480.2763 | 36.5864 | 24.9917 | 999.943 | 0.0067871 | 0.29315 | 617.794 | 665.652
Populus deltoides | LI2C |  | RACiR 300 to 800 | 483.7889 | 36.4963 | 24.9944 | 1000.03 | 0.0067996 | 0.29316 | 621.094 | 668.88
Populus deltoides | LI2C |  | RACiR 300 to 800 | 486.6788 | 36.3957 | 24.9991 | 1000.03 | 0.0067831 | 0.29182 | 624.309 | 671.992
Populus deltoides | LI2C |  | RACiR 300 to 800 | 488.5754 | 36.6954 | 24.9916 | 999.693 | 0.0067696 | 0.29128 | 627.524 | 675.516
Populus deltoides | LI2C |  | RACiR 300 to 800 | 491.8295 | 36.6872 | 24.9829 | 1000.07 | 0.0067654 | 0.29135 | 630.784 | 678.776
Populus deltoides | LI2C |  | RACiR 300 to 800 | 494.9237 | 36.6672 | 24.9838 | 1000.01 | 0.006762 | 0.29108 | 634.005 | 681.988
Populus deltoides | LI2C |  | RACiR 300 to 800 | 496.6423 | 37.0572 | 24.9829 | 999.91 | 0.0067593 | 0.29094 | 637.182 | 685.565
Populus deltoides | LI2C |  | RACiR 300 to 800 | 499.5468 | 37.1597 | 24.98 | 1000.04 | 0.0067592 | 0.29106 | 640.452 | 688.946
Populus deltoides | LI2C |  | RACiR 300 to 800 | 502.8847 | 37.1487 | 24.9764 | 1000.07 | 0.0067598 | 0.29126 | 643.733 | 692.227
Populus deltoides | LI2C |  | RACiR 300 to 800 | 505.6006 | 37.302 | 24.9742 | 1000.02 | 0.00676 | 0.29139 | 646.984 | 695.651
Populus deltoides | LI2C |  | RACiR 300 to 800 | 508.8144 | 37.2794 | 24.9801 | 1000.09 | 0.0067626 | 0.29128 | 650.255 | 698.907
Populus deltoides | LI2C |  | RACiR 300 to 800 | 511.5027 | 37.4579 | 24.9798 | 999.983 | 0.0067655 | 0.29145 | 653.549 | 702.39
Populus deltoides | LI2C |  | RACiR 300 to 800 | 514.0797 | 37.651 | 24.9801 | 999.892 | 0.0067672 | 0.2915 | 656.834 | 705.886
Populus deltoides | LI2C |  | RACiR 300 to 800 | 517.5501 | 37.5467 | 24.9861 | 1000.06 | 0.0067691 | 0.29133 | 660.114 | 709.075
Populus deltoides | LI2C |  | RACiR 300 to 800 | 520.9365 | 37.4948 | 24.9874 | 999.759 | 0.0067713 | 0.29139 | 663.374 | 712.293
Populus deltoides | LI2C |  | RACiR 300 to 800 | 523.3001 | 37.7307 | 24.9916 | 1000.05 | 0.0067731 | 0.29133 | 666.647 | 715.822
Populus deltoides | LI2C |  | RACiR 300 to 800 | 525.9443 | 37.8783 | 24.9975 | 1000.03 | 0.0067749 | 0.29116 | 669.951 | 719.284
Populus deltoides | LI2C |  | RACiR 300 to 800 | 529.4871 | 37.8116 | 24.9963 | 1000.03 | 0.0067757 | 0.29129 | 673.285 | 722.565
Populus deltoides | LI2C |  | RACiR 300 to 800 | 532.315 | 37.9017 | 24.9954 | 1000.01 | 0.0067717 | 0.29117 | 676.544 | 725.926
Populus deltoides | LI2C |  | RACiR 300 to 800 | 535.3858 | 37.8214 | 24.9969 | 1000.08 | 0.0067518 | 0.29029 | 679.808 | 729.108
Populus deltoides | LI2C |  | RACiR 300 to 800 | 538.3865 | 37.8398 | 24.9953 | 999.746 | 0.0067401 | 0.29008 | 683.025 | 732.346
Populus deltoides | LI2C |  | RACiR 300 to 800 | 540.8551 | 38.1015 | 24.9954 | 1000.08 | 0.0067471 | 0.29058 | 686.221 | 735.824
Populus deltoides | LI2C |  | RACiR 300 to 800 | 543.7134 | 38.2662 | 24.9947 | 1000.03 | 0.0067531 | 0.291 | 689.522 | 739.3
Populus deltoides | LI2C |  | RACiR 300 to 800 | 547.1239 | 38.2528 | 24.9975 | 1000.08 | 0.0067598 | 0.29131 | 692.824 | 742.607
Populus deltoides | LI2C |  | RACiR 300 to 800 | 550.8352 | 38.1541 | 24.9987 | 1000.05 | 0.0067652 | 0.29157 | 696.157 | 745.86
Populus deltoides | LI2C |  | RACiR 300 to 800 | 554.5812 | 38.0306 | 24.9996 | 1000.09 | 0.0067699 | 0.2918 | 699.458 | 749.054
Populus deltoides | LI2C |  | RACiR 300 to 800 | 557.2823 | 38.1264 | 25.0049 | 1000.02 | 0.0067715 | 0.29168 | 702.618 | 752.32
Populus deltoides | LI2C |  | RACiR 300 to 800 | 559.5306 | 38.4167 | 25.0026 | 999.774 | 0.0067739 | 0.29195 | 705.795 | 755.807
Populus deltoides | LI2C |  | RACiR 300 to 800 | 562.7037 | 38.4688 | 25.0023 | 999.984 | 0.0067775 | 0.29218 | 709.117 | 759.192
Populus deltoides | LI2C |  | RACiR 300 to 800 | 566.371 | 38.3538 | 25.0045 | 999.997 | 0.0067799 | 0.29221 | 712.463 | 762.437
Populus deltoides | LI2C |  | RACiR 300 to 800 | 568.8418 | 38.5714 | 25.003 | 1000.07 | 0.0067787 | 0.2922 | 715.745 | 765.951
Populus deltoides | LI2C |  | RACiR 300 to 800 | 571.7726 | 38.6873 | 24.9985 | 1000.02 | 0.0067774 | 0.29234 | 719.072 | 769.403
Populus deltoides | LI2C |  | RACiR 300 to 800 | 575.4675 | 38.5732 | 25.0004 | 1000.02 | 0.0067781 | 0.29232 | 722.473 | 772.703
Populus deltoides | LI2C |  | RACiR 300 to 800 | 578.7085 | 38.5942 | 25.0002 | 1000.06 | 0.0067786 | 0.29237 | 725.835 | 776.103
Populus deltoides | LI2C |  | RACiR 300 to 800 | 582.6238 | 38.4543 | 24.9932 | 1000.05 | 0.0067802 | 0.29278 | 729.156 | 779.294
Populus deltoides | LI2C |  | RACiR 300 to 800 | 585.9754 | 38.396 | 24.9956 | 1000.04 | 0.0067826 | 0.29278 | 732.39 | 782.486
Populus deltoides | LI2C |  | RACiR 300 to 800 | 588.1771 | 38.6664 | 24.9993 | 999.938 | 0.0067845 | 0.29272 | 735.611 | 785.993
Populus deltoides | LI2C |  | RACiR 300 to 800 | 591.0781 | 38.7709 | 25.0002 | 1000.05 | 0.0067869 | 0.2928 | 738.905 | 789.401
Populus deltoides | LI2C |  | RACiR 300 to 800 | 594.9375 | 38.6516 | 24.9947 | 999.703 | 0.0067906 | 0.29328 | 742.209 | 792.597
Populus deltoides | LI2C |  | RACiR 0 to 500 | 104.7562 | 4.3198 | 24.9684 | 999.984 | 0.005682 | 0.24661 | 124.919 | 139.084
Populus deltoides | LI2C |  | RACiR 0 to 500 | 107.311 | 4.4146 | 24.9641 | 999.91 | 0.0056824 | 0.24679 | 127.907 | 142.174
Populus deltoides | LI2C |  | RACiR 0 to 500 | 108.9791 | 4.7261 | 24.9583 | 999.981 | 0.0056832 | 0.24704 | 130.869 | 145.456
Populus deltoides | LI2C |  | RACiR 0 to 500 | 110.913 | 4.9739 | 24.957 | 999.972 | 0.0056833 | 0.24712 | 133.855 | 148.703
Populus deltoides | LI2C |  | RACiR 0 to 500 | 113.2601 | 5.1204 | 24.9597 | 999.963 | 0.0056851 | 0.24713 | 136.856 | 151.858
Populus deltoides | LI2C |  | RACiR 0 to 500 | 114.8133 | 5.4702 | 24.9583 | 999.971 | 0.0056859 | 0.24724 | 139.867 | 155.229
Populus deltoides | LI2C |  | RACiR 0 to 500 | 117.1868 | 5.6144 | 24.9642 | 1000.01 | 0.0056865 | 0.24705 | 142.906 | 158.424
Populus deltoides | LI2C |  | RACiR 0 to 500 | 119.8145 | 5.6882 | 24.9688 | 999.964 | 0.0056895 | 0.24701 | 145.903 | 161.504
Populus deltoides | LI2C |  | RACiR 0 to 500 | 121.6808 | 5.9574 | 24.9677 | 999.934 | 0.0056926 | 0.24727 | 148.888 | 164.77
Populus deltoides | LI2C |  | RACiR 0 to 500 | 123.9168 | 6.1374 | 24.971 | 999.973 | 0.0057045 | 0.24772 | 151.868 | 167.942
Populus deltoides | LI2C |  | RACiR 0 to 500 | 126.2023 | 6.3245 | 24.9663 | 999.957 | 0.0057293 | 0.24897 | 154.839 | 171.111
Populus deltoides | LI2C |  | RACiR 0 to 500 | 127.6571 | 6.7053 | 24.9642 | 999.979 | 0.0057412 | 0.24944 | 157.824 | 174.488
Populus deltoides | LI2C |  | RACiR 0 to 500 | 129.48 | 6.9735 | 24.9661 | 999.914 | 0.0057318 | 0.24877 | 160.856 | 177.799
Populus deltoides | LI2C |  | RACiR 0 to 500 | 131.8682 | 7.1111 | 24.9655 | 1000.07 | 0.005727 | 0.2485 | 163.892 | 180.982
Populus deltoides | LI2C |  | RACiR 0 to 500 | 134.0359 | 7.3411 | 24.9634 | 999.948 | 0.0057475 | 0.24945 | 166.932 | 184.264
Populus deltoides | LI2C |  | RACiR 0 to 500 | 136.4619 | 7.5477 | 24.9563 | 999.894 | 0.0057976 | 0.25181 | 169.963 | 187.521
Populus deltoides | LI2C |  | RACiR 0 to 500 | 138.7596 | 7.7345 | 24.9562 | 999.965 | 0.0058227 | 0.25257 | 172.976 | 190.735
Populus deltoides | LI2C |  | RACiR 0 to 500 | 140.3833 | 8.0278 | 24.9608 | 999.945 | 0.0058052 | 0.25121 | 175.995 | 194.054
Populus deltoides | LI2C |  | RACiR 0 to 500 | 142.4482 | 8.2389 | 24.9551 | 1000.28 | 0.0057904 | 0.25054 | 179.048 | 197.328
Populus deltoides | LI2C |  | RACiR 0 to 500 | 144.5548 | 8.4428 | 24.9561 | 1000.01 | 0.0057828 | 0.25004 | 182.098 | 200.588
Populus deltoides | LI2C |  | RACiR 0 to 500 | 145.9139 | 8.8463 | 24.956 | 999.946 | 0.0057794 | 0.24979 | 185.159 | 204.06
Populus deltoides | LI2C |  | RACiR 0 to 500 | 148.0901 | 9.0698 | 24.9446 | 999.954 | 0.0057758 | 0.25 | 188.255 | 207.394
Populus deltoides | LI2C |  | RACiR 0 to 500 | 150.7606 | 9.1564 | 24.9483 | 999.933 | 0.0057762 | 0.24986 | 191.361 | 210.596
Populus deltoides | LI2C |  | RACiR 0 to 500 | 152.8881 | 9.366 | 24.9553 | 999.967 | 0.0057769 | 0.24961 | 194.43 | 213.885
Populus deltoides | LI2C |  | RACiR 0 to 500 | 154.8964 | 9.6031 | 24.9585 | 999.912 | 0.0057784 | 0.24957 | 197.455 | 217.158
Populus deltoides | LI2C |  | RACiR 0 to 500 | 157.077 | 9.8065 | 24.9613 | 1000.22 | 0.0057813 | 0.24964 | 200.501 | 220.416
Populus deltoides | LI2C |  | RACiR 0 to 500 | 158.8608 | 10.1264 | 24.9599 | 999.953 | 0.0057823 | 0.24978 | 203.597 | 223.845
Populus deltoides | LI2C |  | RACiR 0 to 500 | 160.7959 | 10.4072 | 24.9598 | 999.975 | 0.0057844 | 0.24995 | 206.686 | 227.226
Populus deltoides | LI2C |  | RACiR 0 to 500 | 163.0879 | 10.6033 | 24.9587 | 1000.34 | 0.0057894 | 0.25029 | 209.764 | 230.508
Populus deltoides | LI2C |  | RACiR 0 to 500 | 165.192 | 10.8608 | 24.9514 | 1000.04 | 0.0057892 | 0.25058 | 212.903 | 233.915
Populus deltoides | LI2C |  | RACiR 0 to 500 | 167.5121 | 11.0527 | 24.9478 | 999.933 | 0.0057889 | 0.25072 | 216.024 | 237.242
Populus deltoides | LI2C |  | RACiR 0 to 500 | 169.6858 | 11.2579 | 24.961 | 999.961 | 0.0057939 | 0.25047 | 219.13 | 240.562
Populus deltoides | LI2C |  | RACiR 0 to 500 | 171.5471 | 11.5561 | 24.9651 | 1000.3 | 0.0057971 | 0.2505 | 222.236 | 243.977
Populus deltoides | LI2C |  | RACiR 0 to 500 | 173.7218 | 11.7927 | 24.9561 | 999.933 | 0.0057981 | 0.25094 | 225.325 | 247.317
Populus deltoides | LI2C |  | RACiR 0 to 500 | 176.1494 | 11.9685 | 24.9531 | 1000.02 | 0.0058028 | 0.25134 | 228.439 | 250.616
Populus deltoides | LI2C |  | RACiR 0 to 500 | 177.7562 | 12.3542 | 24.951 | 999.978 | 0.0058067 | 0.25167 | 231.567 | 254.138
Populus deltoides | LI2C |  | RACiR 0 to 500 | 180.021 | 12.5521 | 24.9555 | 1000.02 | 0.0058081 | 0.25161 | 234.694 | 257.479
Populus deltoides | LI2C |  | RACiR 0 to 500 | 182.8867 | 12.5964 | 24.9647 | 999.993 | 0.0058149 | 0.25163 | 237.806 | 260.645
Populus deltoides | LI2C |  | RACiR 0 to 500 | 184.9189 | 12.8578 | 24.9655 | 999.982 | 0.0058195 | 0.25188 | 240.886 | 263.997
Populus deltoides | LI2C |  | RACiR 0 to 500 | 186.8237 | 13.1575 | 24.9634 | 999.963 | 0.0058225 | 0.25217 | 243.973 | 267.399
Populus deltoides | LI2C |  | RACiR 0 to 500 | 189.057 | 13.3729 | 24.9652 | 999.957 | 0.0058279 | 0.25243 | 247.069 | 270.719
Populus deltoides | LI2C |  | RACiR 0 to 500 | 190.8598 | 13.6817 | 24.9713 | 1000 | 0.0058294 | 0.25227 | 250.191 | 274.162
Populus deltoides | LI2C |  | RACiR 0 to 500 | 192.815 | 13.97 | 24.9748 | 999.958 | 0.0058325 | 0.25234 | 253.332 | 277.604
Populus deltoides | LI2C |  | RACiR 0 to 500 | 195.5068 | 14.0918 | 24.9722 | 999.961 | 0.0058375 | 0.25275 | 256.479 | 280.882
Populus deltoides | LI2C |  | RACiR 0 to 500 | 197.6554 | 14.3297 | 24.9751 | 999.941 | 0.0058416 | 0.25289 | 259.6 | 284.252
Populus deltoides | LI2C |  | RACiR 0 to 500 | 199.643 | 14.5809 | 24.9823 | 999.98 | 0.0058438 | 0.25274 | 262.678 | 287.596
Populus deltoides | LI2C |  | RACiR 0 to 500 | 202.0591 | 14.7513 | 24.9836 | 999.986 | 0.0058485 | 0.25297 | 265.779 | 290.876
Populus deltoides | LI2C |  | RACiR 0 to 500 | 204.1089 | 15.036 | 24.9803 | 999.993 | 0.0058544 | 0.25347 | 268.891 | 294.285
Populus deltoides | LI2C |  | RACiR 0 to 500 | 206.4221 | 15.243 | 24.9789 | 999.982 | 0.0058581 | 0.25378 | 272.009 | 297.624
Populus deltoides | LI2C |  | RACiR 0 to 500 | 208.5555 | 15.4748 | 24.9836 | 999.949 | 0.0058629 | 0.25385 | 275.1 | 300.957
Populus deltoides | LI2C |  | RACiR 0 to 500 | 210.0183 | 15.8942 | 24.9838 | 999.976 | 0.0058647 | 0.25399 | 278.233 | 304.519
Populus deltoides | LI2C |  | RACiR 0 to 500 | 212.5994 | 16.0232 | 24.9888 | 999.983 | 0.0058675 | 0.25399 | 281.392 | 307.82
Populus deltoides | LI2C |  | RACiR 0 to 500 | 215.6145 | 16.0927 | 24.9935 | 1000 | 0.0058936 | 0.25503 | 284.494 | 311.009
Populus deltoides | LI2C |  | RACiR 0 to 500 | 217.6793 | 16.4153 | 24.9884 | 999.944 | 0.005921 | 0.25639 | 287.539 | 314.395
Populus deltoides | LI2C |  | RACiR 0 to 500 | 219.8856 | 16.6242 | 24.9913 | 1000.04 | 0.0059259 | 0.25635 | 290.637 | 317.718
Populus deltoides | LI2C |  | RACiR 0 to 500 | 221.7967 | 16.8832 | 24.9935 | 1000.02 | 0.0059211 | 0.25594 | 293.726 | 321.072
Populus deltoides | LI2C |  | RACiR 0 to 500 | 223.225 | 17.2737 | 24.9992 | 999.905 | 0.0059195 | 0.25557 | 296.837 | 324.583
Populus deltoides | LI2C |  | RACiR 0 to 500 | 225.7686 | 17.5043 | 24.9965 | 999.954 | 0.0059467 | 0.25688 | 299.997 | 327.997
Populus deltoides | LI2C |  | RACiR 0 to 500 | 228.9302 | 17.6395 | 24.9926 | 999.96 | 0.0059991 | 0.25918 | 303.146 | 331.303
Populus deltoides | LI2C |  | RACiR 0 to 500 | 231.3064 | 17.8438 | 24.9974 | 999.971 | 0.006019 | 0.25947 | 306.311 | 334.691
Populus deltoides | LI2C |  | RACiR 0 to 500 | 233.2811 | 18.0432 | 25.0009 | 999.986 | 0.0060051 | 0.25838 | 309.419 | 338.009
Populus deltoides | LI2C |  | RACiR 0 to 500 | 235.5169 | 18.2195 | 24.9934 | 999.975 | 0.0059936 | 0.25798 | 312.508 | 341.279
Populus deltoides | LI2C |  | RACiR 0 to 500 | 237.4802 | 18.4938 | 24.984 | 999.966 | 0.0059864 | 0.25792 | 315.602 | 344.656
Populus deltoides | LI2C |  | RACiR 0 to 500 | 239.5081 | 18.7336 | 24.9848 | 1000.29 | 0.0059828 | 0.25767 | 318.692 | 347.999
Populus deltoides | LI2C |  | RACiR 0 to 500 | 241.8992 | 18.8697 | 24.9927 | 999.953 | 0.0059821 | 0.25729 | 321.793 | 351.244
Populus deltoides | LI2C |  | RACiR 0 to 500 | 243.7646 | 19.1682 | 24.9961 | 999.983 | 0.0059842 | 0.25727 | 324.885 | 354.646
Populus deltoides | LI2C |  | RACiR 0 to 500 | 246.1266 | 19.3521 | 24.9921 | 999.971 | 0.0059836 | 0.2574 | 327.982 | 357.939
Populus deltoides | LI2C |  | RACiR 0 to 500 | 248.4547 | 19.5269 | 24.9939 | 1000.02 | 0.0059848 | 0.25739 | 331.057 | 361.197
Populus deltoides | LI2C |  | RACiR 0 to 500 | 250.2556 | 19.8523 | 24.9967 | 1000.02 | 0.0059876 | 0.25746 | 334.158 | 364.638
Populus deltoides | LI2C |  | RACiR 0 to 500 | 252.5696 | 20.0564 | 24.9963 | 999.982 | 0.0059895 | 0.25762 | 337.28 | 367.979
Populus deltoides | LI2C |  | RACiR 0 to 500 | 254.8729 | 20.2598 | 24.9998 | 999.976 | 0.0059937 | 0.25774 | 340.4 | 371.31
Populus deltoides | LI2C |  | RACiR 0 to 500 | 256.549 | 20.6501 | 24.9982 | 999.962 | 0.0059981 | 0.2581 | 343.531 | 374.843
Populus deltoides | LI2C |  | RACiR 0 to 500 | 258.6759 | 20.9094 | 24.9987 | 999.992 | 0.0060002 | 0.25815 | 346.707 | 378.294
Populus deltoides | LI2C |  | RACiR 0 to 500 | 261.4518 | 21.0177 | 25.0003 | 1000.01 | 0.0060042 | 0.25827 | 349.937 | 381.643
Populus deltoides | LI2C |  | RACiR 0 to 500 | 263.9174 | 21.2175 | 24.9977 | 999.99 | 0.0060087 | 0.25867 | 353.109 | 385.029
Populus deltoides | LI2C |  | RACiR 0 to 500 | 266.4962 | 21.3408 | 25.003 | 999.967 | 0.0060112 | 0.25866 | 356.238 | 388.297
Populus deltoides | LI2C |  | RACiR 0 to 500 | 269.1422 | 21.4675 | 25.0045 | 1000 | 0.0060164 | 0.25891 | 359.358 | 391.553
Populus deltoides | LI2C |  | RACiR 0 to 500 | 271.1777 | 21.746 | 25.0069 | 999.941 | 0.0060205 | 0.25906 | 362.482 | 394.968
Populus deltoides | LI2C |  | RACiR 0 to 500 | 273.2403 | 22.0032 | 25.0072 | 1000.01 | 0.0060211 | 0.25909 | 365.589 | 398.346
Populus deltoides | LI2C |  | RACiR 0 to 500 | 275.6134 | 22.1842 | 25.0086 | 1000.06 | 0.006022 | 0.25908 | 368.729 | 401.677
Populus deltoides | LI2C |  | RACiR 0 to 500 | 277.2921 | 22.5326 | 25.0133 | 999.996 | 0.0060214 | 0.25885 | 371.891 | 405.197
Populus deltoides | LI2C |  | RACiR 0 to 500 | 279.9673 | 22.6595 | 25.0059 | 999.988 | 0.0060184 | 0.25899 | 375.069 | 408.517
Populus deltoides | LI2C |  | RACiR 0 to 500 | 282.809 | 22.7328 | 25.0014 | 999.971 | 0.0060192 | 0.25923 | 378.177 | 411.707
Populus deltoides | LI2C |  | RACiR 0 to 500 | 284.9927 | 22.9594 | 25.0034 | 999.962 | 0.0060197 | 0.2592 | 381.307 | 415.075
Populus deltoides | LI2C |  | RACiR 0 to 500 | 287.5065 | 23.1265 | 24.9979 | 1000.31 | 0.0060209 | 0.25954 | 384.406 | 418.355
Populus deltoides | LI2C |  | RACiR 0 to 500 | 289.7635 | 23.3605 | 24.9928 | 999.977 | 0.0060247 | 0.25998 | 387.465 | 421.655
Populus deltoides | LI2C |  | RACiR 0 to 500 | 291.3534 | 23.7403 | 24.9972 | 999.942 | 0.0060268 | 0.25991 | 390.6 | 425.184
Populus deltoides | LI2C |  | RACiR 0 to 500 | 293.9157 | 23.914 | 24.9915 | 1000 | 0.0060279 | 0.26023 | 393.777 | 428.55
Populus deltoides | LI2C |  | RACiR 0 to 500 | 296.6686 | 24.0512 | 24.9835 | 1000.05 | 0.00603 | 0.26071 | 396.941 | 431.861
Populus deltoides | LI2C |  | RACiR 0 to 500 | 298.9537 | 24.2853 | 24.9832 | 999.968 | 0.0060316 | 0.26083 | 400.145 | 435.31
Populus deltoides | LI2C |  | RACiR 0 to 500 | 301.9781 | 24.3483 | 24.9757 | 999.993 | 0.0060329 | 0.26123 | 403.329 | 438.574
Populus deltoides | LI2C |  | RACiR 0 to 500 | 304.4862 | 24.4861 | 24.9805 | 1000.02 | 0.0060366 | 0.26124 | 406.431 | 441.822
Populus deltoides | LI2C |  | RACiR 0 to 500 | 306.4277 | 24.7854 | 24.9822 | 999.998 | 0.0060389 | 0.2613 | 409.554 | 445.259
Populus deltoides | LI2C |  | RACiR 0 to 500 | 309.8087 | 24.7417 | 24.9759 | 999.942 | 0.0060416 | 0.26175 | 412.672 | 448.345
Populus deltoides | LI2C |  | RACiR 0 to 500 | 312.2488 | 24.8758 | 24.9812 | 999.981 | 0.0060438 | 0.26166 | 415.725 | 451.542
Populus deltoides | LI2C |  | RACiR 0 to 500 | 313.1009 | 25.4713 | 24.9804 | 1000 | 0.0060459 | 0.26184 | 418.829 | 455.254
Populus deltoides | LI2C |  | RACiR 0 to 500 | 315.5582 | 25.675 | 24.9748 | 1000.28 | 0.0060461 | 0.2621 | 422.023 | 458.669
Populus deltoides | LI2C |  | RACiR 0 to 500 | 318.8196 | 25.6826 | 24.9746 | 999.916 | 0.0060549 | 0.26251 | 425.237 | 461.9
Populus deltoides | LI2C |  | RACiR 0 to 500 | 321.6335 | 25.9094 | 24.9671 | 999.944 | 0.0060832 | 0.26407 | 428.387 | 465.302
Populus deltoides | LI2C |  | RACiR 0 to 500 | 324.2932 | 26.0684 | 24.9688 | 999.983 | 0.0060993 | 0.26455 | 431.546 | 468.645
Populus deltoides | LI2C |  | RACiR 0 to 500 | 326.7442 | 26.1865 | 24.9674 | 999.945 | 0.0060911 | 0.26403 | 434.709 | 471.936
Populus deltoides | LI2C |  | RACiR 0 to 500 | 328.4943 | 26.5313 | 24.9599 | 999.969 | 0.0060845 | 0.26389 | 437.869 | 475.443
Populus deltoides | LI2C |  | RACiR 0 to 500 | 330.4533 | 26.8017 | 24.964 | 999.95 | 0.0060819 | 0.26352 | 441.061 | 478.911
Populus deltoides | LI2C |  | RACiR 0 to 500 | 333.2924 | 26.8606 | 24.9675 | 1000.02 | 0.0060811 | 0.26335 | 444.265 | 482.186
Populus deltoides | LI2C |  | RACiR 0 to 500 | 335.9318 | 27.0248 | 24.9594 | 999.955 | 0.006082 | 0.26376 | 447.422 | 485.523
Populus deltoides | LI2C |  | RACiR 0 to 500 | 338.7932 | 27.1532 | 24.9485 | 999.964 | 0.0060846 | 0.26437 | 450.584 | 488.825
Populus deltoides | LI2C | 3 | Standard | 278.3112 | 22.8388 | 24.9825 | 999.73 | 0.0058704 | 0.25613 | 374.966 | 400.008
Populus deltoides | LI2C | 3 | Standard | 242.3442 | 20.4396 | 24.9981 | 999.963 | 0.0059787 | 0.25965 | 327.625 | 350.023
Populus deltoides | LI2C | 3 | Standard | 209.1915 | 17.5051 | 25.0051 | 1000.03 | 0.0061006 | 0.26524 | 280.824 | 300.042
Populus deltoides | LI2C | 3 | Standard | 176.5242 | 14.3486 | 25.0033 | 1000.03 | 0.0062025 | 0.27081 | 234.212 | 250.014
Populus deltoides | LI2C | 3 | Standard | 144.3802 | 11.0329 | 25.0004 | 999.973 | 0.0063366 | 0.27822 | 187.819 | 200.042
Populus deltoides | LI2C | 3 | Standard | 112.697 | 7.4489 | 24.9949 | 999.942 | 0.0065585 | 0.2861 | 141.648 | 150.026
Populus deltoides | LI2C | 3 | Standard | 80.7458 | 3.7907 | 25.0052 | 1000.08 | 0.0067568 | 0.29477 | 95.627 | 100.064
Populus deltoides | LI2C | 3 | Standard | 48.9679 | -0.10032 | 25.0218 | 999.886 | 0.0068312 | 0.29911 | 49.7599 | 49.9995
Populus deltoides | LI2C | 3 | Standard | 285.645 | 24.6404 | 25 | 999.858 | 0.0070855 | 0.30958 | 372.772 | 400.053
Populus deltoides | LI2C | 3 | Standard | 360.4287 | 30.5871 | 25.0184 | 999.96 | 0.0072362 | 0.31774 | 466.105 | 500.064
Populus deltoides | LI2C | 3 | Standard | 439.0918 | 35.7087 | 24.9864 | 1000.01 | 0.007441 | 0.32569 | 560.148 | 600.024
Populus deltoides | LI2C | 3 | Standard | 524.8323 | 38.9593 | 25.0065 | 1000.04 | 0.0075654 | 0.33064 | 656.174 | 700.094
Populus deltoides | LI2C | 3 | Standard | 615.4996 | 40.4532 | 25.0365 | 999.962 | 0.0075215 | 0.32937 | 753.954 | 800.077
Populus deltoides | LI2C | 3 | Standard | 708.213 | 41.2015 | 25.0361 | 1000.03 | 0.0074367 | 0.32565 | 852.556 | 900.096
Populus deltoides | LI2C | 3 | Standard | 802.5119 | 41.269 | 25.0376 | 999.955 | 0.0072883 | 0.31911 | 951.872 | 1000.08
Populus deltoides | LI2C | 3 | Standard | 897.0344 | 41.0119 | 25.0293 | 1000.04 | 0.0070921 | 0.31012 | 1051.56 | 1100.03
Populus deltoides | LI2C | 3 | Standard | 992.3713 | 40.7205 | 25.0227 | 999.981 | 0.006915 | 0.30274 | 1151.36 | 1200.04
Populus deltoides | LI2C | 3 | Standard | 1084.9318 | 40.9137 | 25.0144 | 1000.24 | 0.0067278 | 0.29412 | 1250.75 | 1300.08
Populus deltoides | LI2C | 3 | Standard | 1181.1001 | 40.637 | 25.0197 | 999.962 | 0.0066043 | 0.28932 | 1350.45 | 1400.01
Populus deltoides | LI2C | 3 | Standard | 1275.4962 | 40.7626 | 25.012 | 999.937 | 0.0065495 | 0.28509 | 1449.78 | 1500.04
Populus deltoides | LI2C | 3 | Standard | 1567.655 | 40.1753 | 25.0282 | 1000.23 | 0.0064427 | 0.28107 | 1748.6 | 1800.04
Populus deltoides | LI2C | 3 | Standard | 1760.9286 | 40.2162 | 25.0214 | 1000.23 | 0.0063747 | 0.27904 | 1947.41 | 2000.04
Populus deltoides | LI2C | 2 | Standard | 250.0519 | 15.0025 | 25.035 | 1000.06 | 0.0027512 | 0.11854 | 383.966 | 400.024
Populus deltoides | LI2C | 2 | Standard | 219.8801 | 13.0581 | 25.0106 | 999.973 | 0.002746 | 0.11897 | 336.06 | 350.041
Populus deltoides | LI2C | 2 | Standard | 192.7853 | 11.0588 | 25.0002 | 999.938 | 0.0028351 | 0.1231 | 288.157 | 300.033
Populus deltoides | LI2C | 2 | Standard | 164.8764 | 8.9782 | 25.0142 | 999.911 | 0.0029238 | 0.12681 | 240.346 | 250.027
Populus deltoides | LI2C | 2 | Standard | 135.6678 | 6.9617 | 25.0012 | 999.954 | 0.0030264 | 0.13128 | 192.48 | 200.024
Populus deltoides | LI2C | 2 | Standard | 106.5564 | 4.8541 | 24.9972 | 999.962 | 0.0031768 | 0.13773 | 144.697 | 150.011
Populus deltoides | LI2C | 2 | Standard | 78.1595 | 2.5178 | 24.9905 | 999.894 | 0.0034094 | 0.14789 | 97.2058 | 100.055
Populus deltoides | LI2C | 2 | Standard | 49.4122 | -0.10157 | 24.9918 | 999.878 | 0.0036557 | 0.15941 | 49.914 | 49.9949
Populus deltoides | LI2C | 2 | Standard | 262.4 | 19.8448 | 24.9796 | 999.997 | 0.0042074 | 0.18245 | 378.558 | 399.994
Populus deltoides | LI2C | 2 | Standard | 331.9254 | 26.756 | 24.9487 | 999.933 | 0.0047153 | 0.20594 | 471.037 | 500.013
Populus deltoides | LI2C | 2 | Standard | 408.6608 | 32.939 | 24.981 | 999.971 | 0.0052409 | 0.22831 | 564.098 | 599.995
Populus deltoides | LI2C | 2 | Standard | 492.6003 | 37.5984 | 24.9588 | 1000 | 0.0056394 | 0.24582 | 658.757 | 700.071
Populus deltoides | LI2C | 2 | Standard | 582.9148 | 40.0392 | 24.9971 | 999.945 | 0.0058463 | 0.2546 | 755.545 | 800
Populus deltoides | LI2C | 2 | Standard | 677.1341 | 41.5205 | 25.0135 | 1000.26 | 0.0059898 | 0.26152 | 853.427 | 900.059
Populus deltoides | LI2C | 2 | Standard | 773.3329 | 41.9938 | 25.0314 | 999.901 | 0.0060431 | 0.26379 | 952.293 | 1000.04
Populus deltoides | LI2C | 2 | Standard | 868.0944 | 42.2037 | 25.0222 | 1000.17 | 0.0060072 | 0.26146 | 1051.56 | 1100.08
Populus deltoides | LI2C | 2 | Standard | 964.0345 | 42.2869 | 25.0183 | 999.821 | 0.0059534 | 0.26 | 1150.89 | 1200.03
Populus deltoides | LI2C | 2 | Standard | 1059.867 | 42.4412 | 25.0037 | 1000.12 | 0.0059429 | 0.25907 | 1250.18 | 1300.05
Populus deltoides | LI2C | 2 | Standard | 1156.8308 | 42.4956 | 24.996 | 999.853 | 0.0059462 | 0.25918 | 1349.54 | 1400.06
Populus deltoides | LI2C | 2 | Standard | 1255.1364 | 42.4007 | 24.9958 | 1000.13 | 0.0059951 | 0.26061 | 1448.94 | 1500.03
Populus deltoides | LI2C | 2 | Standard | 1549.5658 | 41.7143 | 25.0086 | 1000.04 | 0.0059898 | 0.2604 | 1747.68 | 1799.86
Populus deltoides | LI2C | 2 | Standard | 1741.5877 | 41.9322 | 25.0148 | 999.793 | 0.0059452 | 0.2583 | 1946.37 | 1999.88
Populus deltoides | LI2C | 1 | Standard | 303.1934 | 26.3394 | 25.0554 | 999.872 | 0.010144 | 0.44605 | 369.897 | 399.987
Populus deltoides | LI2C | 1 | Standard | 264.4966 | 23.2227 | 25.0466 | 999.892 | 0.010047 | 0.44295 | 323.593 | 350.067
Populus deltoides | LI2C | 1 | Standard | 226.2712 | 19.675 | 25.0393 | 999.803 | 0.0098003 | 0.43097 | 277.654 | 300.05
Populus deltoides | LI2C | 1 | Standard | 189.4182 | 15.9419 | 25.011 | 999.782 | 0.0095651 | 0.42306 | 231.863 | 250.022
Populus deltoides | LI2C | 1 | Standard | 153.4376 | 11.986 | 25.0482 | 1000.09 | 0.0094407 | 0.41393 | 186.268 | 200.012
Populus deltoides | LI2C | 1 | Standard | 118.4123 | 7.9153 | 25.0255 | 999.719 | 0.0092993 | 0.40759 | 140.789 | 150.014
Populus deltoides | LI2C | 1 | Standard | 83.9121 | 3.7942 | 25.0182 | 1000.03 | 0.0091775 | 0.40285 | 95.3726 | 100.042
Populus deltoides | LI2C | 1 | Standard | 49.7961 | -0.39612 | 25.0211 | 999.958 | 0.0090659 | 0.39768 | 49.9368 | 49.9934
Populus deltoides | LI2C | 1 | Standard | 299.9711 | 25.1032 | 25.0251 | 1000.05 | 0.0089831 | 0.39267 | 371.582 | 400.023
Populus deltoides | LI2C | 1 | Standard | 377.7286 | 30.5238 | 25.0367 | 1000.05 | 0.0089699 | 0.3916 | 465.331 | 500.026
Populus deltoides | LI2C | 1 | Standard | 456.5637 | 35.2164 | 25.0438 | 999.747 | 0.0087215 | 0.38353 | 559.943 | 600.044
Populus deltoides | LI2C | 1 | Standard | 542.5203 | 37.6694 | 25.0363 | 1000.14 | 0.0085155 | 0.37434 | 656.79 | 700.052
Populus deltoides | LI2C | 1 | Standard | 632.4606 | 38.7729 | 25.0522 | 999.75 | 0.0082928 | 0.36327 | 755.03 | 800.06
Populus deltoides | LI2C | 1 | Standard | 725.1742 | 39.0386 | 25.02 | 1000.12 | 0.0080901 | 0.35263 | 853.997 | 899.944
Populus deltoides | LI2C | 1 | Standard | 818.5975 | 39.1541 | 25.0245 | 999.99 | 0.0078192 | 0.34185 | 953.4 | 1000.01
Populus deltoides | LI2C | 1 | Standard | 913.082 | 38.8554 | 25.032 | 1000.05 | 0.0076106 | 0.33073 | 1053.19 | 1100.06
Populus deltoides | LI2C | 1 | Standard | 1006.6889 | 38.7074 | 25.0301 | 1000.05 | 0.0073331 | 0.31872 | 1152.98 | 1200.14
Populus deltoides | LI2C | 1 | Standard | 1100.0757 | 38.5652 | 25.0442 | 1000.02 | 0.0071087 | 0.30783 | 1252.52 | 1299.99
Populus deltoides | LI2C | 1 | Standard | 1193.047 | 38.4693 | 25.0301 | 1000.03 | 0.0068345 | 0.29636 | 1352.2 | 1399.91
Populus deltoides | LI2C | 1 | Standard | 1285.8515 | 38.5305 | 25.0235 | 999.997 | 0.0065968 | 0.28631 | 1451.97 | 1500.08
Populus deltoides | LI2C | 1 | Standard | 1574.0534 | 37.8966 | 25.0274 | 999.998 | 0.006305 | 0.27321 | 1751.13 | 1800.07
Populus deltoides | LI2C | 1 | Standard | 1763.559 | 37.9622 | 25.0302 | 999.963 | 0.0060961 | 0.26407 | 1950.18 | 2000.03
Populus deltoides | LI1A |  | RACiR 500 to 0 | 335.8435 | 35.4898 | 25.0125 | 1000.08 | 0.009456 | 0.41975 | 429.009 | 459.451
Populus deltoides | LI1A |  | RACiR 500 to 0 | 333.6684 | 35.4149 | 25.0142 | 999.773 | 0.0094419 | 0.41898 | 426.764 | 457.084
Populus deltoides | LI1A |  | RACiR 500 to 0 | 332.101 | 35.347 | 25.0196 | 1000.09 | 0.0094303 | 0.41813 | 425.177 | 455.413
Populus deltoides | LI1A |  | RACiR 500 to 0 | 330.1838 | 35.317 | 25.0214 | 999.776 | 0.0094506 | 0.41909 | 422.946 | 453.111
Populus deltoides | LI1A |  | RACiR 500 to 0 | 328.9015 | 34.9956 | 25.0215 | 1000.08 | 0.009467 | 0.41984 | 420.707 | 450.533
Populus deltoides | LI1A |  | RACiR 500 to 0 | 327.6134 | 34.9117 | 25.0228 | 1000.07 | 0.009481 | 0.42031 | 419.097 | 448.816
Populus deltoides | LI1A |  | RACiR 500 to 0 | 326.0873 | 34.565 | 25.0202 | 1000.07 | 0.0094629 | 0.41948 | 416.867 | 446.196
Populus deltoides | LI1A |  | RACiR 500 to 0 | 323.6711 | 34.5432 | 25.0188 | 1000.08 | 0.0094349 | 0.4182 | 414.598 | 443.853
Populus deltoides | LI1A |  | RACiR 500 to 0 | 322.302 | 34.3963 | 25.0198 | 999.774 | 0.0094205 | 0.41755 | 412.974 | 442.046
Populus deltoides | LI1A |  | RACiR 500 to 0 | 320.0254 | 34.3267 | 25.0282 | 1000.07 | 0.0094062 | 0.41648 | 410.698 | 439.67
Populus deltoides | LI1A |  | RACiR 500 to 0 | 318.2131 | 34.1846 | 25.0355 | 1000.07 | 0.0094176 | 0.41673 | 408.454 | 437.242
Populus deltoides | LI1A |  | RACiR 500 to 0 | 317.0624 | 34.0125 | 25.0372 | 1000.07 | 0.009417 | 0.4167 | 406.865 | 435.465
Populus deltoides | LI1A |  | RACiR 500 to 0 | 315.4612 | 33.8456 | 25.0312 | 1000.07 | 0.0094254 | 0.41763 | 404.63 | 433.018
Populus deltoides | LI1A |  | RACiR 500 to 0 | 313.8818 | 33.6285 | 25.0278 | 1000.07 | 0.0094367 | 0.41846 | 402.32 | 430.467
Populus deltoides | LI1A |  | RACiR 500 to 0 | 312.2844 | 33.6531 | 25.0315 | 999.751 | 0.0094485 | 0.41879 | 400.685 | 428.843
Populus deltoides | LI1A |  | RACiR 500 to 0 | 310.555 | 33.4683 | 25.0404 | 1000.06 | 0.0094715 | 0.41919 | 398.405 | 426.344
Populus deltoides | LI1A |  | RACiR 500 to 0 | 308.6394 | 33.343 | 25.0405 | 1000.05 | 0.0094755 | 0.41912 | 396.161 | 423.944
Populus deltoides | LI1A |  | RACiR 500 to 0 | 307.4378 | 33.16 | 25.0385 | 1000.06 | 0.0094691 | 0.4188 | 394.553 | 422.116
Populus deltoides | LI1A |  | RACiR 500 to 0 | 305.441 | 32.9815 | 25.04 | 1000.06 | 0.0094518 | 0.41767 | 392.301 | 419.641
Populus deltoides | LI1A |  | RACiR 500 to 0 | 303.0915 | 32.9389 | 25.0447 | 1000.05 | 0.0094383 | 0.41664 | 390 | 417.258
Populus deltoides | LI1A |  | RACiR 500 to 0 | 301.7645 | 32.828 | 25.0437 | 1000.06 | 0.0094333 | 0.41651 | 388.397 | 415.522
Populus deltoides | LI1A |  | RACiR 500 to 0 | 299.6621 | 32.7681 | 25.0424 | 1000.06 | 0.0094301 | 0.41649 | 386.103 | 413.13
Populus deltoides | LI1A |  | RACiR 500 to 0 | 298.0341 | 32.4643 | 25.0394 | 1000.05 | 0.0094055 | 0.4156 | 383.865 | 410.55
Populus deltoides | LI1A |  | RACiR 500 to 0 | 296.2932 | 32.4907 | 25.0397 | 999.747 | 0.0093952 | 0.41521 | 382.221 | 408.891
Populus deltoides | LI1A |  | RACiR 500 to 0 | 294.424 | 32.3419 | 25.037 | 1000.06 | 0.009384 | 0.41504 | 379.972 | 406.455
Populus deltoides | LI1A |  | RACiR 500 to 0 | 292.7564 | 32.2441 | 25.0313 | 1000.05 | 0.009402 | 0.41645 | 377.751 | 404.111
Populus deltoides | LI1A |  | RACiR 500 to 0 | 291.7676 | 32.0594 | 25.0322 | 999.753 | 0.0094173 | 0.41713 | 376.164 | 402.319
Populus deltoides | LI1A |  | RACiR 500 to 0 | 289.927 | 31.9426 | 25.033 | 1000.05 | 0.0094298 | 0.4176 | 373.912 | 399.918
Populus deltoides | LI1A |  | RACiR 500 to 0 | 288.0693 | 31.7558 | 25.0362 | 1000.05 | 0.0094289 | 0.41722 | 371.637 | 397.415
Populus deltoides | LI1A |  | RACiR 500 to 0 | 286.8581 | 31.6086 | 25.0334 | 1000.05 | 0.0094257 | 0.41718 | 370.046 | 395.656
Populus deltoides | LI1A |  | RACiR 500 to 0 | 284.8347 | 31.493 | 25.0367 | 999.738 | 0.0094214 | 0.41666 | 367.798 | 393.248
Populus deltoides | LI1A |  | RACiR 500 to 0 | 283.0338 | 31.2758 | 25.0361 | 1000.05 | 0.0094041 | 0.4159 | 365.568 | 390.77
Populus deltoides | LI1A |  | RACiR 500 to 0 | 281.3057 | 31.3377 | 25.0354 | 1000.04 | 0.0093991 | 0.41579 | 363.969 | 389.196
Populus deltoides | LI1A |  | RACiR 500 to 0 | 279.5496 | 31.173 | 25.0329 | 1000.05 | 0.0093931 | 0.41582 | 361.761 | 386.793
Populus deltoides | LI1A |  | RACiR 500 to 0 | 277.7065 | 31.0415 | 25.038 | 999.741 | 0.009397 | 0.41578 | 359.566 | 384.429
Populus deltoides | LI1A |  | RACiR 500 to 0 | 276.6435 | 30.8486 | 25.0404 | 1000.04 | 0.0094015 | 0.41583 | 358.004 | 382.648
Populus deltoides | LI1A |  | RACiR 500 to 0 | 275.0465 | 30.6326 | 25.0393 | 1000.05 | 0.0094113 | 0.41628 | 355.763 | 380.168
Populus deltoides | LI1A |  | RACiR 500 to 0 | 273.1484 | 30.5521 | 25.0382 | 1000.04 | 0.009425 | 0.41683 | 353.53 | 377.811
Populus deltoides | LI1A |  | RACiR 500 to 0 | 271.7958 | 30.3965 | 25.0392 | 1000.04 | 0.0094134 | 0.41604 | 351.914 | 376.015
Populus deltoides | LI1A |  | RACiR 500 to 0 | 269.706 | 30.2646 | 25.0363 | 1000.04 | 0.0093847 | 0.4148 | 349.674 | 373.599
Populus deltoides | LI1A |  | RACiR 500 to 0 | 268.3104 | 29.8712 | 25.0312 | 1000.04 | 0.0093581 | 0.41399 | 347.425 | 370.912
Populus deltoides | LI1A |  | RACiR 500 to 0 | 266.7119 | 29.9053 | 25.0263 | 1000.05 | 0.0093568 | 0.41445 | 345.783 | 369.277
Populus deltoides | LI1A |  | RACiR 500 to 0 | 264.8846 | 29.7609 | 25.0277 | 1000.04 | 0.0093603 | 0.41468 | 343.52 | 366.837
Populus deltoides | LI1A |  | RACiR 500 to 0 | 262.5998 | 29.8068 | 25.0316 | 1000.04 | 0.0093683 | 0.41489 | 341.26 | 364.582
Populus deltoides | LI1A |  | RACiR 500 to 0 | 261.4616 | 29.617 | 25.0358 | 1000.03 | 0.0093674 | 0.41458 | 339.692 | 362.807
Populus deltoides | LI1A |  | RACiR 500 to 0 | 259.9624 | 29.36 | 25.0332 | 1000.03 | 0.0093688 | 0.41487 | 337.478 | 360.294
Populus deltoides | LI1A |  | RACiR 500 to 0 | 257.4258 | 29.1884 | 25.0292 | 1000.03 | 0.0093673 | 0.41498 | 334.442 | 356.957
Populus deltoides | LI1A |  | RACiR 500 to 0 | 256.4397 | 29.0402 | 25.0278 | 1000.03 | 0.009373 | 0.41541 | 332.997 | 355.431
Populus deltoides | LI1A |  | RACiR 500 to 0 | 254.7844 | 29.1392 | 25.028 | 999.715 | 0.0093934 | 0.4163 | 331.394 | 353.9
Populus deltoides | LI1A |  | RACiR 500 to 0 | 252.9469 | 29.042 | 25.0289 | 1000.02 | 0.0094073 | 0.41665 | 329.223 | 351.612
Populus deltoides | LI1A |  | RACiR 500 to 0 | 251.5445 | 29.017 | 25.0265 | 1000.02 | 0.0094133 | 0.41689 | 327.688 | 350.028
Populus deltoides | LI1A |  | RACiR 500 to 0 | 250.144 | 28.678 | 25.0276 | 1000.03 | 0.0094022 | 0.41601 | 325.585 | 347.546
Populus deltoides | LI1A |  | RACiR 500 to 0 | 248.6715 | 28.3477 | 25.0277 | 1000.02 | 0.009386 | 0.41513 | 323.426 | 345.017
Populus deltoides | LI1A |  | RACiR 500 to 0 | 247.796 | 28.0249 | 25.0296 | 1000.01 | 0.0093771 | 0.41456 | 321.84 | 343.073
Populus deltoides | LI1A |  | RACiR 500 to 0 | 246.0498 | 27.7587 | 25.0229 | 1000.01 | 0.0093504 | 0.41376 | 319.53 | 340.455
Populus deltoides | LI1A |  | RACiR 500 to 0 | 244.3129 | 27.4986 | 25.0173 | 1000.01 | 0.0093276 | 0.41317 | 317.206 | 337.826
Populus deltoides | LI1A |  | RACiR 500 to 0 | 242.9373 | 27.3775 | 25.0137 | 1000.02 | 0.0093121 | 0.41289 | 315.542 | 336.013
Populus deltoides | LI1A |  | RACiR 500 to 0 | 240.9275 | 27.2962 | 25.015 | 1000.01 | 0.0093165 | 0.4133 | 313.217 | 333.566
Populus deltoides | LI1A |  | RACiR 500 to 0 | 239.1234 | 27.1175 | 25.0168 | 1000.01 | 0.0093143 | 0.41335 | 310.925 | 331.066
Populus deltoides | LI1A |  | RACiR 500 to 0 | 237.3238 | 27.2089 | 25.0189 | 1000.02 | 0.0093187 | 0.41347 | 309.289 | 329.491
Populus deltoides | LI1A |  | RACiR 500 to 0 | 235.7166 | 26.9895 | 25.0212 | 999.704 | 0.0093262 | 0.41372 | 307.07 | 327.023
Populus deltoides | LI1A |  | RACiR 500 to 0 | 233.8396 | 26.9206 | 25.0216 | 1000 | 0.0093469 | 0.41462 | 304.84 | 324.691
Populus deltoides | LI1A |  | RACiR 500 to 0 | 233.0283 | 26.6479 | 25.023 | 1000 | 0.009363 | 0.41518 | 303.26 | 322.815
Populus deltoides | LI1A |  | RACiR 500 to 0 | 231.4032 | 26.4181 | 25.0203 | 1000 | 0.0093675 | 0.41534 | 301.013 | 320.309
Populus deltoides | LI1A |  | RACiR 500 to 0 | 229.4564 | 26.2627 | 25.0243 | 999.989 | 0.0093644 | 0.4147 | 298.75 | 317.849
Populus deltoides | LI1A |  | RACiR 500 to 0 | 228.0288 | 26.1322 | 25.026 | 999.992 | 0.0093464 | 0.41361 | 297.144 | 316.089
Populus deltoides | LI1A |  | RACiR 500 to 0 | 225.9797 | 26.0399 | 25.0235 | 999.991 | 0.0093316 | 0.41306 | 294.905 | 313.711
Populus deltoides | LI1A |  | RACiR 500 to 0 | 224.3636 | 25.7981 | 25.0127 | 999.995 | 0.0093131 | 0.41295 | 292.666 | 311.193
Populus deltoides | LI1A |  | RACiR 500 to 0 | 222.8121 | 25.8258 | 25.0036 | 999.996 | 0.0093082 | 0.41339 | 291.07 | 309.601
Populus deltoides | LI1A |  | RACiR 500 to 0 | 221.054 | 25.6616 | 25.0056 | 999.978 | 0.0093125 | 0.41355 | 288.847 | 307.175
Populus deltoides | LI1A |  | RACiR 500 to 0 | 219.1937 | 25.5304 | 25.0065 | 999.985 | 0.0093093 | 0.41337 | 286.652 | 304.815
Populus deltoides | LI1A |  | RACiR 500 to 0 | 218.184 | 25.3011 | 25.01 | 999.991 | 0.0093112 | 0.41323 | 285.081 | 302.988
Populus deltoides | LI1A |  | RACiR 500 to 0 | 216.5398 | 25.0748 | 25.0117 | 999.99 | 0.0093096 | 0.413 | 282.883 | 300.526
Populus deltoides | LI1A |  | RACiR 500 to 0 | 214.8366 | 24.8601 | 25.0139 | 999.988 | 0.0093059 | 0.4127 | 280.661 | 298.058
Populus deltoides | LI1A |  | RACiR 500 to 0 | 213.6729 | 24.702 | 25.0137 | 999.982 | 0.0093061 | 0.41274 | 279.077 | 296.29
Populus deltoides | LI1A |  | RACiR 500 to 0 | 211.7441 | 24.6329 | 25.0094 | 999.981 | 0.009314 | 0.41342 | 276.831 | 293.937
Populus deltoides | LI1A |  | RACiR 500 to 0 | 210.2317 | 24.3732 | 25.0057 | 999.988 | 0.009311 | 0.4135 | 274.634 | 291.453
Populus deltoides | LI1A |  | RACiR 500 to 0 | 208.6691 | 24.3947 | 25.0049 | 999.977 | 0.0093163 | 0.41374 | 273.054 | 289.861
Populus deltoides | LI1A |  | RACiR 500 to 0 | 206.9881 | 24.1587 | 25.0055 | 1000 | 0.0093016 | 0.41297 | 270.87 | 287.407
Populus deltoides | LI1A |  | RACiR 500 to 0 | 204.8554 | 24.1055 | 25.0065 | 999.969 | 0.0092853 | 0.41208 | 268.687 | 285.131
Populus deltoides | LI1A |  | RACiR 500 to 0 | 203.8629 | 23.8721 | 25.0067 | 1000.01 | 0.009278 | 0.41177 | 267.143 | 283.324
Populus deltoides | LI1A |  | RACiR 500 to 0 | 202.3219 | 23.626 | 25.0044 | 1000.01 | 0.0092705 | 0.41165 | 264.977 | 280.881
Populus deltoides | LI1A |  | RACiR 500 to 0 | 200.7085 | 23.4199 | 25.0071 | 1000.01 | 0.0092771 | 0.41191 | 262.785 | 278.444
Populus deltoides | LI1A |  | RACiR 500 to 0 | 199.6806 | 23.2273 | 25.0035 | 1000 | 0.0092768 | 0.41222 | 261.214 | 276.658
Populus deltoides | LI1A |  | RACiR 500 to 0 | 198.0616 | 22.9878 | 25.006 | 999.997 | 0.0092756 | 0.41201 | 259.001 | 274.17
Populus deltoides | LI1A |  | RACiR 500 to 0 | 196.5818 | 22.7042 | 25.0074 | 999.996 | 0.0092782 | 0.41201 | 256.792 | 271.642
Populus deltoides | LI1A |  | RACiR 500 to 0 | 195.001 | 22.6984 | 25.0126 | 999.993 | 0.0092804 | 0.4117 | 255.209 | 270.029
Populus deltoides | LI1A |  | RACiR 500 to 0 | 193.3503 | 22.4965 | 25.0111 | 999.988 | 0.0092775 | 0.41163 | 253.033 | 267.617
Populus deltoides | LI1A |  | RACiR 500 to 0 | 191.4565 | 22.4098 | 25.0057 | 1000.29 | 0.0092744 | 0.41186 | 250.848 | 265.309
Populus deltoides | LI1A |  | RACiR 500 to 0 | 190.5767 | 22.1492 | 25.0011 | 999.943 | 0.009268 | 0.41185 | 249.306 | 263.478
Populus deltoides | LI1A |  | RACiR 500 to 0 | 189.0495 | 21.9057 | 24.9969 | 999.944 | 0.0092636 | 0.41195 | 247.13 | 261.027
Populus deltoides | LI1A |  | RACiR 500 to 0 | 187.387 | 21.6941 | 24.9976 | 999.984 | 0.0092586 | 0.41163 | 244.952 | 258.596
Populus deltoides | LI1A |  | RACiR 500 to 0 | 186.1567 | 21.5243 | 24.9974 | 999.978 | 0.0092424 | 0.41089 | 243.372 | 256.82
Populus deltoides | LI1A |  | RACiR 500 to 0 | 184.1316 | 21.3962 | 25.0015 | 999.982 | 0.0092176 | 0.40953 | 241.164 | 254.443
Populus deltoides | LI1A |  | RACiR 500 to 0 | 182.7533 | 21.0726 | 24.993 | 999.939 | 0.009199 | 0.40944 | 238.958 | 251.871
Populus deltoides | LI1A |  | RACiR 500 to 0 | 181.3631 | 21.0256 | 24.9918 | 999.932 | 0.009201 | 0.40979 | 237.372 | 250.214
Populus deltoides | LI1A |  | RACiR 500 to 0 | 180.0985 | 20.6731 | 24.9935 | 999.966 | 0.0092053 | 0.41017 | 235.16 | 247.618
Populus deltoides | LI1A |  | RACiR 500 to 0 | 178.0605 | 20.6344 | 24.9996 | 1000.27 | 0.009216 | 0.41039 | 232.955 | 245.34
Populus deltoides | LI1A |  | RACiR 500 to 0 | 177.4104 | 20.2829 | 24.9996 | 999.926 | 0.0092172 | 0.41055 | 231.403 | 243.414
Populus deltoides | LI1A |  | RACiR 500 to 0 | 175.6519 | 20.1392 | 25.0002 | 999.926 | 0.0092256 | 0.41102 | 229.193 | 241.024
Populus deltoides | LI1A |  | RACiR 500 to 0 | 174.2439 | 19.8589 | 25.0047 | 999.962 | 0.0092468 | 0.41179 | 226.975 | 238.497
Populus deltoides | LI1A |  | RACiR 500 to 0 | 173.092 | 19.7048 | 25.0098 | 999.924 | 0.0092603 | 0.41201 | 225.396 | 236.743
Populus deltoides | LI1A |  | RACiR 500 to 0 | 171.572 | 19.4835 | 25.0065 | 999.926 | 0.0092837 | 0.41311 | 223.17 | 234.26
Populus deltoides | LI1A |  | RACiR 500 to 0 | 170.0869 | 19.1616 | 25.0078 | 999.951 | 0.0092734 | 0.41226 | 220.964 | 231.699
Populus deltoides | LI1A |  | RACiR 500 to 0 | 168.5072 | 19.1647 | 25.0062 | 1000.26 | 0.0092727 | 0.41221 | 219.362 | 230.072
Populus deltoides | LI1A |  | RACiR 500 to 0 | 166.7117 | 18.9851 | 25.0139 | 999.946 | 0.0092716 | 0.41148 | 217.175 | 227.672
Populus deltoides | LI1A |  | RACiR 500 to 0 | 164.7353 | 18.8911 | 25.0195 | 1000.26 | 0.0092691 | 0.41093 | 214.989 | 225.355
Populus deltoides | LI1A |  | RACiR 500 to 0 | 163.8578 | 18.6042 | 25.0216 | 999.914 | 0.0092624 | 0.41041 | 213.446 | 223.498
Populus deltoides | LI1A |  | RACiR 500 to 0 | 162.1636 | 18.3867 | 25.0169 | 999.94 | 0.0092431 | 0.40983 | 211.239 | 221.037
Populus deltoides | LI1A |  | RACiR 500 to 0 | 160.5826 | 18.1325 | 25.0208 | 999.895 | 0.0092374 | 0.4094 | 209.043 | 218.549
Populus deltoides | LI1A |  | RACiR 500 to 0 | 159.3615 | 17.9789 | 25.0259 | 999.928 | 0.0092335 | 0.40899 | 207.462 | 216.791
Populus deltoides | LI1A |  | RACiR 500 to 0 | 157.5758 | 17.8234 | 25.0351 | 1000.24 | 0.009244 | 0.40902 | 205.252 | 214.388
Populus deltoides | LI1A |  | RACiR 500 to 0 | 155.828 | 17.6596 | 25.0331 | 999.924 | 0.0092417 | 0.4091 | 203.048 | 211.989
Populus deltoides | LI1A |  | RACiR 500 to 0 | 154.3923 | 17.6219 | 25.0323 | 1000.22 | 0.0092405 | 0.40909 | 201.487 | 210.364
Populus deltoides | LI1A |  | RACiR 500 to 0 | 152.8596 | 17.3875 | 25.0339 | 1000.22 | 0.0092467 | 0.40918 | 199.332 | 207.942
Populus deltoides | LI1A |  | RACiR 500 to 0 | 151.4956 | 17.1069 | 25.034 | 999.881 | 0.0092651 | 0.40995 | 197.165 | 205.463
Populus deltoides | LI1A |  | RACiR 500 to 0 | 150.6304 | 16.8689 | 25.0333 | 1000.22 | 0.0092823 | 0.41068 | 195.619 | 203.654
Populus deltoides | LI1A |  | RACiR 500 to 0 | 149.027 | 16.6451 | 25.0299 | 999.874 | 0.0092814 | 0.41062 | 193.434 | 201.215
Populus deltoides | LI1A |  | RACiR 500 to 0 | 147.3327 | 16.4322 | 25.0335 | 1000.21 | 0.0092765 | 0.40996 | 191.246 | 198.774
Populus deltoides | LI1A |  | RACiR 500 to 0 | 146.0579 | 16.2958 | 25.0367 | 999.867 | 0.009261 | 0.40894 | 189.709 | 197.075
Populus deltoides | LI1A |  | RACiR 500 to 0 | 144.2915 | 16.1264 | 25.0384 | 1000.2 | 0.0092476 | 0.40822 | 187.555 | 194.715
Populus deltoides | LI1A |  | RACiR 500 to 0 | 142.8153 | 15.8458 | 25.0371 | 999.898 | 0.0092388 | 0.40793 | 185.376 | 192.218
Populus deltoides | LI1A |  | RACiR 500 to 0 | 141.4535 | 15.7792 | 25.036 | 1000.2 | 0.0092348 | 0.40788 | 183.822 | 190.574
Populus deltoides | LI1A |  | RACiR 500 to 0 | 139.9439 | 15.5333 | 25.0338 | 999.873 | 0.009236 | 0.40817 | 181.638 | 188.107
Populus deltoides | LI1A |  | RACiR 500 to 0 | 138.1511 | 15.3993 | 25.0334 | 999.898 | 0.0092326 | 0.40812 | 179.476 | 185.778
Populus deltoides | LI1A |  | RACiR 500 to 0 | 137.2331 | 15.1465 | 25.0371 | 999.865 | 0.0092309 | 0.40785 | 177.937 | 183.96
Populus deltoides | LI1A |  | RACiR 500 to 0 | 135.6421 | 14.924 | 25.0417 | 999.89 | 0.0092293 | 0.40755 | 175.787 | 181.553
Populus deltoides | LI1A |  | RACiR 500 to 0 | 134.1416 | 14.6678 | 25.0345 | 999.851 | 0.009219 | 0.40758 | 173.61 | 179.083
Populus deltoides | LI1A |  | RACiR 500 to 0 | 132.9988 | 14.4974 | 25.0313 | 1000.19 | 0.0092068 | 0.40728 | 172.042 | 177.318
Populus deltoides | LI1A |  | RACiR 500 to 0 | 131.2397 | 14.3236 | 25.0347 | 999.882 | 0.009196 | 0.40671 | 169.862 | 174.926
Populus deltoides | LI1A |  | RACiR 500 to 0 | 130.1255 | 13.922 | 25.0441 | 999.836 | 0.0092086 | 0.40688 | 167.712 | 172.343
Populus deltoides | LI1A |  | RACiR 500 to 0 | 129.0257 | 13.7793 | 25.0463 | 1000.18 | 0.0092272 | 0.40766 | 166.167 | 170.633
Populus deltoides | LI1A |  | RACiR 500 to 0 | 127.5989 | 13.5375 | 25.04 | 999.841 | 0.009251 | 0.40916 | 163.981 | 168.174
Populus deltoides | LI1A |  | RACiR 500 to 0 | 125.6515 | 13.4745 | 25.0333 | 1000.19 | 0.0092484 | 0.40937 | 161.814 | 165.911
Populus deltoides | LI1A |  | RACiR 500 to 0 | 124.4911 | 13.32 | 25.0329 | 1000.19 | 0.0092442 | 0.40906 | 160.271 | 164.187
Populus deltoides | LI1A |  | RACiR 500 to 0 | 122.8484 | 13.1013 | 25.0367 | 999.841 | 0.0092267 | 0.40789 | 158.146 | 161.808
Populus deltoides | LI1A |  | RACiR 500 to 0 | 121.3035 | 12.8454 | 25.0416 | 999.832 | 0.0092164 | 0.40705 | 156 | 159.368
Populus deltoides | LI1A |  | RACiR 500 to 0 | 120.229 | 12.6556 | 25.0414 | 999.867 | 0.0092046 | 0.40655 | 154.468 | 157.621
Populus deltoides | LI1A |  | RACiR 500 to 0 | 118.7633 | 12.3844 | 25.0419 | 999.823 | 0.009204 | 0.40655 | 152.294 | 155.14
Populus deltoides | LI1A |  | RACiR 500 to 0 | 117.6515 | 11.9785 | 25.0399 | 999.819 | 0.0092042 | 0.40673 | 150.132 | 152.537
Populus deltoides | LI1A |  | RACiR 500 to 0 | 116.2483 | 11.9181 | 25.0415 | 1000.17 | 0.0092036 | 0.40654 | 148.562 | 150.882
Populus deltoides | LI1A |  | RACiR 500 to 0 | 114.5762 | 11.707 | 25.0414 | 999.826 | 0.0091933 | 0.40602 | 146.364 | 148.436
Populus deltoides | LI1A |  | RACiR 500 to 0 | 112.692 | 11.5689 | 25.0449 | 1000.16 | 0.0091696 | 0.40475 | 144.185 | 146.082
Populus deltoides | LI1A |  | RACiR 500 to 0 | 112.01 | 11.2232 | 25.0417 | 1000.16 | 0.0091596 | 0.40464 | 142.628 | 144.151
Populus deltoides | LI1A |  | RACiR 500 to 0 | 110.6459 | 10.8908 | 25.0432 | 1000.15 | 0.0091488 | 0.40427 | 140.426 | 141.58
Populus deltoides | LI1A |  | RACiR 500 to 0 | 109.2958 | 10.5663 | 25.0403 | 1000.14 | 0.0091555 | 0.40498 | 138.184 | 138.978
Populus deltoides | LI1A |  | RACiR 500 to 0 | 108.0169 | 10.4636 | 25.0396 | 999.802 | 0.0091695 | 0.40572 | 136.571 | 137.239
Populus deltoides | LI1A |  | RACiR 500 to 0 | 106.4402 | 10.2394 | 25.0381 | 999.791 | 0.0091893 | 0.40669 | 134.338 | 134.749
Populus deltoides | LI1A |  | RACiR 500 to 0 | 104.7227 | 10.0785 | 25.0355 | 999.83 | 0.009211 | 0.40771 | 132.118 | 132.336
Populus deltoides | LI1A |  | RACiR 500 to 0 | 103.0874 | 10.124 | 25.0319 | 1000.13 | 0.0092133 | 0.40791 | 130.545 | 130.785
Populus deltoides | LI1A |  | RACiR 500 to 0 | 101.5648 | 9.8821 | 25.0264 | 1000.11 | 0.0092089 | 0.40781 | 128.393 | 128.356
Populus deltoides | LI1A |  | RACiR 500 to 0 | 99.8813 | 9.6868 | 25.0257 | 999.81 | 0.0091999 | 0.40725 | 126.221 | 125.953
Populus deltoides | LI1A |  | RACiR 500 to 0 | 99.0877 | 9.3809 | 25.0268 | 1000.11 | 0.0091924 | 0.40671 | 124.682 | 124.082
Populus deltoides | LI1A |  | RACiR 500 to 0 | 97.4785 | 9.1349 | 25.0292 | 999.762 | 0.0091715 | 0.40553 | 122.492 | 121.608
Populus deltoides | LI1A |  | RACiR 500 to 0 | 95.957 | 8.8766 | 25.0279 | 1000.11 | 0.009163 | 0.40523 | 120.307 | 119.129
Populus deltoides | LI1A |  | RACiR 500 to 0 | 94.9784 | 8.6523 | 25.0263 | 999.794 | 0.0091593 | 0.40524 | 118.745 | 117.317
Populus deltoides | LI1A |  | RACiR 500 to 0 | 93.2709 | 8.4697 | 25.0268 | 1000.1 | 0.0091569 | 0.40515 | 116.547 | 114.901
Populus deltoides | LI1A |  | RACiR 500 to 0 | 91.806 | 8.2083 | 25.0237 | 1000.1 | 0.0091629 | 0.40569 | 114.367 | 112.426
Populus deltoides | LI1A |  | RACiR 500 to 0 | 90.3805 | 8.1641 | 25.0251 | 1000.1 | 0.009163 | 0.40564 | 112.802 | 110.792
Populus deltoides | LI1A |  | RACiR 500 to 0 | 88.8553 | 7.9492 | 25.02 | 1000.1 | 0.0091655 | 0.40615 | 110.679 | 108.421
Populus deltoides | LI1A |  | RACiR 500 to 0 | 87.3255 | 7.7245 | 25.0196 | 999.784 | 0.0091602 | 0.40595 | 108.564 | 106.047
Populus deltoides | LI1A |  | RACiR 500 to 0 | 86.3723 | 7.4959 | 25.0163 | 1000.1 | 0.009146 | 0.40544 | 107.042 | 104.271
Populus deltoides | LI1A |  | RACiR 500 to 0 | 84.979 | 7.1761 | 25.0147 | 999.778 | 0.0091141 | 0.4041 | 104.878 | 101.747
Populus deltoides | LI1A |  | RACiR 500 to 0 | 83.4493 | 6.9274 | 25.0172 | 1000.09 | 0.0091136 | 0.4041 | 102.691 | 99.2779
Populus deltoides | LI1A |  | RACiR 500 to 0 | 82.3577 | 6.7631 | 25.0192 | 1000.08 | 0.0091259 | 0.40466 | 101.14 | 97.538
Populus deltoides | LI1A |  | RACiR 500 to 0 | 80.6421 | 6.6169 | 25.0211 | 1000.08 | 0.0091517 | 0.40572 | 98.9769 | 95.1985
Populus deltoides | LI1A |  | RACiR 500 to 0 | 79.1213 | 6.3875 | 25.0192 | 1000.08 | 0.00916 | 0.40603 | 96.8377 | 92.7953
Populus deltoides | LI1A |  | RACiR 500 to 0 | 77.9259 | 6.2662 | 25.015 | 999.765 | 0.0091589 | 0.40609 | 95.3099 | 91.1225
Populus deltoides | LI1A |  | RACiR 500 to 0 | 76.6619 | 5.9327 | 25.0145 | 1000.07 | 0.0091547 | 0.40573 | 93.2007 | 88.646
Populus deltoides | LI1A |  | RACiR 500 to 0 | 75.1396 | 5.6996 | 25.0128 | 999.765 | 0.0091418 | 0.40507 | 91.086 | 86.2636
Populus deltoides | LI1A |  | RACiR 500 to 0 | 74.2285 | 5.4508 | 25.0174 | 1000.07 | 0.0091279 | 0.40401 | 89.5705 | 84.473
Populus deltoides | LI1A |  | RACiR 500 to 0 | 72.6575 | 5.2272 | 25.0179 | 1000.06 | 0.0091143 | 0.40337 | 87.4249 | 82.0677
Populus deltoides | LI1A |  | RACiR 500 to 0 | 71.1795 | 4.9654 | 25.0167 | 1000.07 | 0.009109 | 0.40328 | 85.2587 | 79.606
Populus deltoides | LI1A |  | RACiR 500 to 0 | 70.1316 | 4.7827 | 25.012 | 1000.07 | 0.0091077 | 0.40363 | 83.7166 | 77.8546
Populus deltoides | LI1A |  | RACiR 500 to 0 | 68.5573 | 4.5624 | 25.0127 | 1000.07 | 0.0091116 | 0.40382 | 81.5487 | 75.435
Populus deltoides | LI1A |  | RACiR 500 to 0 | 67.3421 | 4.1924 | 25.0139 | 1000.06 | 0.0091074 | 0.40356 | 79.3861 | 72.8638
Populus deltoides | LI1A |  | RACiR 500 to 0 | 66.0301 | 4.0997 | 25.0176 | 1000.05 | 0.0091059 | 0.40327 | 77.8202 | 71.1812
Populus deltoides | LI1A |  | RACiR 500 to 0 | 64.6211 | 3.8084 | 25.0134 | 1000.05 | 0.0091062 | 0.40358 | 75.6399 | 68.6764
Populus deltoides | LI1A |  | RACiR 500 to 0 | 62.5235 | 3.7869 | 25.0108 | 999.749 | 0.0090821 | 0.40265 | 73.4621 | 66.4396
Populus deltoides | LI1A |  | RACiR 500 to 0 | 61.589 | 3.557 | 25.007 | 1000.05 | 0.0090732 | 0.40259 | 71.929 | 64.6542
Populus deltoides | LI1A |  | RACiR 500 to 0 | 60.1182 | 3.2985 | 25.0087 | 1000.04 | 0.0090547 | 0.40184 | 69.7903 | 62.2197
Populus deltoides | LI1A |  | RACiR 500 to 0 | 58.782 | 2.9839 | 25.0131 | 1000.05 | 0.0090597 | 0.40206 | 67.6277 | 59.7094
Populus deltoides | LI1A |  | RACiR 500 to 0 | 57.7215 | 2.8046 | 25.0166 | 999.742 | 0.0090708 | 0.4025 | 66.0846 | 57.9629
Populus deltoides | LI1A |  | RACiR 500 to 0 | 56.074 | 2.6096 | 25.0159 | 1000.05 | 0.0090827 | 0.40317 | 63.898 | 55.5466
Populus deltoides | LI1A |  | RACiR 500 to 0 | 54.9945 | 2.1919 | 25.0117 | 1000.05 | 0.0090921 | 0.40386 | 61.7359 | 52.9316
Populus deltoides | LI1A |  | RACiR 500 to 0 | 53.7731 | 2.0694 | 25.008 | 1000.05 | 0.0090943 | 0.40412 | 60.1761 | 51.2277
Populus deltoides | LI1A |  | RACiR 500 to 0 | 52.3972 | 1.7568 | 25.0051 | 1000.04 | 0.009085 | 0.40377 | 57.9901 | 48.6894
Populus deltoides | LI1A |  | RACiR 500 to 0 | 50.6627 | 1.5914 | 25.0038 | 1000.04 | 0.0090849 | 0.4038 | 55.8014 | 46.3056
Populus deltoides | LI1A |  | RACiR 500 to 0 | 49.8212 | 1.3107 | 25.0006 | 1000.03 | 0.0090818 | 0.40385 | 54.2368 | 44.4323
Populus deltoides | LI1A |  | RACiR 500 to 0 | 48.5099 | 0.9643 | 25.0034 | 1000.03 | 0.0090814 | 0.40361 | 52.0302 | 41.8423
Populus deltoides | LI1A |  | RACiR 500 to 0 | 47.0116 | 0.68823 | 25.0012 | 1000.02 | 0.0090848 | 0.40385 | 49.8047 | 39.3084
Populus deltoides | LI1A |  | RACiR 500 to 0 | 45.8766 | 0.51617 | 25.0035 | 999.717 | 0.0090811 | 0.40347 | 48.2148 | 37.5161
Populus deltoides | LI1A |  | RACiR 500 to 0 | 44.2476 | 0.29615 | 24.9976 | 1000.02 | 0.0090841 | 0.40396 | 45.9954 | 35.0471
Populus deltoides | LI1A |  | RACiR 500 to 0 | 42.6356 | 0.067269 | 24.9975 | 1000.02 | 0.009082 | 0.40378 | 43.774 | 32.5588
Populus deltoides | LI1A |  | RACiR 500 to 0 | 41.1574 | 0.037673 | 24.9969 | 1000.02 | 0.0090803 | 0.40374 | 42.1879 | 30.9178
Populus deltoides | LI1A |  | RACiR 500 to 0 | 39.5127 | -0.16243 | 24.9968 | 1000.03 | 0.0090787 | 0.40363 | 40.0045 | 28.5005
Populus deltoides | LI1A |  | RACiR 500 to 0 | 37.8759 | -0.35473 | 25.0005 | 1000.02 | 0.0090755 | 0.40323 | 37.8483 | 26.1198
Populus deltoides | LI1A |  | RACiR 500 to 0 | 37.1242 | -0.6547 | 24.9995 | 1000.02 | 0.0090728 | 0.40309 | 36.3266 | 24.2742
Populus deltoides | LI1A |  | RACiR 500 to 0 | 35.7024 | -0.93533 | 24.9948 | 1000.02 | 0.0090668 | 0.40306 | 34.1677 | 21.7976
Populus deltoides | LI1A |  | RACiR 500 to 0 | 34.1056 | -1.151 | 24.9904 | 1000.01 | 0.0090692 | 0.40345 | 31.9957 | 19.3747
Populus deltoides | LI1A |  | RACiR 500 to 0 | 33.1322 | -1.3747 | 24.9944 | 999.7 | 0.0090661 | 0.40306 | 30.4365 | 17.5642
Populus deltoides | LI1A |  | RACiR 500 to 0 | 31.4098 | -1.5457 | 25.0032 | 999.997 | 0.009083 | 0.40321 | 28.2482 | 15.1784